## Slide 1: R. Eichhorn, Cornell
CLASSE facilities are operated by the Cornell Laboratory for Elementary Particle Physics (LEPP) and the Cornell High Energy Synchrotron Source (CHESS) with major support from the National Science Foundation.
Performance (and improvements) of the Cornell ERL injector cryomodule

## Slide 2: The Cornell ERL Injector Module
1.3 GHz RF cavity
HOM beamline absorber at 80K
between cavities
Frequency tuner
Twin Input Coupler
HGRP system with 3 sections
15 feet
Number of 2-cell cavities 5
Acceleration per cavity 1 – 3 MeV
Accelerating gradient 4.3 – 13.0 MV/m
R/Q (linac definition) 222 Ohm
Qext 4.6104 – 4.1105
Total 2K / 5K / 80K loads: 30W / 60W / 700W
Number of HOM loads 6
HOM power per cavity 40 W
Couplers per cavity 2
RF power per cavity 120 kW
Amplitude/phase stability 10-4 / 0.1 (rms)
ICM length 5 m
CW operation of the Cornell SRF injector

## Slide 3
ERL Injector Prototype:
Achievements to date (last high current run Sept’13):
75 mA average current @ 4 MeV
0.3 mm emittance @ 77 pC, 8 MeV
Cornell ERL Injector

## Slide 4
High Currents – Na2KSb
Using a Na2KSb photocathode, ran over 8 hours at 65 mA (2000 C) with a 2.6 day 1/e cathode lifetime.
Reached as high as 75 mA for a short time.
Front surface of the cathode (CsK2Sb on Si) after use.
Active area is offset from the center
Ion damage limited to the central area

## Slide 5
Horizontal Phase Space
x X 1000
data
GPT
x (mm)
Vertical Phase Space
0 pc
19 pc
77 pc
y X 1000
data
GPT
y (mm)
Projected Emittance for 19 (77) pC @ 8MeV:
Data Type | en(100%) [microns] | en(90%) [microns]
Projected (EMS) | 0.33(0.69) | 0.23(0.51 )
GPT | 0.31 (0.72) | 0.19(0.44)
Data Type | en(100%) [microns] | en(90%) [microns]
Projected (EMS) | 0.20(0.40) | 0.14(0.29)
GPT | 0.16(0.37) | 0.11(0.25)
Emittance Results
Vertical
Horizontal

## Slide 6
Beam
Coupler antennas only in x-direction!
SRF Cavity Couplers
No dipole kick, but we see ‘quadrupole’ focusing

## Slide 7
Simulation with Merger
Optimize emittance using a using a genetic algorithm
In the merger, the emittance asymmetry can be (mostly) removed, using the asymmetry provided by the dipoles
Also, optimizations show that a beam energy of 12-13 MeV is needed to maintain minimum emittance in x/y before injection into the main linac
Simultaneous good emittance!

## Slide 8: Performance beyond our parameters
Q (pC) | Ipeak Target (A) | Ipeak (A) | εn Target (95%, mm) | εn (95%, mm) | εn,th /εn
20 | 5 | 5 | 0.25 | H: 0.18, V: 0.19 | 58%
100 | 10 | 11.5 | 0.40 | H: 0.32, V: 0.30 | 80%
300 | 30 | 32 | 0.60 | H: 0.62, V: 0.60 | 70%

## Slide 9
Parameter | Metric | Status | Notes
Average Current | 100 mA |  | 75 mA (1300 MHz)
Bunch Charge | 77 pC |  | Pulsed mode (50 MHz)
Energy | 5 to 15 MeV |  | 14 MeV max (due to cryo limits)
Laser Power | > 20 W |  | > 60 W at 520 nm (1300 MHz)
Laser Shaping | beer can dist. |  | Adequate for now
Gun Voltage | 500 kV |  | Currently operating at 350 kV
Emittance | < 2 mm (norm, rms) |  | Ultimate ERL goal 0.3 mm, with merger
Operational Lifetime | > 1 day |  | Recent improvements with new cathodes
Injector Achievements

## Slide 10
We found that the couplers get quite warm during high power operations.
Temperatures reach equilibrium (150 K) in 4 hours
Booster Coupler Heating (80 K system)
The flow to the HOM’s and couplers are in parallel, but we discovered the coupler tubing is too small. So the HOM’s get more flow, and the couplers less than designed for.

## Slide 11
Analysing the findings
Design parameters:
  Coupler heat load 50 W
  HOM heat load 50 W
Findings:
  Coupler: 75 W
  HOM 5-10 W
  In addition: some design flaws in choosing piping
The issue:
  HOM absorbers and couplers are cooled in parallel
  Mass flow ratio was not 1:1 but rather 1:3 in favor of the HOM absorber

## Slide 12: Understanding the coupler heating in the ICM
Supply
Return
HOM absorbers
Couplers
80 K jumper system
Added heat to the HOM branch

## Slide 13
Understanding the coupler heating in the ICM
Q1
Q2
Q1 = 200 W
Q2 = 200-400W
=> The cooler channel steals the flow

## Slide 14: Fluid dynamics calculations
*
Calculate pressure drop and heat exchange rate
Calculate flow properties
Update pipe temperature
Iterate, record, reset with new pipe temp.

## Slide 15: Recalculation of the 80 K system
Pipe | Inner Diameter | Length | Mass
HOM | 5.94 mm | 68.6 cm | 200 g
Coupler | 3.88 mm | 83.8 cm | 200 g
Inlet | 2.00 mm | 50.0 cm | N/A
Regime | HOM Heat Load | Coupler Heat Load | Input Pressure | Mass Flow Rate | Initial Fluid Temp
No inlet | 5 W | 50 W | 20 bar | 9 g/s | 80 K
Inlet | 5 W | 50 W | 20 bar | 4.5 g/s | 80 K
High Heat (inlet) | 5 W | 120 W | 20 bar | 6 g/s | 80 K
HOM absorber
Coupler
Flow restrictor pipe
Inlet flow restrictor: 2.3 mm, 87 cm long

## Slide 16: Inlet flow restrictor

## Slide 17
HOM inlet flow restrictor, 120 W heat load on couplers
Including the inlet pipe allows more cryogen to flow through the couplers, greatly improving system performance

## Slide 18: ICM rebuilt 2015/16
Cooling pipes
Adding the flow restrictor took only 2 days. However, disassembling the module and reassembling it after the modification was 6 months of labour.

## Slide 20: Typical ICM cool-down

## Slide 21: ICM heat exchanger can modifications
Cool-down lines added,
separation of the 5K and the 80K mass-flow

## Slide 22: Outlook
ICM has been rebuilt and improved
Achieving 100 mA should be feasible
Cooldown happened June 2016
Installation of beamlines continue
Processing of the DC gun is in progress
Beam operations to resume in August

## Slide 23
Thanks