## Slide 1: Retreatment of Superconducting Cavities for the European XFEL
Detlef Reschke – DESY for the EU-XFEL Accelerator Consortium
Notes: How to edit the title slide
Upper area: Title of your talk, max. 2 rows of the defined size (55 pt)
Lower area (subtitle): Conference/meeting/workshop, location, date, your name and affiliation, max. 4 rows of the defined size (32 pt)
Change the partner logos or add others in the last row.

## Slide 2: VT-CM comparison: USABLE GRADIENT
Vertical Test – min of
Maximum gradient (quench, RF power)
FE limit (top/bottom X-ray)
Q0 limit (= 1010)
Cryomodule – min of
Maximum gradient (quench)
FE limit (front/back X-ray)
RF power limit at 31 MV/m
Acceptance criteria for “Usable gradient” in vertical test
  INITIAL: > 26 MV/m (10% margin to required average design operating gradient)
  NOW (after analysis of retreatment results in May 2014): > 20 MV/m (for optimized number of retreatments and retests)
Notes: No update necessary

## Slide 3: Definition of “Usable Gradient” in VT
Notes: No update necessary

## Slide 4
Results: Usable Gradient “As received”
re-treat & retest
accept for module
Both vendors well above Spec
RI shows ~ 3MV/m in average more than EZ: a) final EP b) low gradient quenches at EZ
Several cavities with < 20 MV/m accepted “as received”, especially if a) limitation = “bd“ + b) no FE
Average Q-values for both vendors: Q0(4 MV/m) = 2.1·1010 Q0(23.6 MV/m) = 1.3·1010
Missing ~90 cavities “as received”?
Notes: updated

## Slide 5
Why cavities w/o “as received” test? (“as received”: first test after delivery to DESY + cavity conform to Spec)
Three main categories for retreatment (rework):
  Non-Conformities after delivery from vendor (~ 90 Tests) (mechanical, vacuum, defects, …) => retreatment/rework at DESY or vendor depending on NC-type
  Performance (~160 Tests) (acceptance criteria for gradient or Q-value not met) - mainly by field emission (+ “Quench”, “low Q”) => first + successful retreatment High Pressure Ultrapure Water Rinse <Eacc, usable>: 19 MV/m => 26 MV/m <Q0(4 MV/m)>: 2.1 · 1010 => 2.4 · 1010
  Non-Conformity during string assembly (21 Cv)
Retreatments
Notes: - Categories for test purpose + formalized decisions according to test result

## Slide 6
Now performance-related retreatments, only
In the following comparison analysis, "before" and "after" refer to a pair of vertical tests for a given cavity before and after retreatment
Analysis limited to "before" tests being "as received" (i.e. no second pass treatments are included)
129 data sets for “HPR only” shown
Retreatment results: HPR

## Slide 7
Retreatment performance: HPR Max gradient

## Slide 8
Retreatment performance: HPR Usable gradient

## Slide 9
Retreatment performance: HPR Q-value at 4 MV/m

## Slide 10
Retreatment performance: HPR Q-value at 23.6 MV/m

## Slide 11
Now performance-related retreatments, only
In the following comparison analysis, "before" and "after" refer to a pair of vertical tests for a given cavity before and after retreatment
Analysis includes second, third,… pass treatments ("before" tests being "as received“
18 data sets available
Retreatment results: BCP+HPR+120C bake

## Slide 12
Based on 18 tests
Retreatment results: BCP+HPR+120C bake Usable Gradient

## Slide 13
Based on 18 tests
No improvement on Q-value!
Retreatment results: BCP+HPR+120C bake Q-value at 4 MV/m
Mean
RMS

## Slide 14: Thanks to all colleagues of:
- IFJ-PAN Krakow (esp. J. Swierblewski, M. Wiencek)
- CEA Saclay
- INFN Milano (esp. L. Monaco)
- E. Zanon
- Research Instruments
- Daher Transkem
- Alsyom
- DESY (esp. V. Gubarev, J. Schaffran, L. Steder, N. Walker, M. Wenskat, S. Yasar)

## Slide 15: :
- Back-up
xxxxxxxx

## Slide 16: Cavity Surface Preparation
Two schemes for the final surface treatment: - E. Zanon: Final 10µm BCP (“BCP Flash”) - Research Instr.: Final 40µm EP
Successful mechanical production and surface preparation at both vendors!
No performance guarantee resulted in DESY taking responsibility for:
  the risk of unexpected low gradient or field emission
  retreatment (good cooperation/agreements with both vendors)
Final EP
BCP Flash

## Slide 17: Transportation + Incoming Inspection
17
Transportation:
  “Cavity ready for test” requires well-defined transport concept
  Transport under vacuum => avoid particle transport
  Dedicated boxes for horizontal transport by truck (Vendor => DESY => Saclay)
  No performance degradation observed
Incoming inspection at DESY
  Basic mechanical, electrical, RF checks + final vacuum leak check before test
  Idea: Check for damages during transport
  Found: - Assembly errors + contaminations - Mechanical + electrical damages - Few leaks
  Incoming Inspection is mandatory
=> 54 Cv‘s back to vendor
Notes: Remark:
Additional and comprehensive information about NC‘s + experiences with QC concept during production extend this talk => aware that this is missing

## Slide 18: Vertical tests at AMTF

## Slide 19: Statistics of Cavities + Vertical Acceptance Tests
19
Cavity production finished with last delivery in Mar 2016:
  (800 + extra 4) Series Cavities
  24 ILC “HiGrade”-Cavities (w/o He-tank; QC)
    8 since equipped with He-tank and used for XFEL
  16 Cavities for infrastructure commissioning
    4 since used for XFEL
    => 816 cavities available and vertically tested for XFEL
Analysis of vertical acceptance tests includes
  Series Cavities
  ILC “HiGrade”-Cavities (w/o He-tank)
  NO infrastructure commissioning tests
Stable average vertical test rate ~40 tests/month achieved
1225 vertical tests (few additional tests due to returns from string assembly possible)
Notes: - updated

## Slide 20: Analysis of Cavity Results
20
Fully statistical approach; (nearly) no individual cavity analysis => simplified vertical test procedure at 2K: - Q0(Eacc), - fundamental mode spectra, - HOM-check (partially)
Work in progress: Analysis and possible correlations for Eacc,max, Q-value + x-rays
  Nb-material: Vendor dependency
  Cavity with vs. without He-tank
  Surface treatment (not vendor)
  Dependence on test infrastructure (RF test stand, cryostat, test inserts, …)
  Cool down procedure => no special procedure applied, bad instrumentation, …
  “Processing” vs. degradation in case of field emission; degradation after quench
  Optical inspection, local repair
  …
  Error analysis
WEPMB007
Notes: - updated

## Slide 21
Trend of Usable Gradient “As received”
Good stability of average usable gradient over full production period

## Slide 22: Finally: “Send to Saclay“
22
(nearly) all cavities shipped to CEA Saclay for string assembly
average usable gradient: ~ 30 MV/m
includes cavities for three modules below 20 MV/m
  Mainly low gradient quenches => retreatment not successful => retreatment would run out of schedule
  Consequence of missing performance guarantee
> 810
Notes: - updated