## Slide 1: Carrousel Biennale 2016 Pise
Didier Laporte
ATRIUM-130278

## Slide 4: Filter Exchange System – Découpage
Loader
Auto-Changer
Focal Plan
Cryostat
Filter Control System
Shutter
Carousel
On-line Filter
LPNHE
LPSC
CPPM
APC

## Slide 5: Spécifications
Durée de vie : 15 ans (C-EXCH-001)
Un filtre doit pouvoir être changer sans démonter la caméra du télescope (C-EXCH-096)
Le système changeur de filtre doit accepter des filtres de masse différente (30 à 44 kgs) (C-EXCH-009)
Un filtre doit être changer en moins de 89 secondes (C-EXCH-013)
Température de fonctionnement : -5°C à 30 °C (C-EXCH-143)
Température de survie : -15°C à 40° C (C-EXCH-144)
Le carrousel doit peser moins de 250 kgs (C-EXCH-046)
La position du centre de gravité du carrousel est contrainte (C-EXCH-047)
Le carrousel doit avoir 5 positions de filtre disponible (C-EXCH-049)
Le carrousel doit pouvoir faire 35 000 rotations (C-EXCH-050)
Le carrousel doit pouvoir verrouiller et déverrouiller les filtres 30 000 fois (C-EXCH-106)
Sismique
Sécurité
Maintenance

## Slide 8: Carousel – maintien filtre
Ball joint
Sliding ball joint
Point
Z

## Slide 9: Carousel – mécanisme de déverouillage
Design baseline
  Fixe actuator on the flange
    Detailed design done
    Maintenance: entire actuator replacement (2 screws + 1 connector)
  Magnetic coupling transmit power on the carousel
    Misalignment allowed → successfully tested
    0.6 Nm magnetic coupling can pull 80N thru a reducer
  Unlocking force transmitted to clamps by cable

## Slide 10: Guidage, slip-ring et moteur/freins

## Slide 11: Le carrousel dans le temps
Fin 2006 : Discussion avec Kirk Gilmore des contributions techniques IN2P3 à la caméra LSST
01/2007 : R&D carousel au LPNHE
06/2010 : Totalité du changeur de filtre à l’IN2P3
12/2012 : Installation du single filter test au CPPM
02/2015 : Commandes des 2 rings
04/2015 : Revue validation design  des points à éclaircir
10/2015 : Delta Revue validation design
05/2016 : Arrivée au LPNHE du dummy backflange
08/2016 : le carrousel commence à tourner
09/2016 : Revue validation calculs

## Slide 12: Validation de la conception : FEA
Pour valider certaines conceptions, nous avons besoin de faire des calculs.
Jusqu’à fin septembre 2015, les calculs étaient réalisés au laboratoire,
01/2016 : négociation avec l’équipe de SLAC pour faire les calculs du carrousel
09/2016 : Présentation des calculs et clôture des réserves sur la conception (RFA)

## Slide 13: Carousel – Structure Analysis: Natural modes
Natural frequencies
  Filters (simplified)
    Around 25 Hz
  Structure
    Higher than 44 Hz
  Ring gear
    Around 45 Hz
Mode | Frequency [Hz] | 
1 | 24.723 | Filter bending
2 | 24.848 | Filter bending
3 | 29.634 | Filter bending
4 | 29.731 | Filter bending
5 | 33.335 | Filter bending
6 | 44.573 | 
7 | 49.206 | 
8 | 49.466 | 
9 | 54.697 | 
10 | 55.267 | 

## Slide 14: Carousel FEA Model
Full Carousel assembly modeled including THK Rail and carriages, bearing stiffnesses, tubes, welds, fasteners, pins
Every bolted and welded joint is modeled with a rigid element connecting coincident nodes.
Since the full camera model is used, all of the interface loads into the Carousel are realistic, with realistic interfaces, stiffnesses and masses
Carousel FEA model incorporates all of the latest design features, including the 3mm filter offset
Total Carousel model represents approximately 125K DOF (~15% of Integrated model)
Camera Body Housing
Camera Body Shroud
L1-L2 Assy
Utility Trunk
Back Flange
By SLAC team

## Slide 15: Carousel FEA Model
THK Rail
Motors and Brakes
Back Flange
Rotating Assembly
Carriages (5)
Due to the rotating parts of the carousel, many structural connections require separate local coordinate systems (very critical point)
In the whole camera, the carousel has the most complex connections and load paths of any structure
Pentagon Rods (5)
Support Posts (5)
Synchro Plates (10)
Shear pins
Aluminum Ring
Ring Gear
By SLAC team

## Slide 16: Carousel Dynamic Modes
Camera Mode2@ 24.8 Hz
Carousel Lateral X
Camera Mode18@ 37.17 Hz
Carousel Rotation
By SLAC team

## Slide 17: Camera Body Deflections under 1GX
X Component
dmax=747mm
@ Filter
Unsymmetric loading results in greatest displacements
Y Component
dmax=628mm
@ Filter
Symmetric Loading
Z Component
dmax=30.9mm
@ Filter
Gravity in X direction
Gravity in Y direction
Gravity in Z direction
Min distance from filter tip to cryostat = 35.6 mm
By SLAC team

## Slide 18: Allocations: natural frequency
Although it does not strictly meet subsystem requirement, it meets integrated camera requirement.
After extensive parametric analyses, increasing the lateral mode is difficult with the dead weight of four filters
By SLAC team

## Slide 19: Stresses due to Survival Seismic Loads
Stress risers at clamp interfaces; 57.8 MPa
Pentagon tube fittings see maximum stress of 68.2 MPa
Significant shear load in Z-coupling plate; 39.7MPa
Load paths are not very efficient due to space constraints (posts take significant moment loads)
The stress fringes shown are maximum stresses from combined survival seismic inertial loads, still healthy margins even with no-test safety factors.
Fty=241MPa
Ftu=290MPa
By SLAC team

## Slide 20: Interface Forces
The forces between the Carousel joints are outputted for the bolted joint analysis
  Every bolted joint is modeled, so bolt forces can be easily extracted and used in the bolt calculation (for preload against gapping and strength calculations)
  All fastener interfaces are designed to be robust and redundant, with high margins of safety
  The vector components of the interface forces are combined in linear superposition to derive the worst case forces for all of the static load cases described above
329N Pentagon tube pin shear @ 1gZ
2390N Post fastener prying force @ 1gX

## Slide 21: Weld calculations Line load calculation
By SLAC team

## Slide 22: 1 – Support Tubes
Support tubes: Investigate steel posts for the support tubes, welded to or bonded to the end fittings, to simplify the design and allow design issues to be resolved quickly.
After the DFDR, the posts were changed to welded steel tubes
The analysis reflects the current design (5mm tube wall)
The current design supports the stiffness and strength requirements
By SLAC team

## Slide 23: 2 – Pentagon Rods
Pentagon rods: Finalize the design of pentagon rod connections of the tubes to end fittings (welded or bonded) and their bolted connections to corner brackets.
The pentagon rod design is complete, being aluminum tube with welded end fittings
The analysis reflects the current design
The current design supports the stiffness and strength requirements
Fastener head pull through
MS = High
By SLAC team

## Slide 24: 8 – Ring gear analysis write up
Complete and write up ring gear analysis: show deflection of final ring gear design and effect on pinion gear teeth meshing, given the non-standard mounting method and design geometry.
Updated analysis shows the ring gear macro deformations are insignificant under normal operation and will maintain correct gear mesh during operation
Detailed stress analysis of tooth and contact was deemed to be the same as in the previous analysis, so further analysis was not needed
1gX
1gY
1gZ
Max Disp = 16.2 mm
Diametral Dev. = ~1.5 mm
Max Disp = 59.8 mm
Diametral Dev. = ~2.5 mm
Max Disp = 79.0 mm
Diametral Dev. = ~3.8 mm
Notches do not cause any significant loss in ring gear circularity
Displacements shown at 100X
By SLAC team

## Slide 25: Pendant ce temps au LPNHE

## Slide 26: 9 – THK Rail Stiffness
THK rail stiffness: run tests to empirically determine rotational stiffness of the carriages on the rails. Use this to update the FEA model.
This has been answered in: 4-FEA boundary conditions, i.e.
Stiffness testing of the THK Rail and carriage was performed, and the FEA model was correlated to the test results
We have a high confidence on the fidelity of the FEA model stiffness at this critical interface
THK Rail
Position indicators
Loads
Chart: Load [N.m]
Model | 0 | 0.0335 | 0.058213 | 0.081481 | 0.104 | 0.14774 | 0.19043 | 0.31512 | 0.43708 | 0.55749 | 0.67685 | 0.79547
By LPNHE team

## Slide 27: 10 – Pin Flexure Testing
Pin flexure testing: test case-hardened pin to define allowable loads to be used for the pin. Don’t rely on the analysis to develop this, since stress margins were very small and material properties of the case-hardening were not clear.
This has been answered in 7-Pin Analysis, i.e.
Pin analysis was deemed impractical. Instead a strength test simulating the worst case pin loading was performed and shown to have significant margin of safety
Load cell
Test pin
Position indicators
Chart: Load [N]
| 12 | 440 | 561 | 681 | 801 | 922 | 1042 | 1239 | 1435 | 1799 | 2163 | 2528 | 2892 | 3256 | 3366 | 3422 | 3532 | 3652 | 3756 | 3867 | 3932 | 4036 | 4147 | 4214 | 4323 | 4434 | 4538 | 4930
Loaded Angular | -3.125e-05 | 0.0024438 | 0.0028 | 0.0031187 | 0.00365 | 0.0041938 | 0.0046063 | 0.0051687 | 0.0062125 | 0.0073625 | 0.0089625 | 0.010725 | 0.011719 | 0.014581 | 0.015713 | 0.013537 | 0.015188 | 0.014481 | 0.014531 | 0.014837 | 0.017306 | 0.01605 | 0.015956 | 0.017163 | 0.021406 | 0.022713 | 0.022719 | 0.025731 | 0
Unloaded Angl. | -2.5e-05 | 5e-05 | 5e-05 | 3.75e-05 | 3.75e-05 | 3.125e-05 | 2.5e-05 | 3.75e-05 | 3.125e-05 | 3.75e-05 | 3.75e-05 | 3.125e-05 | 3.75e-05 | 0.00013125 | 0.000175 | 0.0001875 | 0.000275 | 0.00028125 | 0.0002875 | 0.00029375 | 0.0004 | 0.00041875 | 0.0004625 | 0.0025688 | 0.0027563 | 0.00325 | 0.0034188 | 0.0055687 | 0
Loaded Vert. | 0.00225 | 0.04975 | 0.058 | 0.06675 | 0.075 | 0.08325 | 0.09075 | 0.10175 | 0.1145 | 0.1295 | 0.1465 | 0.1615 | 0.17725 | 0.19375 | 0.2 | 0.1995 | 0.21 | 0.20975 | 0.21325 | 0.2185 | 0.22525 | 0.228 | 0.23175 | 0.237 | 0.26275 | 0.2755 | 0.27875 | 0.29875 | 0
Unloaded Vert. | 0 | 0.0025 | 0.0025 | 0.0035 | 0.003 | 0.00325 | 0.003 | 0.003 | 0.00325 | 0.0035 | 0.004 | 0.00375 | 0.004 | 0.00575 | 0.006 | 0.0065 | 0.007 | 0.00725 | 0.008 | 0.00775 | 0.0085 | 0.00925 | 0.009 | 0.02025 | 0.02125 | 0.024 | 0.02575 | 0.03875 | 0
No visible permanent deformation before 3000N
No rupture before 5000N (> 1655N of the test plan)
By LPNHE team

## Slide 28: Conclusions FEA
A comprehensive system structural analysis of the Carousel Assembly has been completed
The current design is proven to have adequate margins of safety against stress, and meets stiffness criteria
Analysis issues explained in the Tier 1 and Tier 2 RFA’s have been addressed and are found to not be an issue
All RFA’s related to the structural analysis of the Carousel can now be closed
By SLAC team

## Slide 29: Le futur du carrousel
24/10/2016 : Bâti (LPC Clermont) dispo dans le hall (LPNHE)
10/2016 : Unbalanced test
10/2016 : Début commande pieces finales à Paris
10/2017 : backflange final à Paris pour intégration
04/2018 : Intégration finale à SLAC
12/2019 : Opérationel au Chili