## Slide 1
Discussion on the integrator specifications
with comparison of the 3 options
Debriefing meeting at LPC , 28 April 2016
François Vazeille
Summary of the TileCal Calibration systems in the upgrade scheme.
TileCal data for the Cesium source moving.
TileCal data for the Luminosity measurements.
Chicago integrator.
Argonne integrator.
Clermont-Ferrand integrator.
Summary of specifications.

## Slide 2
 Summary of the TileCal Calibration systems
in the upgrade scheme
4 complementary systems acting at various levels of the whole electronic chain
+ likely a 5th system for the Back end electronics not shown here.
1. At the Tile/Fiber level: Cesium radioactive source and p-p Minimum Bias events.
2. At the Light Mixer/PMT level: Laser.
3. At the Very Front End electronics level: Charge Injection System.
4. At the Front End Board digital level: Digital tests.
1. PMT gain adjustment/calibration and long term monitoring,
plus calibration transport from the ATLAS Test Beam.
2. Short term monitoring and calibration.
3. Electronics calibration in pC.
4. Working tests of the Main Board/Daughter board communication.

## Slide 3
PMT
Divider
FATALIC
All-in-One
Main board
Daughter
Board
Back End
electronics
PMT Block
Drawer
Low
Voltage
High
Voltage
Mixer
Tile and Fiber
Minimum Bias events
Cesium source
Laser
Charge injection system
Digital tests

## Slide 4
 TileCal data for the Cesium source moving
Source speed: 30 cm/s.
TileCal scheme
Steel Master = 5.0 mm
Steel Spacer = 4.05 mm
Plastic Tile = 3 mm
Period = 18.325 mm
= 5 + 4.05 + 5 + 4.05 + G
= 18.10 + G G = Glue
18.325
- Transit time:
In a Tile: 0.3/30 = 0.01 s = 10 ms.
From a Tile to the following one: 1.8325/30 = 61.08 ms.
- The optimum integration time is 10 ms,
that can be increased by two means:
- Digital sum without any problem.
- Analog integration using different time constants.

## Slide 5
▪Orders of magnitude from Ilya’s talk
By taking into account the non-replacement of the sources (Decreased activity)
Cells A to D: 60 to 90 nA.
Cells E1&E2:  4 nA.
Cells E3&E4:  0.06 nA.
but no Cs in them
Improvements of scintillators/Dividers are possible.
Comment: for the E cells, the constraint of 10 ms with respect to adjacent Tiles
is no longer relevant  digital sums are possible,
and/or use of a 20 ms time in the analog mode
in order to reach these low values.
Beams
Collisions
E1
E2
E3
E4

## Slide 6
 TileCal data for the Luminosity measurements
▪ Orders of magnitude from Ilya’s talk
 Maximum value at the highest Luminosity:
- A13 cell: 8 µA.
- E cells: 100 µA  no saturation for Argonne and Clermont-Ferrand.
Minimum value for the vdM scan at very low Luminosity:
up to 0.02 nA to cover almost all the A and E cells and part of B cells,
if not 0.05 nA.
with which accuracy?
1% is impossible,
5 % ? 3% ?

## Slide 7
 Chicago integrator
▪ Principle (Present ATLAS scheme):
in 2 steps
Integration time
from feedback capacitance/resistor
Amplification
because 1% signal used
1%
Combination of switches to control
Timing and Gain.
DC injection for calibration.
T = 10 + RC in ms, with C=0.1 10-9
From Ilya (Present 3-in-1): R from 2.7 to 100 M
 10. 27 ms to 20 ms. (Ilya said 10.3 to 20).

## Slide 9
▪ Performances
Present ATLAS scheme (from TIPP 2011 Chicago)
- Range 0.01 nA-1.4 µA.
- Non linearity < 1%.
- 12 bit ADC.
- Signal noise 0.003 nA
Not credible
(never demonstrated)
The PMT dark current is  2 nA at 800 V, well above this noise !

## Slide 10
- 6 gains … not given in current
(as for the present ATLAS)
0: 0-750
1: 0-200
2: 0-300
3: 0-160
4: 0-110
5: 0-80
0-2.5 V
0-5.0 V
DAC
counts
- Does not correspond to python soft
sent by Kelby very recently
maxval = [3200,180,300,120,160,90]
minval = [64,6,4,4,4,9]
Upgrade scheme

## Slide 11
Hi Francois, yes, these points are averaged over about 10 measurements
each having the int.time of 20ms.
So effectively the int.time is about 200ms on this plot.
Best,
Ilya
On Fri, Apr 8, 2016 at 1:20 PM, François Vazeille <vazeille@clermont.in2p3.fr> wrote:
Hello Ilya, looking at your very interesting and useful talk on the Cs system,
I have a question about the VdM scan through the cell A13 (slide 6) where you measured
a current of 0.5 nA.
What was the integration time of this measurement?
Did you add several measurements in order to increase the integration time?
With my best regards,
François
For times > 20 ms  summations
Example of a reply of Ilya to a question of FV about current of 0.5 nA
in the current ATLAS scheme.

## Slide 12
▪ Up to 8 µA: external
- Integration time of 2 ms, but will be fixed later at a constant value (10 ms ?).
- 16-bit ADC range, having 122 pA/count  That does mean it is the accuracy.
 Argonne integrator
 40% signal used

## Slide 13
▪ Above 8 µA: internal from digital sum
- This plot goes up to 6 nA, but with an error of about 4%.
- 0.2% above 100 nA but the noise level is very optimistic.
- Used in fact above 8 µA.
- It is a simulation.
- The integration time
is not indicated,
but it is 10 ms from Ilya.

## Slide 14
 Clermont-Ferrand integrator
▪ Accuracy calculations on digital sum from the present noise performance
Noise over the whole frequency spectra: 8 fC. (and not 7 fC HF noise only).
For a realistic pulse shape (triangular), it corresponds to 400 nA, or 50 nA/fC.
A “sum number” of 1 corresponds to the sum of 400 000 samples at 40 MHz.
Cesium scan | Cell value | Time (ms) | Sum number | Accuracy %
 | A to D 60 nA | 10 | 1 | 1.05
 | A to D 90 nA | 10 | 1 | 0.70
 | E1-E2 < 4 nA | 10 | 2 | > 11.2
 | E3-E4  0.06 | Impossible, no Cs |  | 
 | Limit 1% 63 nA | 10 | 1 | 1.0039
 | Limit 1% 45 nA | 10 | 2 | 0.9938
- For cesium scan, OK for A to D above 45 nA with an accuracy of 1% or better.
- E scintillators are likely thicker than Tiles
 a time > 20 ms could be chosen , but not too much.

## Slide 15
Luminosity scan | Cell value | Time (ms) | Sum number | Accuracy %
 | HL minimum > 100 nA from (B-C) | 10 | 1 | < 0.63
 | vdM A-B 0.02 nA | 10 | 12000 | 28.9
 | vdM A-B 0.05 nA | 10 | 12000 | 11.5
 | A, Limit 1% 0.58 nA | 10 | 12000 | 0.9954
 | A, 3% 0,192 nA | 10 | 12000 | 3.007
 | A, 5% 0.115 nA | 10 | 12000 | 5.02
- For Luminosity scans, larger times can be used.
- We must consider two extreme case: HL-LHC (7 1034 cm-2 s-1 )
and vdM scans (some 1030 ).
- For vdM scans, we take a time of 2 minutes imposed by Luminosity blocks.
- No problem for HL scans, up to the 100 µA value on E cell without saturation.
- vdM scan possible for A cells with an accuracy  3%.

## Slide 16
▪ Requirements on Analog measurements
 For Cesium scans with times of 10 or 20 ms
- Must overlap the Digital approach  Maximum value of 150 nA.
- Must reach low values up to some nA for E1-E2.
 For Luminosity scans at Low Luminosity:
to reach 0.05 nA in order to scan the B cells with an accuracy of 3%
over durations of 2 minutes.
Dynamics from 0.05 nA on means to 150 nA in 10 or 20 ms.

## Slide 18
 TileCal data for the Cesium source moving
▪ Recommendations from Ilya’s talk

## Slide 19
▪ Our specifications for the Analog approach
- A fourth current copy
with the ADC inside the chip.
- A DAC calibration
via current injection.
- Possibility of 2 integration times: 10 ms and 20 ms.
- Is it possible to use the same DAC ?
- Dynamics for Cesium scans
Maximum: 150 nA.
Minimum: 0.5 nA for E1-E2 cells.
- Dynamics for Luminosity scans
Maximum: no constraint  we keep 150 nA and saturation accepted above.
Minimum: means are possible over 2 minutes  to reach 0.05 nA.