## Slide 1: Current and future activities at the Medical University of Vienna
1,2Andreas Resch, 1,2Hermann Fuchs, 1Fatima Padilla, 1,2Dietmar Georg
1Division Medical Radiation Physics Department of Radiation Oncology / Medical University Vienna & AKH Wien
2Christian Doppler Laboratory for Medical Radiation Research for Radiation Oncology

## Slide 2: Benchmarking Gate/Geant4 for 16O ion beams
Objective: Determine accuracy and optimal settings
Reference data:
  Depth Dose Distributions measured at HIT
    Good agreement for all 5 tested hadron configurations
    Good agreement for EMY and EMZ
  Fragmentation (Charge changing CS) measured at GSI (670MeV/u)
    QMD model best  suggested
Gate 7.2

## Slide 3: Pencil Beam algorithm for 16O ion beams
Base data created with GATE (in water)
  Dose in 3D
    using radial symmetry
      Increased statistics  speed up
      Reduced memory
        Scoring grid: ,
        Phantom (cylinder) size: ,
    total and proton contribution scored separately
  dose average LET: in 1 dim

## Slide 4: Pencil Beam algorithm for 16O ion beams
Parametrization of GATE results
  2 fit functions evaluated
    Triple Gaussian (GGG)
    Gaussian + Gaussian + Rutherford (GGR)
  X2 fit very sensitive to
    start values
    dose threshold
    MC noise
  Weighted X2 fitting stable
    Regularization
Lateral dose distribution of a 430 MeV/u 16O ion beam in water

## Slide 5: Pencil Beam algorithm for 16O ion beams
Implementation of PB algorithm into stand alone research tool
  Generate arbitrary PB spot distributions
  Calculate dose in 3D using
    Parametrization (Triple Gaussian or Double Gaussian with Rutherford)
    GATE base data (discrete energies only)
  Create Gate macro files automatically
Validation ongoing

## Slide 6: 16O ion beams: field size effects
10x10cm2 PB field simulated with GATE (gps)
Build up of dose contributed by protons towards the center
Next steps:
  Spread out Bragg Peak simulations
    Tool to create spot intensities

## Slide 7: Proton Pencil Beam dose algorithm in presence of magnetic fields
Base data generated with GATE 7.1:
  Proton energies 40 – 250 MeV
  Homogeneous magnetic fields (0, 0.5, 1.5 and 3T)
  Direct parametrization of the water lateral profiles
    Gaussian central beam
    Independent exponential tailing functions for halo region
C++ / Root implemented
F. Padilla
0 T
3 T

## Slide 8: First stage validation: slab phantom (MC)
F. Padilla

## Slide 9: Second stage validation: patient data (MC) + experimental
  Ongoing evaluation period:
  MC simulations: realistic patient data. Implementation of the algorithm in an existing treatment planning software foreseen
  Experimental data
    Research beam line @MedAustron
    1T electromagnet - 3% homogeneity
    Experimental design using GATE
    Field map implementation?
F. Padilla

## Slide 10: Recently started
@MedAustron, dual energy flat panel
  cone beam CT in use
  Cooperation with MedPhoton in modelling the imaging ring system
  Head scatter evaluation
    M.Sc. thesis
Research room @MedAustron started operation
  First MC beam model available (thanks to Alessio Elia, MedAustron)
  Validation ongoing M.Sc. thesis
Investigation on the impact of the treatment nozzle on beam parameters
H. Fuchs

## Slide 11: GATE annotations

## Slide 12: GATE startDAQCluster changed
When using gjs (GATE job splitter), simulations are split using
  startDAQCluster
In Gate 8.0 the usual cluster approach stopped working
Simulations are no longer stopped after the defined time interval
Best way to fix?
  For RT application, split by particle number should work
    Future developments might include time
  (How) should we solve this issue?
H. Fuchs

## Slide 13: GATE physics list builder
What works?
  Some commands do not work (Gate 7.2):
    /gate/physics/processes/ProtonInelastic/setModel G4BinaryCascade
  Some do work (Gate 7.2):
    /gate/physics/addProcess RadioactiveDecay
Documentation up to date?
    repeatedly an issue in Gate mailing list
Re-establish disable/enable physics processes from macro file?
  Current workaround:
    create custom physics list in Geant4 + re-compile
    tedious & cumbersome

## Slide 14: Actors: Diligence
Some actors limited to certain geometries
  Cylindrical symmetry dose actor
    Beam direction (0 0 1) required
  GateCrossSectionProductionActor
    Beam direction (0 0 1) required
    Generalized to all (x, y, z) , but not
No warning or error if requirement violated
    For safety: GATE should (ideally) abort if requirements violated (effort)
    Documentation of the limitations

## Slide 15: Tools
Various tools available on our git repository
  .mhd, .txt and .root reader & merger (python or matlab)
  Automated cluster use:
    Split & Submit to cluster
    Merge: check for missing files, merge and re-submit missing splits
Contact:
  andreas.resch@meduniwien.ac.at
  hermann.fuchs@meduniwien.ac.at

## Slide 16: GATE workshop @ ESTRO
Joint operation with Uwe and Loic
Applied for workshop at ESTRO
New workshop format
17-18 of November
The multidisciplinary Gate/Geant4 Monte Carlo platform in medical physics: clinical applications and research
Looking for
  Support
  Presenters/Participants
H. Fuchs

## Slide 17: Acknowledgements
http://www.meduniwien.ac.at/hp/radonc/
and many others ……