## Slide 1
Centre de Recherche en Cancérologie de Toulouse
UMR 1037 INSERM
Gustavo Coelho Alves Costa
Supervisors:
Manuel Bardiès
Daniel A. B. Bonifácio
Modelling and validating simulations of the
GE Discovery 670 SPECT/CT

## Slide 2: Previous study
Context: Nuclear Medicine dosimetry
Modelling whole body planar acquisitions with XCAT
Reduce simulations time.
  ARF,
  Kill Actor,
  Reduced spectrum,
  Regionalized parameterization …
An article was submitted for publication (Physica Medica).

## Slide 3: Previous study
Assessing the “goodness” of the model?

## Slide 4: Introduction
European project MRTDosimetry
Metrology for clinical implementation of dosimetry in molecular radiotherapy
  http://mrtdosimetry-empir.eu
  Scintigraphic images standardization.

## Slide 5: Introduction
Work Package A3.1.4
  MC modelling to replicate a range of experimental SPECT images:
  Phantom and activity in air
  Model explicitly attenuated photons
  Model explicitly scattered photons
Work packages A3.1.5
  Metrics for comparing acquired images and modelled images?
  Qualitative and quantitative?
Work packages A3.1.6
  Generation of time sequence images from a given pharmacokinetics

## Slide 6: Materials and Methods
Simulations
The simulations were performed using GATE 7.2 on a Mac cluster containing 240 cores of 2.66 GHz Intel Xeon and 16GB ram
The acquisition protocol followed the “Protocol for calibration of a SPECT/CT camera for use in Quantitative Imaging” proposed in the project
  SPECT/CT with MEGP collimators
  Elliptical Jaszczak phantom
  Spherical source with 60MBq of 177Lu
  6 energy windows (see table 1 below)
  128 x 128 matrix size
  60 projections (120 total)
  30s/projection
Upper main scatter window [Width] | 241 keV [10 %]
Main energy window [Width] | 208 keV [20 %]
Lower main scatter window [Width] | 178 keV [10 %]
Upper secondary scatter window [Width] | 131 keV [10 %]
Second Energy Window [Width] | 113 keV [20 %]
Lower secondary scatter window [Width] | 98.7 keV [6 %]

## Slide 7: Materials and Methods
Acquisitions: elliptic Jaszczak test object.
  In air, source a the centre of the phantom.
  In water, source at the centre of the phantom.
  In water, source as far as possible from the centre.

## Slide 8
Modelling:
GE Discovery 670 installed in Lund (Sweden)
Crystal of 15.8 mm (5/8”)
Medium energy Collimator
Materials,
Dimensions, etc.
Materials and Methods

## Slide 9: Results: Energy spectrum
MIRD Pamphlet No. 26: Joint EANM/MIRD Guidelines for Quantitative 177Lu SPECT Applied for Dosimetry of Radiopharmaceutical Therapy .
Ljungberg et al 2017
XCAT phantom (Simind)
Jaszczak phantom

## Slide 10: Results: Air
Experimental
GATE
113 keV
208 keV

## Slide 11: Results: Water Centred
Experimental
GATE
113 keV
208 keV

## Slide 12: Results: Water Off Centre
Experimental
GATE
113 keV
208 keV

## Slide 13: Conclusion
The camera is well modelled as shown by the spectrum and the images.
However some differences need to be explained.
The images will be returned to partner Lund in order to discuss discrepancies between modelled and experimental projections.

## Slide 14: Future Work
Solve the validation issues.
Model explicitly attenuated and scattered photons
Implement GATE from the previous study
  ARF
  FFD + ARF ?
Task 3.1.6
  Human phantom acquisitions

## Slide 15
THANK YOU !