## Slide 1
SG - Superconducting Gravimetry
Monitoring of time-gravity changes using superconducting gravimeters
Séverine Rosat
Institut de Physique du Globe de Strasbourg; UMR 7516, Université de Strasbourg/EOST, CNRS

## Slide 2
Cryogenic dewar
Liquid He (4.2 K)
The levitation force is produced by the interaction between the magnetic field from the coils and the currents induced on the surface of the superconducting sphere.
Magnetic feedback  very low instrumental drift
(a few µGal/year, where 1 µGal ~ 10-8 m/s²)
Very high sensitivity at the nanogal level (~ 10-12 g)
Components of the SG sensor (GWR Instruments, Inc.)
Superconducting gravimetEr: SENSOR

## Slide 3
least-squares adjustment
Scale factor = -451 ± 3 nm/s²/V (relative error of 0.7 %)
 Using absolute values of gravity
Very small instrumental drift ~ 0.8 µGal/year
Superconducting gravimetEr: calibration and instrumental drift
Example of iOSG-24 with FG5#206

## Slide 4
1 µm/s²
100 nm/s²
10 nm/s²
1 nm/s²
0.1 nm/s²
0.01 nm/s²
~10-9g
~10-12g
Hinderer, Crossley and Warburton, 2007
SPECTRUM OF TIME-VARYING GRAVITY CHANGES

## Slide 5
International Geodynamics and Earth Tide Service (IGETS)
http://igets.u-strasbg.fr/
Superconducting gravimetEr NETWORK
~ 30 SGs worldwide

## Slide 6
GWR Instruments, Inc.
Early commercial model TT instruments
~1981 - 1994
200 L dewar
Observatory SG
2003 - 2012
42 L dewar
Compact SG
1994 - 2002
125 L dewar
iOSG
Since 2012
35 L dewar
Niobium sphere: 2.54 cm diameter, mass ~4.3g
Superconducting gravimetEr
Size, difficulty of installation and price
iGrav
Since 2012
16 L dewar
Field SG

## Slide 7
Today’s Superconducting gravimetErs
GWR Instruments, Inc.
iOSG
Observatory SG
Super stable: Drift rate < 0.5 µGal/month ; Scale factor constant to < 0.01 % over decades
Super precise: 0.1 nanoGal (10-3 nm/s²) resolution in frequency domain; < 0.3 nm/s² resolution in the time domain for 2 minute averaging
Super low noise: < 1 (nm/s²)²/Hz in seismic band (1 to 8 mHz)
iGrav
Field SG
Super stable: Drift rate < 0.5 µGal/month ; Scale factor constant to better than 1 part in 104 for years
Super precise: 1 nanoGal (10-2 nm/s²) in frequency domain; 0.05 µGal in the time domain for 1 minute averaging
Super low noise: 0.3 µGal/(Hz)½ ~ 1 (nm/s²)²/Hz in seismic band
Smaller and cheaper (~ 250 k€)
GWR Instruments, Inc.

## Slide 8
where ω0 is the natural frequency of the oscillator, Q its quality factor and m is the mass of the oscillating sphere (Warburton et al., 2010); kB is the Boltzmann constant and T the temperature.
Increase mass of levitated sphere  decrease noise due to Brownian motion
(as long as field at surface of sphere remains << critical value)
From magnetic gradient (between upper and lower coils)
Damping constant proportional to velocity
SG as a damped harmonic oscillator:
 Brownian noise given by:
x: relative displacement of the sphere wrt its equilibrium position
Superconducting gravimetEr
k
b
m

## Slide 9
Superconducting gravimetEr
Mass of the levitating sphere
~17.7g for iOSG (ex. at LSBB and J9)
Rustrel, LSBB
Since July 2015
iOSG-24
MIGA
Since Feb. 2016
iOSG-23
RESIF
BFO
Dual-sphere
Since Oct. 2009
OSG056
2 iOSGs
+ prototype OSG-56 (BFO)
Strasbourg, J9

## Slide 10
Rosat et al. 2016
Heavier sphere
SEISMIC NOISE LEVELS: iOSGs VS. SGs (5 QUIETEST DAYS)
(Modified from Rosat & Hinderer 2011)
iOSG-24 (and OSG-56_L)  lowest noise level in the seismic band
New Low Noise Model of Peterson (1993) = seismological noise model

## Slide 11
1 (nm/s²)²/Hz
Seismic noise levels at j9: SG vs. other instruments
Rosat et al. 2018, soumis
FG5: absolute ballistic gravimeter
gPhone-54, CG5, L&R ET-11: relative spring gravimeters

## Slide 12
Rosat et al. 2018, soumis
Seismic noise levels at j9: SG vs. other instruments
gPhone-54, CG5, L&R ET-11: relative spring gravimeters

## Slide 13
Rosat et al. 2018, soumis
self-noise extracted from a 3-channel correlation analysis
~0.1 nGal
1S1: inner core Slichter mode
[Sleeman et al. 2006]

## Slide 14
1 nGal
0.5 nGal
Rosat et al. 2003, GRL
Superconducting gravimetEr: seismic modes
Ex: High-resolution frequency analysis after the 2001 Mw 8.4 Peru earthquake

## Slide 15
1 nGal
Rosat et al. 2003, GRL
Superconducting gravimetEr: seismic modes
Ex: 1st detection of 2S1 mode after the 2001 Mw 8.4 Peru earthquake

## Slide 16
Superconducting gravimetEr: seismic modes
Ex: STS-1 vs. SG after the 2001 Mw 8.4 Peru earthquake
1 nGal
1 nGal

## Slide 17
Superconducting gravimetEr: seismic modes
Ex: STS-1 vs. SG after the 2004 Mw 9.1 Sumatra-Andaman earthquake

## Slide 18
Superconducting gravimetEr: seismic modes
Ex: STS-1 vs. SG after the 2011 Mw 9.0 Tohoku earthquake
Sutherland (South-Africa)
STS-1
SG

## Slide 19
Superconducting gravimetEr: hum
Ex: iOSG-24 at LSBB

## Slide 20
conclusion
SGs:
Perform better than long-period seismometers at frequencies lower than 1 mHz
Have a linear response to local atmospheric pressure changes  better noise reduction than STS-1
Resolution close to 0.1 nGal at seismic frequencies
Application to gravity-based earthquake early warning:
Successful detection after 2011 Tohoku EQ (Montagner et al. 2015) but some difficulties for future early warning system:
Poor geographical coverage of SGs
Signal to be detected is small and transient  needs to be separated from other possible disturbances

## Slide 21
Thank you
Severine.Rosat@unistra.fr

## Slide 22
Superconducting gravimetErs
Strasbourg, J9
iOSG23 and iGrav30 parallely recording

## Slide 23
Hydrological effect on iOSG-24 using MERRA2 hydrological model
Local hydrological effect anti-correlated with SG residuals
SG residuals after removing the total hydrological effect
SG residuals after removing tides, polar motion and atmospheric pressure effects
Superconducting gravimetEr

## Slide 24
Statistical Noise levels of some worldwide SGs
New Low Noise Model of Peterson (1993) = seismological noise model
Rosat et al. 2016
Superconducting gravimetEr: seismic noise level