## Slide 1
Development of a
Gaseous Proton-Recoil Detector
for neutron flux measurements between
0.2 and 2 MeV neutron energy
Centre d’Etudes Nucléaires de Bordeaux Gradignan,
CNRS/IN2P3, Université de Bordeaux, FRANCE
P. Marini, L. Mathieu, M. Aïche, S. Czajkowski, B. Jurado, I. Tsekhanovich

## Slide 2
I- Context and motivations
Neutron flux measurement
An independent measurement
Useful to normalize cross section measurements, for neutron beam line characterization, for dosimetry investigations…
Usually, standard reactions are used : 235U(n,f), 238U(n,f), 237Np(n,f)
These standards are known with an accuracy of few % (from 0.5% to 10%)
1H(n,p) reaction is a primary standard known with an accuracy around 0.2% in the MeV region
Especially, its application and detection are completely different from the other standards
(e.g. recoil proton detection VS fission fragments detection)
A real improvement in term of independence
Evaluators require independent measurements

## Slide 3
I- Context and motivations
Principle
Experimental setup :
Ep
Ep
Ep
En
neutrons
recoil protons
Si
H-rich foil
proton peak (Nproton)
Usual background :
The Si detector may see other signals which have to be subtracted
Need for a background measurement (masking H-rich foil)
To obtain a precise neutron flux (DNproton~1%), the background has to be small enough
Ep ~ En
Well defined efficiency
Si spectrum after usual background subtraction
Ta screen

## Slide 4
II- Issue at low energy
A tremendous background
At low energy, the background is hundred times higher than the proton peak
Non-ideal background subtraction
Ep(keV)
Counts/20keV
En=1.1MeV
En=0.9MeV
En=0.7MeV
En=0.5MeV
Remaining background
Impossible to determine Nproton better than 1% below 1 MeV
P. Marini et al., Nucl. Instr. and Meth. A, 841 (2017)

## Slide 5
II- Issue at low energy
Background source
An experiment was carried out @AIFIRA to investigate this background :
e- flux for a 1 MeV gamma-ray source
(MCNP simulations)
Si
(50 µm)
p
e-
e-
electrons
less ‘MeV/µm’ but :
. whole Si depth used
. possible tangent tracks
. more multiple events
Very high signal at few 100 keV (quickly decreasing
with E)
- present as soon as the beam is ON
- present even when the Si detector is masked
Emitted by gamma-ray interaction

## Slide 6
II- Issue at low energy
Background source
An experiment was carried out @AIFIRA to investigate this background :
Si
(50 µm)
p
e-
e-
Very high signal at few 100 keV (quickly decreasing
with E)
- present as soon as the beam is ON
- present even when the Si detector is masked
These gamma-rays are mainly produced by (n,g) reaction on the neutron source material
They cannot be avoided
Emitted by gamma-ray interaction
electrons
less ‘MeV/µm’ but :
. whole Si depth used
. possible tangent tracks
. more multiple events

## Slide 7
III- A new detector, phase I
Low sensitivity to electrons: thickness adapted to proton range
Gaseous detector:
gas pressure adapted to proton range
- much more complex
Main features
Thinner Si detectors:
need to be changed for each energy
sensitive to irradiation
more difficult to calibrate
Background events rejection
segmented detector to perform track analysis
Well defined efficiency
a collimator is placed inside the detector => two chambers
100% efficiency required for good tracks (“no missed event”)

## Slide 8
Designed with track calculations
(nbr e- generated, different En, different Pgas)
Constraint :
as low material as possible to avoid interference with cross section measurement (e.g fission)
The Gaseous Proton Recoil Detector (GPRD)
III- A new detector, phase I

## Slide 9
H-film
64 pads
smaller pads in DE region (lower e- generation)
Segmented detector (CEA/Irfu)
The Gaseous Proton Recoil Detector (GPRD)
collimators
rotating sample disk
Picture of the GPRD prototype
(test experiment)
collimators
divided in two chambers: DE-E
E region
DE region
well defined efficiency (collimated sample)
III- A new detector, phase I

## Slide 10
Energy calibration
III- A new detector, phase I
- signal amplitude:
Signal amplitude
f(exp.conditions,xyz)
f(gap, gas)
Nevertheless, due to ecollection, it is very hard to determine the proton energy via the signal amplitude
Test on constant energy deposition (very long alpha track) showed a very constant gap
hgain = f(gas)
same for all pads!

## Slide 11
Energy calibration
- track length:
obtained thanks to the segmentation
Issue: no information on the z-axis
III- A new detector, phase I
very poor energy resolution
proton energy
Issue: smooth end-of-track
Chart: Range (cm)
energy deposition reconstruction
(thanks to constant hgain, despite ecollection)
can we see this ?

## Slide 12
Efficiency
III- A new detector, phase I
H-film
collimators
E region
DE region
Calculation
(simulation)
100 % ?
not yet proven
- detector dead time:
a high counting rate could lead to missed events
. the recoil proton counting rate is usually quite low
. OK if there is not a lot of background events

## Slide 13
Test experiment on the AIFIRA facility in 2016
Validations
irradiation with neutrons (En down to 300 keV)
- same conditions than previous experiments (distance, reaction)
- track measurement
III- A new detector, phase I
track for En = 1 MeV
« straight » tracks using the whole length of the GPRD
pads amplitudes

## Slide 14
Test experiment on the AIFIRA facility in 2016
Validations
- very low sensitivity to e- or g
- track measurement
test without H-film gives no signal
(each pad is not sensitive enough to see signal from background electrons)
only few cosmic rays seen
irradiation with neutrons (En down to 300 keV)
- same conditions than previous experiments (distance, reaction)
III- A new detector, phase I

## Slide 15
Validations
- discrimination between direct and scattered neutrons
- very low sensitivity to e- or g
- track measurement
shorter “DE-E” tracks rejected
(protons from neutrons with lower energy / higher angle)
irradiation with neutrons (En down to 300 keV)
- same conditions than previous experiments (distance, reaction)
III- A new detector, phase I
Test experiment on the AIFIRA facility in 2016

## Slide 16
Validations
- discrimination between direct and scattered neutrons
- very low sensitivity to e- or g
- low detection energy limit: 300 keV (at least)
- track measurement
no test carried out below 300 keV
track for En = 300 keV
irradiation with neutrons (En down to 300 keV)
- same conditions than previous experiments (distance, reaction)
III- A new detector, phase I
Test experiment on the AIFIRA facility in 2016

## Slide 17
Validations
- discrimination between direct and scattered neutrons
- very low sensitivity to e- or g
- low detection energy limit: 300 keV (at least)
- general functioning (electronics, acquisition, gas regulation, rotating disk…)
- eintr = 100% not proven
need a quantitative and accurate experiment
- track measurement
irradiation with neutrons (En down to 300 keV)
- same conditions than previous experiments (distance, reaction)
III- A new detector, phase I
Test experiment on the AIFIRA facility in 2016

## Slide 18
Issues
- Static sparks
CF4 gas at few 10 mbars => Paschen regime
=> drop of the breakdown voltage
removal of some conductive pieces
hence the adhesive !
- Non homogeneous electric field
Weak signal on the side pads
+ perturbation due to the collimator
future addition of a field cage to constraint the electric field
III- A new detector, phase I

## Slide 19
Chart: Time (min)
| 0 | 0.34401709401709413 | 0.58974358974358998 | 0.98290598290598297 | 1.2286324786324778 | 1.4743589743589753 | 1.7200854700854704 | 1.9658119658119666 | 2.4572649572649592 | 2.9487179487179533 | 3.4401709401709422 | 3.9316239316239296 | 4.4230769230769225 | 4.9145299145299095 | 5.4059829059829063 | 5.8974358974358871 | 6.3888888888888875 | 7.3717948717948723 | 8.3547008547008641 | 9.8290598290598297 | 10.811965811965814 | 11.5
colimateur d'entrée | 2600 | 2300 | 2100 | 1800 | 1600 | 1300 | 1200 | 1100 | 1000 | 800 | 600 | 500 | 350 | 300 | 250 | 200 | 100 | 100 | 100 | 100 | 150 | 200
Chart: Time (min)
| 0 | 0.34401709401709407 | 0.58974358974358987 | 0.98290598290598297 | 1.2286324786324778 | 1.474358974358976 | 1.7200854700854702 | 1.9658119658119673 | 2.4572649572649592 | 2.948717948717956 | 3.440170940170943 | 3.9316239316239283 | 4.4230769230769225 | 4.9145299145299095 | 5.4059829059829063 | 5.8974358974358836 | 6.3888888888888875 | 7.3717948717948723 | 8.3547008547008694 | 9.8290598290598297 | 10.811965811965814 | 11.5
colimateur d'entrée | 2600 | 2300 | 2100 | 1800 | 1600 | 1300 | 1200 | 1100 | 1000 | 800 | 600 | 500 | 350 | 300 | 250 | 200 | 100 | 100 | 100 | 100 | 150 | 200
colimateur d'entrée relié cathode et masse | 2200 | 2000 | 1800 | 1500 | 1400 | 1300 | 1200 | 1200 | 1100 | 1000 | 900 | 800 | 700 | 700 | 650
Issues
- Signal loss after few minutes of irradiation
Charge accumulation on insulators
(from e- coming « from » the beam) => electric field distortion
future addition of a field cage to constraint the electric field
grounding some mechanical elements reduce the issue
III- A new detector, phase I

## Slide 20
Electric field simulations
With the OPERA code (ex-Tosca)
IV- A new detector, phase II
sample disk
collimators
micromegas detector
cathode

## Slide 21
Electric field simulations
With the OPERA code (ex-Tosca)
IV- A new detector, phase II
top view of the GPRD
Equipotential lines :
micromegas detector
sample disk
macor structure
motor axis
cathode
aluminum structure
ground ring
electric field distortion !
collimators
(ground)
(-500V)
(-2000V)
(ground)
(ground)

## Slide 22
Equipotential lines :
field cage
IV- A new detector, phase II
Electric field simulations
With the OPERA code (ex-Tosca)
much better electric field

## Slide 23
Electric field simulations
With the OPERA code (ex-Tosca)
IV- A new detector, phase II
top view of the GPRD
A static electric charge can be added
sample disk
(electric charge -2nC)
strong electric field distortion !

## Slide 24
Electric field simulations
With the OPERA code (ex-Tosca)
IV- A new detector, phase II
top view of the GPRD
A static electric charge can be added
sample disk
(electric charge -2nC)
weak electric field distortion

## Slide 25
A different gas
With very good insulating properties: 70% N2 – 30% CO2
IV- A new detector, phase II
- used in accelerator tank (as is SF6)
- very cheap
Much slower than CF4:
no background events
=> a fast gas is not mandatory anymore
Chart: E/P (V/cm/torr)
| 0.05 | 0.1 | 0.15 | 0.2 | 0.3 | 0.4 | 0.5 | 0.75 | 1 | 1.5 | 2
P10 | 2.5 | 4.8 | 5.5 | 5.5 | 5 | 4.4 | 4 | 3.2 | 2.9 | 2.6 | 2.6
CF4 | 0.275 | 0.93 | 2.61 | 3.85 | 5.95 | 7.2 | 8.05 | 9.1 | 10.5 | 11.6 | 12 | 12.6 | 13 | 13.2 | 13.1 | 12.5 | 11.3 | 10.7 | 10.2 | 9.6 | 9.5 | 9.6 | 9.8 | 10 | 10.4 | 11.3 | 11.9 | 12.8 | 13.9 | 15.9 | 18.1 | 20.3
Aligal12 | 0.7 | 1.6 | 3 | 3.6 | 4 | 5 | 6 | 8 | 12 | 20 | 26
P= 500 mbar
DV= 1500V
(En = 1,5 MeV)
P= 40 mbar
DV= 1250V
(En = 0,2 MeV)
P= 60 mbar
DV= 1500V
(En = 0,5 MeV)
P= 200 mbar
DV= 1500V
(En = 1,0 MeV)

## Slide 26
Test experiments on the AIFIRA facility in 2018
Validations
irradiation with neutrons (En down to 200 keV)
- same conditions than usual experiments (distance, reaction)
- good static electric field behavior
no more electron leakage on the side pads (except the last raw)
IV- A new detector, phase II

## Slide 27
Test experiments on the AIFIRA facility in 2018
Validations
irradiation with neutrons (En down to 200 keV)
- same conditions than usual experiments (distance, reaction)
- good static electric field behavior
- good electric field behavior under irradiation
no loss of signal amplitude with irradiation !
Chart: Irradiation time (min)
IV- A new detector, phase II

## Slide 28
Test experiments on the AIFIRA facility in 2018
Validations
- good static electric field behavior
- good electric field behavior under irradiation
- 3D tracks reconstruction
lower drift velocity => significant time difference depending
on z-axis
improvement of Eres via track length determination
IV- A new detector, phase II
z-coordinate (cm)
irradiation with neutrons (En down to 200 keV)
- same conditions than usual experiments (distance, reaction)

## Slide 29
Test experiments on the AIFIRA facility in 2018
Validations
- good static electric field behavior
- good electric field behavior under irradiation
- 3D tracks reconstruction
- low detection energy limit: 200 keV (at least)
limitation due to the
H-rich sample thickness
(0,5µm max at 200keV)
track for En = 200 keV
IV- A new detector, phase II
irradiation with neutrons (En down to 200 keV)
- same conditions than usual experiments (distance, reaction)

## Slide 30
Test experiments on the AIFIRA facility in 2018
Validations
- good static electric field behavior
- good electric field behavior under irradiation
- 3D tracks reconstruction
- low detection energy limit: 200 keV (at least)
- eintr = 100% not proven
need a quantitative and accurate experiment
IV- A new detector, phase II
irradiation with neutrons (En down to 200 keV)
- same conditions than usual experiments (distance, reaction)

## Slide 31
Next experiment
IV- A new detector, phase II
Goals
- prove eintr = 100%
- better understanding of the detector
- determine the detector rate limit
Direct proton beam experiment
A new chamber has been designed:
direct proton micro beam
(few p/s to few 10 p/s)
GPRD shift in (x;y)
to irradiate different parts
Si detector
to monitor the counting rate

## Slide 32
Conclusion
A difficult energy range for recoil proton
very high background due to n->g->e-
prevent an accurate counting of recoil protons
The Gaseous Proton Recoil Detector
built in 2016
designed for a low sensitivity to g/e- background
test experiments @AIFIRA facility :
. validation of the low sensitivity to background
. low energy limit of 200 keV (limited by the thinness of the H-film)
. good behaviour under irradiation
- a direct proton experiment planned to investigate the efficiency
. track reconstruction
Perspectives : The GPRD will be completed and used to measure the 242Pu(n,f) cross section below 1 MeV (submitted to EURATOM WP2018)

## Slide 33
Thank you for your attention