## Slide 1: The Collimated Beam Projector (CBP)
Nicholas Mondrik (Harvard)
Collaborators:
Michael Coughlin (Caltech)
Christopher Stubbs (Harvard)
John Woodward (NIST)
Paris LSST Calibration Meeting, May 23-25, 2018

## Slide 2: Outline
Purpose of the CBP
CBP design and operation
Collecting data with the CBP
Results
Troubles, lessons, and the future

## Slide 3: The flat field is not enough
Traditional broadband flat fields use a non-collimated light source to illuminate the focal plane
Can use monochromatic flat fields to estimate instrumental throughput as a function of wavelength
Primary issue:
  Ghosted light paths are different than those produced by collimated beams
  Wavelength dependent ghosting can masquerade as QE variation

## Slide 4: Enter the Collimated Beam Projector (CBP)
Tunable laser
Integrating sphere
Monitor photodiode and spectrograph
Image mask
Collimating optic
To telescope
Variable ND filter

## Slide 5
Tunable laser
CBP mounted on platform attached to dome

## Slide 6: Tunable light source
https://ekspla.com/product/nt242-series-tunable-wavelength-nanosecond-lasers/
Approx. 400nm
Ekspla NT-242:
1 kHz rep rate
1064nm pump

## Slide 7: LSST filters

## Slide 8: Key point:
Use the NIST-traceable photodiode as metrology standard, NOT light source
(X correction factor)
Challenge: measuring the transfer function of the CBP system

## Slide 9: The Data

## Slide 10: Taking data: SMARTS 0.9m, CTIO, Chile
Open telescope shutter
Open CBP shutter
Expose
Close CBP shutter
Close telescope shutter
Change wavelength

## Slide 11: Charge homogeneity
Goal: avoid divide-by-zero errors

## Slide 12
Primary spot (5x5)
Ghost spot
(due to filter)

## Slide 13: The results (so far)

## Slide 14: Current Difficulties and Places for Improvement

## Slide 15: Harvard CBP Specific

## Slide 16: Challenges for the Harvard CBP
Alignment (equatorial vs AltAz + dome rotation)
Contamination from signal into idler beam
Temperature monitoring
Change from refracting to reflecting optics?
Fix tunable laser (or find new light source)

## Slide 17: Questions and Challenges for the LSST CBP

## Slide 18: General CBP challenges
Vignetting* corrections between days (more after this)
Output of tunable light source in blue (<400nm)
Spot-to-spot comparison map
Using a CBP to flat-field Ronchi data for atmospheric transmission
How broad can light source FWHM be, and still achieve effective bandpass measurements (alternatively, a method for including observed FWHM in bandpass extraction)
Impact of OPO degeneracy point on filter scans of r and i
Where should we take the CBP next?

## Slide 19: “Vignetting” correction issue
In order to make accurate synthetic colors, we need to be able to put 2 different filters (w/ scans taken on different nights) onto same flux scale. That is, we need to know that the peak I-band throughput is 80% of the peak R-band throughput (don’t care about what the overall multiplicative offset is, though).

## Slide 20: Spot Matching
Create a map that links same spots on CBP mask between different runs
No-filter scan image taken on 2017-10-11
No-filter scan image taken on 2017-10-04

## Slide 21: Inconsistent raw throughput ratios
Plot is ratio of raw throughputs of each spot
Note that not only are there large (>20%!) offsets between nights, but also not consistent spot-to-spot
Nofilter/nofilter

## Slide 22: How does it affect throughput measurements
Back-to-back scans taken without moving spots, CBP, or Dome
Our “flagship” R-band curve
Same curve, but made by dividing an R-band scan by the 2017-10-11 no-filter scan
BAD

## Slide 23: Can we correct for it?
Same pair as before, but this time, each spot has a “vignetting”(=overall multiplicative scaling) applied. Correction factor is found by MEDIAN(spot_A_Nofil_dayX[common wavelengths]/spot_A_Nofil_dayY[common wavelengths])
In this case, the nofilter 2017-10-11 image is corrected by data from the nofilter 2017-10-04 run

## Slide 24: Application to I band
Spot-spot ratio of an I band scan from 2017-10-07 to the no filter scan on 2017-10-11
Same, but with a vignetting correction (based on a single wavelength no-filter image taken on 2017-10-07)

## Slide 25: Conclusion
Large steps in the maturity of the CBP
There are a number of systematics in the existing CBP data that have been identified, and are in the process of being quashed
Unfortunately, the set of data taken at CTIO in 2017 is probably not going to provide us with final filter curves with uncertainties <1%
There also exist a number of broader challenges with regard to CBP data, which hopefully DESC can provide some input on