Calibration from 137Cs
Luc Pourqué
Vincent Boudry*

---
Long Slab 8 ASUs
Source 137Cs, 370 kBq
on 1 ASIC with babywafer @ D ~ 12 mm
Threshold scan from 250 to 460, step 30
Acquisition time = 25200 s

---
Fit for Chip 58 (ASU 4) ch 11, High Gain, SCA0, threshold 280 :
- Erf function caracterises the trigger resp.
- Erfc function caracterises the Compton Edge ≡ 477 keV
- Exp function modelises the spectrum

---
Fit parameters for all 16 channels, SCA0, high gain, threshold 280
Good fit for the Erf, less so for Erfc (OK if good stat)

---
- We can check the DAC-ADC linearity with a mean of all channels
- We observe a potential saturation beyond thr400

---
DAC-ADC linearity relation for all 16 channels

---
- Linearity parameters for the previous fit
- The outlier is caused by an error in the pol1 fit

---
Compton Edge
- The erfc position should remain stable relative to the threshold value
- These variations can be explained by a bad fit or statistical variations due to the relatively small number of events

---
The same work can be done with the low gain

---
DAC-ADC linearity, low gain

---
DAC-ADC linearity, 16 channels, low gain

---
Distribution of the linearity parameters, 16 channels, low gain

---
Optimal ACQ
Data for threshold values 250 (blue), 340 (green) and 460 (red) ; High Gain; SCA0
Statistics at end spectrum is largely DAQ limited (not source)

---
Higher statistics for Compton Edge
Same measurement, threshold 240, acquisition time 60000 s

---
Fit for Chip 90 (ASU6) ch 4, High Gain, SCA0, threshold 240 :
- Erf function caracterises the trigger resp.
- Erfc function caracterises the Compton Edge ≡ 477 keV
- Exp function modelises the spectrum

---
Fit parameters for the 16 channels

---
Fit parameters for all 16 channels, SCA0, high gain, threshold 240
Good fit for the Erf, less so for Erfc (OK if good stat)

---
A measurement was also made on the stack prototype
- Source 137Cs, 37MBq
- D~10 cm
- Acquisition time = 60000 s
- Threshold 240

---
- Fit for the sum of all SCAs
- Degrades the quality of the fit of Erfc

---
A simulation of the energy deposited in the detector was realised with Geant4

---
Different values of resolution for the convolution
%sigma_energy = 30 over sqrt(1000)
%sigma_energy = 70 over sqrt(1000)
%sigma_energy = 50 over sqrt(1000)

---
Next...
- Extract calibration for LS (almost there)
- Extract all parameters for stack
  - Pedestals per CH and SCA.
  - Thr. = a + b×DAC
    - σ of thr. (should be cst)
  - Edge position & σ
    - Complement mip (CE ~ 4 mips)
- Understand slope discrepency data/MC