## Slide 1
Développement d’un dispositif
de récupération et contrôle
du Xénon en anesthésie
Jose Busto: Pr.; Aix-Marseille Université/CPPM
Greg Hallewell: IRHC (HDR); CPPM
Alexandre Martinez: Pr.; Institut des Sciences Moléculaires de Marseille/ CPPM
Avec collaboration de Prs. John Dingley et David Williams;
professeurs en anesthésie: Univ. Wales Medical School: Morriston Hospital,
Swansea, Pays de Galles, GB
Experts en anesthésie xénon et instrumentation associé.
Présentations des activités interdisciplinaires
au CPPM, le 9 janvier 2020

## Slide 2
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020
Anesthésie par inhalation
(gaz ou vapeur ajoutée a l’oxygène)
X
X

## Slide 3
Notre but: augmenter l’efficacité et sélectivité de récupération de Xe suite aux anesthésies:
prix réduite  no. interventions avec les avantages neuroprotecteurs Xe etc. augmenté
Strategie: capture de Xe suite aux interventions sur cartouches régénérables
de matériel nano-poreux proche a la température ambiante…  brevet?
Développement: - avec fournisseur de gaz et/ou cartouches d’adsorbeur – business model
pour recyclage de Xe gaz + cartouches d’adsorbant
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020
Etat de l’art
(circulation et (non)-recuperation)

## Slide 4
“XePOHCAS” Xenon for Neuroprotection During Post-Cardiac Arrest Syndrome in Comatose Survivors of an Out of Hospital Cardiac Arrest - Tabular View - ClinicalTrials.gov
(https://clinicaltrials.gov/ct2/show/NCT03176186)
Cohort: 1400 patients in multiple hospitals in the US.
Likely duration of 50%Xe/50%O2 ventilation + poss. hypothermia @ 33 ⁰C 24 hrs
Anaesthesia
Neuroprotectant
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020
En cas de succès….
Augmentation de demande Xe:
Prix de Xe <<< soins de longue durée
pour patients autrement handicapées
suite au privation d’oxygène

## Slide 5
Current Xenon – Oxygen monitoring in anaesthesia:
Current technology is “split”:
electrochemical sensor for O2,
thermal conductivity (very non-linear) for Xe: (INSOVT GKM-03 analyzer):
retro-adaptation of existing FDA-approved ultrasonic flowmeter (Drager Spirocell) add improved flow algorithm and simultaneous gas analysis: (single technology),  new instrument (more modern electronics, firmware, optimal acoustic cell)
for Xe/O2 monitoring in respired and recovered gas.
5
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 6
Driver
m-controller
Threshold
generator
Communication
Interface
Acoustic
transducers
Up/Down
Transducer excitation
(simultaneous transmission mode)
Bias voltage
D, U
Received
Sound signals
Transit
time TU
Gas flow
Amplifier/
comparator
Multi - MHz clock
Transit
time TD
Threshold
xducer bias
generator
Gas flow velocity determination depends on TU – TD difference in sound transit times
measured in two opposite directions in flowing gas:
Gas mixture determination depends on sound velocity calculated from TU,TD average
Simultaneous Ultrasonic Gas Analysis & flowmetry using
bi-directional ultrasound transmission
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020
In the range (40-50% Xe in O2)
slope of composition / SoS
curve is ~ -0.8% Xe/ ms-1
Xe conc. uncertainty = Sound velocity uncert.*local gradient dmix=dVs*m: dv of 0.1ms-1  d %Xe = 0.1%

## Slide 7
Circuit fermé pour créer et circuler les mélanges xénon-oxygène (mesures simultanées débit & conc.)
“Split” technology GKM-03
Xe-O2 analyzer
Combined Ultrasonic
flowmeter / analyzer
7
Bancs de mesure au CPPM (1)
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 8
(1) : Banc de mesure d’adsorption de xénon sur les nanomatériaux (au CPPM)
(développé pour mesures d’absorptivité des nano-materiaux pour 222Rn: J. Busto).
Abs_Xe = (Pinitiale - Pfinale)/Pinitiale* (Voltube– Volads)/mads
Bancs de mesure au CPPM (2)
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 9
Recycling and recovery of Xenon
Preferable adsorbents
Modern activated carbons & Cryptophanes are probably best candidates for Xe recovery:
High binding constant >104 at higher temperatures
Tuneable cavity
Easy adsorption and desorption of selected gas
stability in working environment
Cryptophane
Some cryptophane congeners
Polymer/Silica Supported Cryptophane
Selective Xe binding cryptophane
Xe (xenon) 216 pm
Ra (radon) 220 pm
Silica Supported

## Slide 10
O2
CO2
Ultrasonic instrumentation for simultaneous Xe-O2 gas mixture analysis
O2
CO2
Recovered
Gas
Flow meter
= Xenon (Xe)
= Carbon dioxide (CO2)
= Oxygen (O2)
Adsorbent use and regeneration plan
Cartridge A
Cartridge B
Cartridge A
=Adsorbent for CO2 and O2 and H2O trapping
Cartridge B
=Adsorbent for Xe trapping
Cartridge A
Cartridge B
SodaLime
Cartridge CO2 removal in closed circle
Ultrasonic Xe-O2 analysis to detect cartridge saturation
Cartridge B for regeneration
Ultrasonic Xe-O2 analysis
+ CO2 measurement (patient safety)
PostOp Gas
recovery line
Industrial
Clinical

## Slide 11
Work plan & planning (proposition DOC2AMU 2018: dossier non- retenu)
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 12
D. Williams, G. Hallewell, E. Chakkarapani & J. Dingley;
“Real-time measurement of Xenon in a binary gas mixture using
ultrasound time-of-flight:a feasibility study”;
Anesthesia & Analgesia: October 2019 - Volume 129 - Issue 4 - p 985-990
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 13
Financement recu ou en en cours
Équipements 21200 €
Fonctionnement: 15 K€
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 14: Financement de 15kE par R. Clédassou: achat consommables: échantillons d’adsorbeur + précurseurs chimiques pour synthèse de cryptophanes + leurs supports
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020
Adsorbants identifiées
 Phase Pré-maturation

## Slide 15
Statut du projet xénon anesthésie
(1) Nous avons reçu/allons recevoir financement pour petits équipements et consommables
pour identifier des adsorbants prometteurs (bonne efficacité/sélectivité Xe
proche a la température ambiante);
(2) Nous avons développé un algorithme pour les mélanges Xe/O2 par intermédiaire
d’un débitmètre a ultrasons (Drager Spirocell): publication en journal d’anesthésie 2019;
(3) Un succès dans l’activité (1) pourra ouvrir la possibilité d’un brevet:
soutien de programme Prématuration CNRS + SATT…
(4) Brevet pour protéger la P.I. avant d’ouvrir discussions avec fournisseur de gaz pour
business model recyclage Xe + cartouches d’adsorbants pour régénération et réutilisation;
(5) Nous avons identifie un programme de travail d’env 3 ans pour développer
l’instrumentation ultrasonique et tester les adsorbants aux débits caractéristique a la
vitesse de récupération de Xe a la fin d’un anesthésie;
On espère que ce projet aura bientôt suffisamment de soutien pour pouvoir attirer
et embaucher les doctorants et postdocs: besoin de manpower.
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 16
Reserve
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 17
Work plan & planning (proposition DOC2AMU 2018: dossier non- retenu)
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 18
Driver
m-controller
Threshold
generator
Communication
Interface
Acoustic
transducers
Up/Down
Transducer excitation
(simultaneous transmission mode)
Bias voltage
D, U
Received
Sound signals
Transit
time TU
Gas flow
Amplifier/
comparator
Multi - MHz clock
Transit
time TD
Threshold
xducer bias
generator
Gas flow velocity determination depends on TU – TD difference in sound transit times
measured in two opposite directions in flowing gas:
threshold & electronic delay usually cancel
Gas mixture determination depends sound velocity calculated from TU – TD average
threshold & electronic delay effects are important
Simultaneous Ultrasonic Gas Analysis & flowmetry using
bi-directional ultrasound transmission
Typical total acoustic path in a medical instrument is just a few cm:
The cell and transducers need to be sterilizable
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 19
19
Using Cp, Cv data for ALL expected constituents calculate concentration vs. measured sound velocity over expected temp., press. & CO2, H2O range, to make database of fit equations
In practice, work in the 100%RH regime:
 H2O molar concentration given by svp:
Xe/O2 anesthetic gas mixture can contain CO2 if soda-lime cannister ineffective.
implement Xe/O2 gas concentration measurement algorithm in T, P, CO2 space
3D (T, P, CO2) database grid intersections are ‘cans’ containing polynomial fit parameters for Xe/O2 conc vs. measured sound velocity, which are interpolated to measured (T, P and CO2) by 3-step collapse. Can also go to 4D (tessaract) etc.
In the range (40-50% Xe in O2)
slope of composition / SoS
curve is ~ -0.8% Xe/ ms-1
Xe conc. uncertainty = Sound velocity uncert.*local gradient dmix=dVs*m: dv of 0.1ms-1  d %Xe = 0.1%
Binary analysis with look-up tables of SoS. vs. composition for Xe/O2 in presence of 3rd party contamination (particularly CO2 , H2O)
19
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020
Notes: 19

## Slide 20
Calculate Cp, Cv for the individual
component gases over the expected
temperature, pressure & CO2 range
20
(2) Calculate concentration vs. measures sound velocity over expected temp., press. & CO2 range, to make database of fit equations
Xe/O2 anesthetic gas mixture can contain CO2 (and perhaps H2O) if soda-lime cannister ineffective.  implement new « 3-D » gas concentration measurement algorithm: new Xe/O2, T, P, CO2 database
3D (T, P, CO2) database grid intersections are ‘cans’ containing polynomial fit parameters for Xe/O2 conc vs. measured sound velocity, which are interpolated to measured (T, P and CO2) by 3-step collapse.
In the range (40-50% Xe in O2)
slope of composition / SoS
curve is ~ -0.8% Xe/ ms-1
Xe conc. uncertainty = Sound velocity uncert.*local gradient dmix=dVs*m: dv of 0.1ms-1  d %Xe = 0.1%
Example: Look-up tables of SoS. vs. composition for Xe/O2
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020

## Slide 21: References
[1] J. Dingley, R. King, L. Hughes et al; “Exploration of xenon as a potential cardiostable sedative: a comparison with propofol after cardiac surgery”. Anaesthesia. 2001; 56 (9):829-835.
[2] T. Goto, H. Saito, M. Shinkai, Y Nakata, E. Ichinose & S. Morita; “Xenon provides faster emergence from anesthesia than does nitrous oxide–sevoflurane or or nitrous oxide–isoflurane”. Anesthesiology. 1997; 86 (6):1273-1278
[3] D. Ma, S. Wilhelm, M. Maze & N. Franks; “Neuroprotective and neurotoxic properties of the “inert” gas, xenon”. Br J Anaesth. 2002; 89(5): 739-746.
[4] Y. Ishiguro; “Cardiovascular effects of xenon”. Int. Anesthesiol Clin. 2001:39(1):77-84.
[5] B. Lachmann, S. Armbruster, W. Schairer et al; “Safety and efficacy of xenon in routine use as an inhalational anaesthetic”. Lancet. 1990; 335 (8703):1413-1415.
[6] “XenonPOHCAS”: Clinical trial using xenonnon for neuroprotection during post-cardiac arrest syndrome in comatose survivors out-of-hospital cardiac arrest https://clinicaltrials.gov/ct2/show/NCT03176186
[7] J. Dingley & R. Mason; “A cryogenic machine for selective recovery of xenon from breathing system waste gases”. Anesthesia & Analgesia 2007; 105:1312–8
[8] N. Burov et al. Anesteziol Reanimatol. 1997 Jul-Aug;(4):71-4.
[9] D. Williams, G. Hallewell, E. Chakkarapani & J. Dingley; “Real-time measurement of Xenon in a binary gas mixture using ultrasound time-of-flight: a feasibility study”; Proc. Euroanaesthesia2017, Geneva, June 3-5, 2017; http://euroanaesthesia2017.esahq.org/. Anesthesia & Analgesia: October 2019 - Volume 129 - Issue 4 - p 985-990 doi: 10.1213/ANE.0000000000003806
[10] J. Busto et al; “Separation of rare gases and chiral molecules by selective binding in porous organic cages”.
Nature Materials, 13, 954-960 (2014)
[11] R. Noel Thèse doctorale – Université d’Aix Marseille, 30/11/2015. « Etude et développement de substrats microporeux pour l'adsorption du radon et son application en physique du neutrino »
[12] R. Noel et al. Proc. 5th International Workshop in Low Radioactivity Techniques. 18–20/3/2015 Seattle, WA.
[13] D. Zhang, A. Martinez & J.-P; Dutasta; “Emergence of hemicryptophanes: from synthesis to applications for recognition, molecular machines and supramolecular catalysis”; Chem. Rev. 2017; 117: 4900–4942.
[14] T. Brotin and J.-P. Dutasta “Cryptophanes and Their Complexes—Present and Future”
Chem. Rev. 2009, 109 : 88-130.
[15] C. Rossi et al. “Applications and perspectives of gas analysis based on sonar Instrumentation’’
Thermal Science and Engineering Progress 2019, 9: 169-176

## Slide 22
D. Williams, G. Hallewell, E. Chakkarapani & J. Dingley;
“Real-time measurement of Xenon in a binary gas mixture using ultrasound time-of-flight:
a feasibility study”;
Proc. Euroanaesthesia2017, Geneva, June 3-5, 2017; http://euroanaesthesia2017.esahq.org/. Anesthesia & Analgesia: October 2019 - Volume 129 - Issue 4 - p 985-990
Présentations des activités interdisciplinaires au CPPM, le 9 janvier 2020