## Slide 1
PICMIC: The First Results
FCPPN/L2024 10th June 2024, Bordeau, FRANCE
H. Abreu(1), E. Bechetoille(1),G. Bertolone(2), G. Claus(2), C. Colledani(2), C. Combaret(1), G. Doziere(2), C. Hu-Guo(2), I. Laktineh(1), H. Mathez(1), H. Pham(2), M. Specht(2), I. Valin(2), L. Zhang(3), Y.Zhao(2,5). , D. Breton(4), J. Maalmi(4)
(1) IP2I, Institut de Physique des 2 infinis, CNRS/IN2P3
(2) IPHC, Institut Hubert Curien, CNRS/IN2P3
(3) SDU : Shandong University, China
(4) IJCLab - Orsay : Laboratoire de Physique des 2 Infinis CNRS/IN2P3.
( (5) Nankai University, China )

## Slide 2
Overview
Concept/Introduction
Description of the Test Bench (HW+FW)
Pedestal thresholds calibration “S-Curves”
Signal Injection ”Hat-Curves”
Direct injection on the chip surface (Probes)
Proof-of-concept Demonstrator
2

## Slide 3
PICosecond subMICron detector:
https://t.ly/Iiw7G
Imad’s presentation last year
3

## Slide 4
SAMPIC
FirmWare
SAMPIC
HW
ti-tj
=>
3ps Time precision + 1mm position estimation
5µm position, in 400ns frame resolution
(Row,Col)(35,103)(38,63)(37,12)
(tracker-like)
(electron absorber)
E. Bechetoille, PICMIC, TWEPP, 1–6 oct. 2023, Geremeas, Sardinia, Italy
Proof-of-concept Demonstrator
Notes: Today, the MCP, is not cabable of generating 3ps resolution signal, but SAMPIC measure at the precision.
4

## Slide 5
PICosencod subMICron detector:
Imad’s presentation last year
5
L/D = 40-200
L is the plate thickness and D the tube diameter
D=3-25 µm
Gain/MCP ≂ 10000
it allows the detection of charged particles but also neutral ones such as γ and neutrons when equipped with appropriate converters (photocathode)
8-point PCB for the time measurement.
PCB to be in contact with the grid
SAMPIC: WTDC ASIC allowing to reach ~3 ps time resolution
MCP
TIME+position
- position determination using triangulation.
DT
y[mm]
x[mm]

## Slide 6
PICosencod subMICron detector:
6
POSITION+time
To reach few microns resolution with a limited number of electronic channels, a genuine concept was developed, successfully tested and patented.
- Connect the pixels in woven strips
- Two neighboring pixels are connected to two different strips
- Each strip is connected to one electronic channel
- Share the charge among a few ones
- Crossing the fired strip to determine the position
Reduction of electronic channels, power consumption.

## Slide 7
PICosencod subMICron detector:
7
6 metal layer technology:
3xdirectionsx852 hexagon(5μm)-strip-lines, in TOP metal interconnector Metal-5-4-3
[53,128] matrix of Analog+Digital comparator readout in M2 & M1, each line gets a comparator.
POSITION+time
Analog readout
Digital readout
Very front-end:
- Current mirror + current comparator
- Current reference: global 8-bit tuning DAC + 3-bit local
- protection diodes
i2c Slow control :
- global registers + each readout-cell locally
- Column readout by priority encoder Parallel readout :
- 13-bit data (touched pixel address) + Marker
- 40MHz // readout speed
- 16 sample by frame : 16* 25ns => 400ns frame
https://doi.org/10.1088/1748-0221/19/05/C05015
TWEPP2023 paper

## Slide 8
Test Bench
Test bench HW developed by IPHC - New design for sensor board & Arduino shield - Most of the test bench HW reused from a previous project - I2C: Arduino DUE board - DAQ: NI PXI6562 board 16 I/O up to 200 MHz
PICMIC DAQ Sequence Schematic:
8

## Slide 9
Freeze a set of DAC(Global) values driven by : - small dispersion between s-Curves - bigger number of sCurves in a given space of phase vs VRefN(quite sensitive)
Then, additional scan in local values are performed: - iadj<3bits>, sw<2bits>, EN_CM, EN_CC.
To determine value @50% inflection point.
To select the closest sCurves@50% to VRefN
Pedestal threshold calibrations
PICMIC Very Front-End synoptic(Analog part)
(1/2)
The first tries were performed choosing a few channels (among the 2556 available):
[Protocol]
2D scan [iadj vs VRefN] results following steps1-3 of the protocol
Matched sCurves(VRefN) for sets of values in iadj
9

## Slide 10
Freeze a set of DAC(Global) values driven by : - small dispersion between s-Curves - bigger number of sCurves in a given space of phase vs VRefN(quite sensitive) -- VRefP=50, VBP=90, VNB=45, VBN_adj=128
Then, additional scan in local values are performed: - iadj<3bits>, sw<2bits>, EN_CM, EN_CC.
To determine value @50% inflection point.
To select the closest sCurves@50% to VRefN. The software’s interface of the test bench modified to allow up to 5-Dimensional scans.
Pedestal threshold calibrations
(2/2)
Full calibration of 2556 channels. Global fixed, varying local: 32 configurations per channel: ~81800 sCurves to analyse.
[Protocol]
*Excluded 38 Yellow channels (Metal 3), noisy due to coupling between memory flush signal in Metal 2
full matched sCurves(VRefN)
10

## Slide 11
Injections Studies
PICMIC allows external signal injection trough the (PAD): - R424[R68,C28],R425[R103,C35] and R426[R6,C15]. - * it also allows to excite whole columns in the channel matrix (no shown in above layouts).
VPULSE
1pF
‘1’
iComp
iRefN
ith+iin
No edges
Vout
0
1
Dout
0
1
iComp
iRefN
ith+iin
fall edge
Vout
0
1
Dout
0
1
Dout
iComp
iRefN
ith+iin
No edges
Vout
0
1
Dout
0
1
Voltage injection through external 1pF, then scan of the threshold of Current Comparator.
Below the current threshold, the falling edge sensitive does not trig.
The signal is recorded on falling edge, so above a certain value of iRefN added to threshold current, the comparator does not trig anymore-→ chapeau!
(1/2)
VRefN [udac]
Efficiency [%]
11

## Slide 12
Injections Studies (Probes)
Having the full pedestal thresholds under control (PICMIC calibrated). We proceeded to accomplish an important mainstone. Direct inject on the surface.
(2/2)
Probe scan from bottom - top
Clustering and centroid estimation
left-probe
right-probe
both-probes
Two probes injection
12

## Slide 13
Proof-of-concept demonstrator
First results: Position detector validated.
13
- SAMPIC WTDC integrated with the PICMIC0. - New Calibration (pedestal threshold) in-situ, to avoid potential noise from the setup.
- concept validated!. - Work in progress to reduce channels dispersion. Fine tune at the calibration level. - Developing coincidence protocols (SAMPIC+PICMIC0) - Looking forward to move and test with beams.
Using a mezzanine to interconnect time+position sensors (IJClab-Orsay)

## Slide 14
Future plans:
- Hit rate improvement from 2.5Mhz to 10Mhz or 100MHz (expected)
- Sensor size to increase - Hexagon pitch from 5um to 1.25um - migrated in a smaller tech node like 65nm or less (with more metal layers)
- tile 4 die - narrow electron shower to keep PICMIC0 target resolution → Going beyond of MCP - Idrogen (high rate readout) : Liroc + picoTDC

## Slide 15
Summary
PICMIC concept validated:
    Principle of very precise time measurement already validated.
    Principle of very precise positions validated.
A first demonstrator using SAMPIC for time measurement and combining both time and position measurements based on SAMPIC DAQ is being used to combine both very precise time and position measurements.
A new detector NCP is being developed to reach unprecedented resolutions in both time and position.
A board integrating a low-jitter timing preamplifier and discriminator (LIROC-OMEGA) as well as a precise TDC (picoTDC-CERN).
15

## Slide 16
Summary
PICMIC concept validated:
    Principle of very precise time measurement already validated.
    Principle of very precise positions validated.
A first demonstrator using SAMPIC for time measurement and combining both time and position measurements based on SAMPIC DAQ is being used to combine both very precise time and position measurements.
A new detector NCP is being developed to reach unprecedented resolutions in both time and position.
A board integrating a low-jitter timing preamplifier and discriminator (LIROC-OMEGA) as wellas a precise TDC (picoTDC-CERN).
16

## Slide 17
BACKUP
17

## Slide 18
Single pixel s-curves
18

## Slide 19
s-curves [update]
Picmic has 5-DAC [VRefN,VRefP,VBN,VBN_adj and VBP]. Global variables.
2 local variables → Iadj<4bits> + SW<2bits>
We started to performing studies to understand relationship between DACs[Globals] +and Local one.
Two-Dimensional scans :
    VRefN and VrefP Scan : VRefN and VBN_adj Scan
      In both cases, nice linear trend
19

## Slide 20
s-curves [update]
Now, Three-dimensional scan. Everything fix except VRefN, VRefP and VBN_adj :
It allows to see previous results and to define the following for the protocol:
  Vg= VrefN*400nA-VrefP*80nA-VBN*80nA*10*(1.5 - 0.75*SW<0> - 0.5*SW<1>) (step 1)
  Vl= iadj<2>*0.2+iadj<1>*0.6+iadj<0>*1 (step2)
  – step1 → compute the mean value at 50% inflection point Vg – step2 → from step1 estimate values in DACs, then fine tune using locals .
    As already shown in the last meeting
20

## Slide 21
Effects # of ways in sCurve Scan
In order to reduce the iteration time (data production), we studied the effect as is affected the sCurve for a way according the number of ways added together
21

## Slide 22
Effects # of ways in sCurve Scan
In order to reduce the iteration time (data production), we studied the effect as is affected the sCurve for a way according the number of ways added together
In agreement within 1 [udac]
22

## Slide 23
Injection [update]
We roughly know how to melt a few curves. So, lets try again signal injection for three pixels studied previously (R424, R425 and R426) his effect over there.
S-curves in three ways after calibration.
23
23

## Slide 24
Temperature Tests
Temperature lows and stables during DAQ scans (ON during several days).
Effects within 1[udac] in VBN_adj(low current step) value scans.
24

## Slide 26
Confidential

## Slide 27
DT
Confidential