## Slide 1: Focal plane
Pierre Antilogus, Claire Juramy

## Slide 2: Camera assembly
Achieved 10 µm planeity across focal plane
4k x 4k CCDs
  21 ‘science rafts’ (3 x 3)
  4 ‘corner rafts’ (2 + 1)
Two vendors: 13 T-e2v rafts, rest with ITL

## Slide 3: Full focal plane
First light September 2020, several runs

## Slide 4: Sensor and electronics effects
Intrinsic sensor effects:
  Defects: bright, dark
  Quantum Efficiency, fringes
  Static effects: tree rings, edges, midline
  Dynamic effect: brighter-fatter
Sensor management effects:
  Clock-induced: tearing
  Charge Transfer Efficiency: clocking, traps, PTC dip, full well
  Clock injections: noise, bias
  Bias drift
  CCD output stage: gain stability, noise
  Readout speed: noise
Readout electronics:
  Gain stability
  Linearity
  Crosstalk
  Memory
  Bias shift
Acquisition :
  ADU favoritism
  DAQ scrambling

## Slide 5: Quantum efficiency
Scan of all sensors, rafts
Not absolutely calibrated
CCOB

## Slide 6: Parallel transfer efficiency
Parallel clocks in e2v: 4 phases, crossing
Delay rise / fall, propagation over sensor area
No overlap in parallel clocks
1 µs overlaps

## Slide 7: Dynamic range and PTC dip
Parallel (mainly) transfer efficiency at high flux
Voltage range + timing
Measured on flat fields, narrow beam CCOB?

## Slide 8: Dark corner
Appeared in early focal plane frames with light
Propagation of parallel clocks
Simulate glitch time after flat exposure (0 / 200 ns / 500 ns / 1 µs)
Bug in sequencer management

## Slide 9
Skyline pattern: observed at high deltaP (> 9.5V), builds up from middle of sensor
Thin tearing: excess of charges comes before deficit in readout: not an issue with parallel transfer
Divisadero
9
Original tearing patterns

## Slide 10
Bipolar mode: variety of structures depending on working conditions and sensor
Without inversion (< -7V), same conditions as unipolar mode
Testing on full focal plane

## Slide 11
Bias 2D structure: all sensors
Specific to some e2v sensors: time-dependent structure
Bias drift

## Slide 12: Clock injection in ITL rafts
Individual clock injections: S1 > S2 > S3
Affects bias level: combination of clock injection
Injects noise

## Slide 13: Bias shift
Seen from bright defects
Very long relaxing time
Threshold ~ 190 kADU (raw)
Adjustement to ASPIC gain + RC + integration time

## Slide 14: Electronic gain variations
Fine reconstruction of monitoring values
Current consumption increased in ASPIC when not running
Need to clamp (restore ASPIC first stage) regularly

## Slide 16: Scan de la lecture

## Slide 17: Horloges et scan
Raw OS
Integrated
ASPIC signals
ADC trigger
CCD serial clocks