## Slide 1: CNAO accelerators and experimental room
Pullia M.

## Slide 2: Comparison of the depth dose profiles
Protons/Carbon
Electrons
Photons
Tumour
Beam

## Slide 3: Spread Out Bragg Peak
To treat thick tumours
a "spread out Bragg peak"
is created by overlapping
many pristine Bragg peak
at different depths,
corresponding to different
energies

## Slide 4: Treatment delivery
(courtesy of Siemens Medical)

## Slide 5: CNAO dual (multi) particle accelerator system
2 ECR sources
1 RFQ + LINAC
1 synchrotron
Slow extraction
Protons
60-250 MeV
< 1010 p/spill
Carbon
120-400 MeV/u
< 4 x 108 C/spill

## Slide 7: Machine cycle
Start cycle
Multiturn
Rf trapping
Acceleration
Rf jump
Rf off
Start betatron
Stop betatron
Start hysterestis
End cycle
BPM signal
BETATRON
# PARTICLES
DIPOLE field
Bmin
Bmax
Bextr
Binj
Hysteresis cycle

## Slide 8: Treatment execution
Sync
Extraction magnets
Injection magnets
1 s
Extracted beam

## Slide 9: Multi Energy Extraction in progress
Sync
1 s
Saved
time
Saved
time
(Courtesy K. Noda – CAS Medical)
Use remaining beam after slice completion

## Slide 10: HEBT
60-250 MeV p
120-400 MeV/u C
1010 p/spill (~2nA)
4 108 C/spill (~0.4nA)
3 Treatment rooms
2 H
1 H+V

## Slide 11: Dose delivery
Giordanengo et al.,NIM A613
Dose driven
Real time measurement
Feedback on scanning magnets
< Δt > = 35.1 ± 3.5 µs between 20% to 80%
ΔI /Δt ~ 170 kA/s or ~ 85 T/sec
In House developed
certified medical product
=> Continuous development

## Slide 12: Experimental room
60-250 MeV p
120-400 MeV/u C
1010 p/spill (~2nA)
4 108 C/spill (~0.4nA)
3 Treatment rooms
2 H
1 H+V
+1 Experimental Room

## Slide 13: Experimental room
60-250 MeV p
120-400 MeV/u C
1010 p/spill (~2nA)
4 108 C/spill (~0.4nA)
3 Treatment rooms
2 H
1 H+V
+1 Experimental Room
Isocenter 1
Isocenter 2
Isocenter 3
Isocenter 4
Fixed beam (single spot)
Fixed beam (single spot)
Beam measured and controlled
135 x 135 mm2
Scanned beam
200 x 200 mm2
Scanned beam

## Slide 14: Available beams
Protons
60-250 MeV
< 1010 p/spill
Carbon
120-400 MeV/u
< 4 x 108 C/spill
~2s
t
I
~1s
En.1
En.2
En.3
FWHM ~10 mm (particle and energy dependent)
Irradiation field
< 200 x 200 mm2 [iso4]
Max distance downstream target < 5.5 m
[iso1]
Low intensities (103 part/s) possible

## Slide 15: Scanning magnets and DDS on rail

## Slide 16: Isocenter 3

## Slide 17
The CNAO experimental facility: a unique opportunity to perform research activities at 360°
CNAO is also a Centre of Research and Development, whose activities range from clinical and radiobiological research to translational research with the objective of providing continual improvements in the capacity to cure.
For researchers a dedicated experimental irradiation room is available, using the CNAO beams, in time slots not impacting on patients treatment, but specifically devoted to research purposes: i.e. some night shifts (22:00 - 4:00) during the week and some full shifts (8h) during weekends.
Developed and built in collaboration with INFN

## Slide 18: CNAO radiobiology laboratory
Nearby university:
Animal house
Centro grandi strumenti
Biological laboratory with all the necessary equipment including biological laminar flows, incubators, centrifuges, fluorescence microscope, cell counters is provided. Furthermore, in the next 2 years the research area will be expanded and rooms dedicated to microscopy, cell handling, cytology/histology and small animals preparation will be accessible

## Slide 19: Upgrades in progress
From 2023 at CNAO will be available
New research area and bio-lab
Third source/new ions:
He, Li, O (and Fe)

## Slide 20: CNAO expansion project
Expansion Area A
Project approved and financed in 2019 by Ministry of Health

## Slide 21
PROTONTHERAPY

## Slide 22: Single room facility with gantry for protons
Synchrotron room
Treatment room
Contract signed with Hitachi: December 5th, 2019
Operational end 2023
360°isocentric gantry (Field size: 30x40 cm2)

## Slide 23: BNCT
BNCT AREA

## Slide 24: BNCT
Boronated drug that selectively reaches the tumour cells and avoids the healthy tissues
Accelerator driven neutron production
First Approved BNCT Drug (in Japan), May 2020
Notes: STELLA PHARMA will launch
Steboronine®, the World's First BNCT Drug,
on May 20, 2020

## Slide 25: Similar installation, Clinic in Xiamen (Fujian, China)
(Sergey Taskaev, Frascati 2019)
Notes: https://agenda.infn.it/event/19830/

## Slide 26: Accelerator based BNCT
Proton energy 2.5 MeV
Intensity 10-15 mA
p-Li reaction
Collaboration agreement signed September 2020
Operational end 2023
Notes: neutron yield at 10 mA
8.9 × 1012 [s-1]
Avg neutron energy 0.55 MeV at 0°
[/cell], [m-2] (30 minutes irradiation by flux [n/cm2/s ]):
(mostly fired)
(the half is not fired)
neutron yield at 10 mA
8.9 × 1012 [s-1]
Avg neutron energy 0.55 MeV at 0°
𝑁=〖10〗^9 [/cell], Φ=1.8×〖10〗^16 [m-2] (30 minutes irradiation by flux 〖10〗^9[n/cm2/s ]):
𝑃(0)=𝑒^(−6.84)=0.001 (mostly fired)
𝑃(0)=(⟨𝑛⟩^0 𝑒^(−⟨𝑛⟩ ))/0!=𝑒^(−⟨𝑛⟩ ) =" " 𝑒^(−𝑁𝜎Φ)
𝑁=〖10〗^9 "[/cell], " Φ=1.8×〖10〗^15 " [m−2] (30 minutes irradiation by flux " 〖10〗^8 "[n/cm2/s ]):"
𝑃(0)=𝑒^(−0.684)=0.505 (the half is not fired)

## Slide 27: Research and development are fundamental
Research and development on many subjects
  Machine operation improvement (RFKO, MultiEnergy)
  Improved dose delivery (organ motion management, 4D DDS)
  Improved treatment delivery (new ions, LET painting)
  Dose delivery and range verification (Online PET, prompt particles and gamma monitoring, improved beam monitoring)
  Improved research facility (XPR, new ions, new lab)
  Center upgrade (proton SRF, BNCT, ion gantry)
  …

## Slide 28
Thank you for your attention
“Physics is like sex: sure, it may give some practical results, but that's not why we do it. ”
R. Feynmann