## Slide 1
LOREM IPSUM DOLOR
SIT AMET
The CNAO dose delivery system and the DDS 4.0
M. Donetti

## Slide 2
Introduction
The Dose Delivery System (DDS) is the unit of the accelerator complex that acts as interface between the accelerator control system and the patient (OIS).
During treatment delivery is interfaced with master timing (MTG).
The treatment modality in use is the intensity modulated spot scanning.
Energy selection (z) from the accelerator, x and y with scanning magnets controlled by DDS.
In CNAO we use our DDS built in collaboration with the Torino section of the Istituto Nazionale di Fisica Nucleare.

## Slide 3
History
1992 Collaboration between TERA and INFN Torino to develop detectors for hadrontherapy.
2002 Collaboration between CNAO and INFN Torino to build the DDS (version 1).
2011/12 First patients with proton and carbon beam.
2012 Contract signed with MedAustron for DDS CE labelled (version 2) (3H, 1V and 1 for proton gantry).
2016 Collaboration between CNAO and GSI to develop a version for moving target (version 3).
2021 Start of development of version 4 (version 3 + CE) toward clinical use.

## Slide 4
Main components and interfaces
Detectors inside the nozzle.
Cabinet inside the technical room.
Detectors
Cabinet

## Slide 5
PXI
The DDS software is based on PXI technology by National Instruments.
Control program is executed by a real time application (LabVIEW).
Control loops for monitor DAQ and beam scanning during treatment handled by FPGA.

## Slide 6
Monitors
Two redundant monitors (BOX1 and BOX2).
5 ionization chambers:
Two large area;
Two strip (x,y);
One pixel.
Sealed with Nitrogen flow.
HV, P, T interlocks
Integral 2:
intensity;
1 MHz.
BEAM
Integral 1 :
intensity;
1 MHz.
Strip:
position;
dimension;
intensity;
20 kHz.
Pixel:
position;
dimension;
intensity;
fluence 2D;
10 kHz.
BOX1
BOX 2

## Slide 7
Monitors
Strip and Pixel
21x21 cm2
Integral
24x24 cm2
128 strip
pitch 1.65 mm
128 strip
pitch 1.65 mm
1024 pixel
pitch 6.6 mm
Segmented anodes are
glued on frames.
Very thin (25µm).
Total thickness < 0.8 mm H2O.

## Slide 8
Electronics
The front-end electronics is based on chips designed on purpose by INFN for segmented ionization chambers.
Tera chip is now in version 9.
64 channels
i → f conversion
No dead time
TERA06:
strip and pixel
negative currents
16 bit
fmax = 5 MHz
100 fC < Qc> 800 fC
TERA08:
integral
bidirectional currents
32 bit
fmax = 20 MHz
50 fC < Qc> 350 fC

## Slide 9
DDS 4.0
The DDS 4.0 is the evolution of the CE marked Medical Device in use in CNAO and MedAustron.
In collaboration with GSI. It will be able to treat moving organs.
The new irradiation modality is based on different plans for the various tumor positions.
It is called multi phase 4D delivery.
The switch between the plans is done in real time increasing the beam-on duty cycle.
In addition a tracking correction can be applied if needed.
The time needed to handle the beam energy variation has been reduced.
Possible use with multi-energy acceleration cycles or with cyclotrons.
DDS 4.0
Dose calculations
based on
measurement
results done
In CNAO
Single field
(interplay)
Multi-phase
4D delivery

## Slide 10
Publications
Strip ionization chambers as 3-D detector for hadrontherapy | Nuclear Instruments and Methods in Physics Research A 389 (1997), 499-512
Performances of a VLSI wide dynamic range current-to-frequency converter for strip ionization chambers | Nuclear Instruments and Methods in Physics Research A 405 (1998), 111-120
Dosimetric characterization of a large area pixel segmented ionization chamber | Medical Physics 31 (2), February 2004, 414-420
A pixel ionization chamber used as beam monitor at the Institut Curie – Centre de Protontherapie de Orsay (CPO) | Nuclear Instruments and Methods in Physics Research A 565 (2006) 833-840
Online monitor detector for the protontherapy beam at the INFN Laboratori Nazionali del Sud-Catania | Nuclear Instruments and Methods in Physics Research A 572 (2007) 1094-1101
Design and test of a 64-channel charge measurement ASIC developed in CMOS 0.35 mm technology | Nuclear Instruments and Methods in Physics Research A 583 (2007) 461-468
Characterization of a front-end electronics for the monitoring and control of hadrontherapy beams | Nuclear Instruments and Methods in Physics Research A 586 (2008) 270-275
Performances of the scanning system for the CNAO center of oncological hadron therapy | Nuclear Instruments and Methods in Physics Research A 613 (2010) 317-322
Design and characterization of the beam monitor detectors of the Italian National Center of Oncological Hadron-therapy (CNAO) | Nuclear Instruments and Methods in Physics Research A698 (2013) 202–207
The CNAO dose delivery system for modulated scanning ion beam radiotherapy | Medical Physics 42 (1 ) January 2015, 263-275
Real-time optical tracking for motion compensated irradiation with scanned particle beams at CNAO | Nuclear Instruments and Methods in Physics Research A 827 (2016) 39-45
Commissioning of the 4-D treatment delivery system for organ motion management in synchrotron-based scanning ion beams | Physica Medica 32, 2016, 1667-1671
A modular dose delivery system for treating moving targets with scanned ion beams: Performance and safety characteristics, and preliminary tests | Physica Medica 76, 2020, 307-316
A modular system for treating moving anatomical targets with scanned ion beams at multiple facilities; pre-clinical testing for quality and safety of beam delivery | Frontiers in Oncology 11 (2021), 620388
Extension of RBE-weighted 4D particle dose calculation for non-periodic motion | Physica Medica 91, 2021, 62–72
Dosimetric validation of a system to treat moving tumors using scanned ion beams that are synchronized with anatomical motion. | Frontiers in Oncology 11 (2021), 712126