## Slide 1
LOREM IPSUM DOLOR
SIT AMET
BNCT@CNAO: experimental and environmental dosimetry
S. Agosteo, S. Altieri, V. Conte, M. Ferrarini

## Slide 2
Neutron sources
Basical features of a BNCT field
Epithermal field | 
Energy | 0.4 eV – 10 keV
Fluence rate (cm-2s-1) | > 109
Fast neutron contamination | < 2 10-13
Thermal neutron contamination | < 0.05
Gamma contamination | < 2 10-13
Collimation | > 0.7
Beam diameter (cm) | 12-14
Epithermal field | 
Energy | 0.4 eV – 10 keV
Fluence rate (cm-2s-1) | > 109
Fast neutron contamination | < 2 10-13
Thermal neutron contamination | < 0.05
Gamma contamination | < 2 10-13
Collimation | > 0.7
Beam diameter (cm) | 12-14
(IAEA TecDoc-1223 2001)

## Slide 3
Proton beam monitoring
Transmission ionization chamber;
Faraday cup.
Irradiation field monitoring
Twin ionization chambers: gamma and neutron (no B) absorbed dose;
Fission chamber (fluence rate);

## Slide 4
IRRADIATION FIELD CHARACTERIZATION: GAMMA RAY DOSIMETRY
Graphite and CO2 ionization chamber for measuring the gamma dose in the BNCT field
INFN ENTER_BNCT project

## Slide 5: IRRADIATION FIELD CHARACTERIZATION: NEUTRON ACTIVATION TECHNIQUES
Main reaction for thermal and epithermal neutron detection (bare and Cd covered target):
  197Au(n,γ)198Au: T1/2=2.69 d, σ(0.025 eV)=98.5 b;
  115In(n,γ)116mIn: T1/2=54.15 min, σ(0.025 eV)=157 b;
  <other materials: Dy, Co, Cu, Ag.
 | Gold | Indium
Half-life | 2.695 d (198Au) | 54.15 min (116mIn)
σ0 (0.025 eV) | 98.8 b | 157 b
RI | 1560 b | 2600 b
Neutron spectrometry (fast neutrons) through activation techniques.
Reactions:
  58Ni(n,p)58Co Eth = 1.9 MeV
  59Co(n,α)56Mn Eth = 5.2 MeV
  54Fe(n,p)54Mn Eth = 2.2 MeV
  58Ni(n,2n)57Ni Eth = 13.0 MeV
  115In(n,n’)115mIn Eth = 0.339 MeV
  32S(n,p)32P Eth = 2.0 MeV
  12C(n,2n)11C Eth = 20 MeV
  27Al(n, α)24Na Eth = 4.9 MeV
  27Al(n, p)27Mg Eth = 3.8 MeV
If the fast fluence rate is high enough for giving a measurable activity.

## Slide 6
IRRADIATION FIELD CHARACTERIZATON NEUTRON SPECTROMETRY
  Superheated emulsions;
  ACSPECT (recoil-proton silicon telescope);
  DIAMON (moderator + dopes Si diodes);
  NCT-WES (moderator + SiC detectors);

## Slide 7
MICRODOSIMETRY: AVAILABLE DETECTORS
Mini-TEPC (LNL)
Silicon Telescope (PoliMi)
MicroPlus probe (LNS)
Silicon Carbide Detector (LNS)

## Slide 8
They can simulated tissue sites down to 25 nm
AVALANCHE-CONFINEMENT TEPC
Microdosimetry
Nanodosimetry
Sensitive volume:
Right cylinder
(d = h = 13 mm)
DRIFT REGION
(Cathode-Helix)
MULTIPLICATION REGION
(Helix-Anode wire)
3 electrodes independently biased
(Anode, Helix, Cathode)
Vanode > Vhelix > Vcathode
HELIX = electronic avalanche confinement

## Slide 9
Interchangeable
cathodes
(different 10B concentration)
Sensitive volume:
Right cylinder
(d = h = 13 mm)
TEPC with 10B doped walls
Measurements with and without 10B allow to determine the radiaton field components.
Può simulare siti fino a 25 nm
MICRODOSIMETRY FOR BNCT

## Slide 10
BNCT MICRODOSIMETRY: TWIN MINI-TEPC
Two TEPCs in a single device:
with bare and 10B-doped walls;
10B concentration: 100 ppm
Sensitive zone diameter: 0.9 mm
TEPC at LENA (PV)

## Slide 11
Radiation field components in the irradiation room
Neutrons from the Li target and moderated by the BSA;
Photons generated from the target;
Photons from neutron capture in the room (also @ high energies);
Photons from neutron activation.
ENVIRONMENTAL DOSIMETRY: MAJOR CONCERNS FOR BNCT BASED ON THE 7Li(p,n)7Be REACTION

## Slide 12
>Very high neutron fluence rate (~109 s-1 cm-2)
Spectrum peaked @10 keV
The ambient dose equivalent (H*(10)) rate inside the treatment room is of the order of a few Sv/h
ENVIRONMENTAL DOSIMETRY: NEUTRON SPECTRUM FROM THE BSA

## Slide 13
Instruments should face:
  High intensity fields (=>saturation)
  Complex fields (=> pile-up, instrument response against other field components)
  Unusual energy range (e.g. epithermal neutrons, high energy photons)
ENVIRONMENTAL DOSIMETRY: MAJOR CONCERNS FOR MEASUREMENTS

## Slide 14
Neutrons
LINUS
Commercial rem-meters (Wendi II – Thermo)
Lupin (FPGA based electronics and logarithmic amplifier – intense fields)
Based on micro-structured diodes
Misure di fotoni (dosimetria/spettrometria)
Active rem-meters
ENVIRONMENTAL DOSIMETRY: INSTRUMENTS
Passive rem-meters
Based on track detectors (CR-39)

## Slide 15
Bonner Sphere Spectrometers
Active thermal neutron detectors:
  LiI(Eu) scintillators (original BSS);
  BF3 proportional counters;
  3He proportional counters.
Passive thermal neutron detectors:
  TLDs;
  activation foils (Au mainly);
  track detectors coupled to 10B radiators.
ENVIRONMENTAL DOSIMETRY: INSTRUMENTS
Neutrons
Photons
Scintillators/HpGe/ionization chambers

## Slide 16
ADDITIONAL SLIDES

## Slide 17: IRRADIATION FIELD CHARACTERIZATON SUPERHEATED EMULSIONS
Courtesy of F. d’Errico, Yale Univ. and DMNP Pisa Univ.
“Superheated emulsion” is the name adopted by ISO and ICRU for detectors based on a superheated liquid suspended in a gel, also known as bubble detectors or superheated drop detectors.
  The suspended droplets consist of an over-expanded halocarbon and/or hydrocarbon which vaporizes upon exposure to the high-LET recoils from neutron interactions.
  The superheated emulsion is contained in a vial and acts as a continuously sensitive, miniature bubble chamber.
  The total number of bubbles evolved from the radiation-induced nucleation of drops gives an integrated measure of the total neutron exposure.

## Slide 18: SUPERHEATED EMULSIONS
Am-Be
252Cf
All figures in this slide:
Courtesy of F. d’Errico, Yale Univ. and DMNP Pisa Univ.
Superheated emulsions are currently used either as personal and environmental dosemeters or as neutron spectrometers.
  Neutron spectrometry is performed by exploiting the different response to neutron energy against temperature or pressure of the superheated liquid.
  Dosemeters: one of their advantages is the possibility of determining an average ambient dose equivalent rate in a pulsed neutron field.
  They are completely insensitive to low-LET radiation, X and  rays as well as muons, which is a clear advantage when measuring the neutron component in mixed fields.

## Slide 19: MEASUREMENTS AT THE LNL (1996)
Activation techniques (bare and Cd covered In foils)
Netron spectrometry with superheated detectors

## Slide 20: IRRADATION FIELD CHARACTERIZATION ACSPECT (NEUTRON SPECTROMETRY)
A silicon telescope neutron spectrometer based on:
  A plastic scintillator acting as a recoil proton conveter which measures the deposited energy;
  A silicon momolithic telescope (∆E-E stage);
  Only forward directed recoil-protons (0°-7°) are accepted, thus avoiding the use of unfolding techniques;
  γ-neutron discrimination;
  Low efficiency, for high intensity neutron fields.

## Slide 21: ACSPECT: neutrons from 9Be(p,xn)9B
[1] W.B. Howard, S.M. Grimes, T.N. Massey, S.I. Al-Quraishi, D.K. Jacobs, C.E. Brient, J.C. Yanch, J.C., Measurement of the Thick-Target 9Be(p,n) Neutron Energy Spectra, Nucl. Sci. Engin. 138(2) (2001) 145-160.
Neutrons from 5 MeV protons on su beyllium
Neutrons from 4 MeV protons on beryllium

## Slide 22: IRRADIATION FIELD CHARACTERIZATION DIAMON – Survey meter and spectrometer
dosimetric quantities of interest (H*(10) and h*(10)) | field quantities (fluence, field fractions) | 
Neutron spectrometry |  | Angular response
OUTPUTS
DIAMON can store up to 10000 different measurements.
Web access, PC, smartphone or tablet access.
User-friendly and easy to be positioned.

## Slide 23
IRRADIATION FIELD CHARACTERIZATION NEUTRON SPECTROMETRY
CONES
COmpact NEutron Sensors
A compact detector for characterizing the neutron spectrum in a phantom
NCT-WES
(Neutron Capture Therapy – Wide Energy Spectrometer)
The spectrometer is made of a cylindrical polyethylene moderator with 6 cavities along the central axis housing 6 SiC detectors. A cylindrical collimator on the detector front side rules the directional response.
INFN ENTER_BNCT project

## Slide 24
DOSIMETRY VS MICRODOSIMETRY
It measures the energy deposited in microscopic volumes per single event.
It measures the total mean energy deposited in a macroscopic volume.
DOSIMETRY
MICRODOSIMETRY

## Slide 25: MICRODOSIMETRY
Microdosimetric distributions and derived quantities (mean, secon moment or higher-order moments) can be used for assessing the radiation quality.
The radiation components can be discriminated (photons, fast and thermal neutrons, 4He e 7Li from BNC reactions).
Microdosimetry can predict the RBE of mixed radiation fields.
It measures the energy deposited in microscopic volumes per single event.
19
1 μm

## Slide 26: SILICON TELESCOPE
Monolithic silicon telescope (ST-Microelectronics, Catania, Italy) single ∆E stage
ΔE thickness : ~1.9 μm
E thickness: ~ 500 μm
Sensitive area: 1 mm2
∆E stage
E stage
9 μm
14 μm
CATANA distal to the Bragg peak
Pixelated silicon microdosimeter (telescope):
  constituted by a matrix of cylindrical ∆E elements (about 2 µm in thickness) and a single residual-energy E stage (500 µm thick);
  the nominal diameter of the ∆E elements is about 9 μm and the width of the pitch separating the elements is about 41 µm.
  more than 7000 pixels are connected in parallel to give an effective sensitive area of about 0.5 mm2.
  minimum detectable energy is limited to about 20 keV by the electronic noise.
  the ∆E stage acts as a microdosimeter and the E stage plays a fundamental role for assessing the full energy of the recoil-protons, thus allowing to perform a LET-dependent correction for tissue-equivalence.

## Slide 27: SI TELESCOPE: SECONDARY NEUTRONS @ CATANA
10B converter
Bare

## Slide 28
BNCT MICRODOSIMETRY: TWIN MINI-TEPC
Two TEPCs in a single device:
with bare and 10B-doped walls;
10B concentration: 100 ppm
Sensitive zone diameter: 0.9 mm
TEPC at LENA (PV)

## Slide 29
MICRODOSIMETRY AT THE LENA REACTOR (PAVIA)
P. Colautti et al., Microdosimetric measurements in the thermal neutron irradiation facility of LENA reactor. Applied Radiation and Isotopes, 2014.

## Slide 30
Lee et al. “A Monte Carlo dosimetry-based evaluation of the 7 Li (p,n) 7 Be reaction near threshold for accelerator boron neutron capture therapy
Medical Physics 27, 192 (2000); doi: 10.1118/1.598884
Up to 17 MeV photons are generated in the target
ENVIRONMENTAL DOSIMETRY: PHOTONS GENERATED IN THE TARGET

## Slide 31
Photons from neutron capture
Map of the dose from photons generated in the BSA (μSv/h, H*(10))
Photons are also generated in the BSA materials
(the target generates ~1013 neutrons s-1, out of which only a part exits from theBSA)
Neutrons from the BSA generate photons in the treatment room through absorption in the housed materials
ENVIRONMENTAL DOSIMETRY: PHOTONS FROM NEUTRON CAPTURE
Map of the dose from photons generated in a polyethylene target (μSv/h, H*(10))

## Slide 32
Also high energy photons are produced (~8 MeV)
X-sec libraries of the MC codes should be checked carefully.
NNDC Data – Prompt gamma spectrumin natural Pb.
ENVIRONMENTAL DOSIMETRY: PHOTONS FROM NEUTRON CAPTURE