## Slide 1
DETECTION IN HIGH POWER LASER ENVIRONMENT
Excitations Nucléaires par Laser: ENL
Emmanuel Atukpor

## Slide 2
Context: High power Lasers for ion acceleration
  Solid targets to Gas Jet targets
Main work:
  Online proton detection
    From Imaging Plates to CMOS detector
  Gamma Spectroscopy in Laser environment
    Scintillators
    Semiconductors
Conclusions and Perspectives
CONTENT SUMMARY

## Slide 3
LASER ION ACCELERATION
New kind of target for high repetition rate
1 shot per 1 hour
~1Hz
SOLID TARGET: TNSA
GAS TARGET: CSA
H+
Puyuelo-Valdes, Proc. SPIE 11037, Laser Acceleration of Electrons, Protons, and Ions V, 110370B (2019)
Puyuelo-Valdes et al., Phys Plasma, 26, 123109 (2019)
ROTH, Met al. Ion Acceleration—Target Normal Sheath Acceleration. CERN Yellow Reports, [S.l.], v. 1, p. 231, feb. 2016.

## Slide 4
4
Solenoid valve + nozzle
J.L. Henares, et al. Rev. Sci. Instrum. 90, 063302 (2019)
J L Henares, J. Phys. : Conf. Ser. 1079 012004 (2018)
Interféromètre Mach-Zehnder
With gas jet targets
Service instrumentation
LASER DRIVEN ION SOURCE ACCELERATION

## Slide 5
  ONLINE PROTON DETECTION
Imaging plates are a common tool for detection of laser accelerated protons
High repetition rate implies online detection
Imaging plates are limited with high repetition rate Lasers
Not suitable for online detection
CMOS detector
Thomson Parabola + Imaging Plates IP
IP
Thomson Parabola + CMOS detector

## Slide 6
  ONLINE PROTON DETECTION
CMOS RadEye detector calibration done at AIFIRA
Energy: 3MeV
Beam diameter: 1µm
Mean current ~1pA
Front
CMOS detector sensitive area
Quartz
Pixel dimension: 20µm
Nano beamline to fully control the sweep through different pixels of the detector
CMOS from IPHC Strasbourg
Collaboration
First test was in April 2021, LULI PICO 2000
EMP challenges
http://www.iphc.cnrs.fr/PLUME.html

## Slide 7
  ONLINE PROTON DETECTION
Obtained Calibration
y=0.009x+ 181
La France

## Slide 8
LaBr3
PMT
LaBr3
HPD
Inject LED in HPD
It is necessary to perform offline treatment and analysis
F. Negoita et al., AIP Conference Proceedings vol. 1645, p. 228, 2015
M.Tarisien, et al., IEEE Transactions on Nuclear Science, Vol 65, issue 8, p.2216-2219 (2018)
DETECTOR TEST in high power Laser environment: 90Zr(p,n)90mNb
  γ=122 keV, =63 µs,
  γ=257 keV, =6ms
NUCLEAR PHYSICS IN PLASMA: EXPERIMENTAL STUDY
Lanthinum Bromide+ PMT
~3ms
Some µs
Detector recovery time
Few µs
Full oscillogram is necessary
Scintillator Detector
Lanthinum Bromide+ Hybride Photodiode
Lots of X-rays

## Slide 9
X-ray signal baseline subtraction
OFFLINE ANALYSIS 1/2
Subsequently:
Noise treatment
Gain correction
Peak detection
Histogramming
1
2
Oscillogram
NUCLEAR PHYSICS IN PLASMA: EXPERIMENTAL STUDY

## Slide 10
OFFLINE ANALYSIS 2/2
122 keV detected
More statistics are needed
New generation of lasers >1Hz
optimised detection: Digitizer, …
=> Pileup
NUCLEAR PHYSICS IN PLASMA: EXPERIMENTAL STUDY
Service electronique et aquisition

## Slide 11
Delay, dt
MANAGE PILEUP OCCURRENCES OFFLINE
Pileup below 60ns
analysis after detection is ready!
Currently at some µs with scintillators
Semiconductors could help improve the time of detection?
Result is improved for beyond 4ns
Correction via template pulse
PILEUP SIMULATION: From LaBr3 Scintillator Captured By An Oscilloscope
LaBr3 properties
Rise time: 2ns
Decay time: 16ns
1
2
3

## Slide 12
EXPLORING SEMICONDUCTOR: Cadmium telluride (CdTe)
How it works: Drift current
no. of e-h pairs =
e-h pairs are created when particle interact in detectors
Under an Electric field, they drift to their respective electrodes
Current is induced in the process =>Ramo-shockley: I α v(e,h)
CdTe is efficient around 100 keV of gamma
Poor mobility of holes 10 times less than electrons; =100 , =1100
Cathode
Anode
+70 V
To Preamplifier
electrons
holes
e Drift
h Drift
Radiation
+
-

## Slide 13
With ATLAS SILVACO: CdTe+ γ (different positions in Crystal)
Exposing the Cathode electrode will help reduce signal duration
Near the anode
In the middle
Near the cathode
5µs
3µs
1µs
Holes travel longer distance=>longer time
cathode
Anode
3 points of γ interaction
2mm
4.83mm
EXPLORING SEMICONDUCTOR: Cadmium telluride (CdTe)
Gamma 122 keV
Transient time (s)
Anode current (A)
PILEUP CONCERNS!!!
+
-

## Slide 14
X-flash signal (100MeV of deposited x-ray energy)
Gamma signal 122 keV
Ideal case: CdTe+X-flash+γ
Current A
Time s
When can the 122 keV signal be found on the flash signal if it arrives after the flash?
cathode
Anode
2mm
4.83mm
Homogeneous distribution of charges from X-ray flash
Gamma 122 keV
X-flash
100MeV
t≈0
t+Δt
cathode
Anode
2mm
4.83mm
EXPLORING SEMICONDUCTOR: Cadmium telluride (CdTe)
+
-

## Slide 15
Long duration=>Hole drift and electron diffusion current
Within some µs for detection (for 100MeV)
undetected 122 keV
@10ns and 100ns
detected 122 keV@3µs
X-ray flash signal tail
EXPLORING SEMICONDUCTOR: Cadmium telluride (CdTe)
Real case: CdTe+X-flash+γ
X-flash(t=0)+γ(10ns)+γ(100ns)+γ(3µs)
Transient time (s)
Anode current (A)
T=0

## Slide 16
CONCLUSIONS AND PERSPECTIVES
Back to scintillators=> Filter emission light from Scintillator
Use 14 bits ; 2Gs/s ; 1GB memory digitizer
Try silicon=> with more thickness (good hole and electron mobility)
Further test experiments next year: CELIA/LULI/CLPU
Suggestions are welcomed
Online particle detection is necessary for high repetition rate lasers
CMOS (IPHC+RadEye) detector calibration is mandatory for online proton detection
Large Quantities of Xrays =>A major limitation of 𝜸 spectroscopy in Laser Plasma
With scintillators => Deadtime: 10s of µs
CdTe considered:
    detection possible within few µs
    Pileup concerns

## Slide 17
THANK YOU FOR LISTENING
Service electronique et aquisition
Service instrumentation
Service Mecanique/BE
Service Administration (Nadine )
MERCI

## Slide 18
Backup slides

## Slide 19
Carroll, D. C et al.,New Journal of Physics 12 (2010) 045020 (15pp)
Energies reached for C and H ions
Macchi et al. Rev. Mod. Phys, 85, 751 (2013)

## Slide 20
  ONLINE PROTON DETECTION
CMOS detector calibration done at AIFIRA
Picture of cmos
Some irradiation
Calibration

## Slide 21
OFFLINE ANALYSIS
PMT
HPD
HPD has a wider range compared to PMT

## Slide 22
OFFLINE ANALYSIS
Combining successive LEDs we can restore each gamma pulse to their original amplitudes
X-ray signal baseline subtraction
Noise and EMP perturbation treatment
Gain correction

## Slide 23
  ONLINE PROTON DETECTION
CMOS detector calibration done at AIFIRA
Energy: 3MeV
Beam diameter: 1µm
Mean current ~1pA
back
Front
Proton beam
Motor to move detector vertically
CMOS detector sensitive area
CMOS detector
Quartz
Detector Holder
Pixel dimension: 20µm
In the experimental Chamber
Nano beamline to fully control the sweep through different pixels of the detector

## Slide 24
LaBr3
PMT
LaBr3
HPD
Inject LED in HPD
It is necessary to perform offline treatment and analysis
F. Negoita et al., AIP Conference Proceedings vol. 1645, p. 228, 2015
M.Tarisien, et al., IEEE Transactions on Nuclear Science, Vol 65, issue 8, p.2216-2219 (2018)
DETECTOR TEST in high power Laser environment: 90Zr(p,n)90mNb
  γ=122 keV, =63 µs,
  γ=257 keV, =6ms
NUCLEAR PHYSICS IN PLASMA: EXPERIMENTAL STUDY
Lanthinum Bromide+ PMT
Scintillator Detector
Lanthinum Bromide+ Hybride Photodiode
90Zr
Ps/fs Laser 1
10-90J
84mRb
µm target
C ions
Detector
Lots of X-rays
~3ms
Some µs
Detector recovery time
Few µs
Full oscillogram is necessary

## Slide 25
WITH SEMICONDUCTOR HOW CLOSE CAN WE GET FOR 122 keV?
Ωcm
50Ωcm
Ωcm
Exploring Cadmium telluride (CdTe) : 90Zr(p,n)90Nb

## Slide 26
SILVACO SOLVES THE FOLLOWING EQUATIONS
Modelling CdTe detector with SILVACO
Electron current
Hole curent
Displacement current
cathode
Bulk crystal in blue
Radiation pass through the anode Electrode

## Slide 27
No detector saturation
Long duration=>Hole drift and electron diffusion current
undetected 122 keV
@10ns and 100ns
detected 122 keV@3µs
X-ray flash contribution
EXPLORING SEMICONDUCTOR: Cadmium telluride (CdTe)
Real case: CdTe+X-flash+γ
X-flash(t=0)+γ(10ns)+γ(100ns)+γ(3µs)
Transient time (s)
Anode current (A)
T=0
Hole drift + e diffusion current
e drift + h drift

## Slide 28
Long duration=>Hole drift and electron diffusion current
Within some µs for detection (for 100MeV)
undetected 122 keV
@10ns and 100ns
detected 122 keV@3µs
X-ray flash contribution
EXPLORING SEMICONDUCTOR: Cadmium telluride (CdTe)
Real case: CdTe+X-flash+γ
X-flash(t=0)+γ(10ns)+γ(100ns)+γ(3µs)
Transient time (s)
Anode current (A)
T=0

## Slide 29
76Ge
Ps/fs Laser 1
10-90J
84mRb
ns Laser 2
30kJ
µm target
C ions
Detector
Lots of X-rays
Modification of nuclei apparent lifetime in plasma: 76Ge(12C,p+3n)84mRb
NUCLEAR PHYSICS IN PLASMA: EXPERIMENTAL STUDY
We need carbon ions for this reaction
High power lasers can also accelerate ions and produces Plasma
Huge quantities of X-rays!!!
Our detectors will be blind
HAPPY!!!