## Slide 1: Laser Ablation ICPMS technique and potentialities - Status of the round table
Frédéric Perrot for WP2 group
LP2i Bordeaux / IN2P3-CNRS / University of Bordeaux
GDR DUPhy, 19-21 octobre 2022, Nantes

## Slide 2: Natural radioactivity : our main ennemy!
Primordial radioactivity, cosmogenics, man-made radioactivity
Long-lived isotopes~109 y (U/Th/K)
NAA, ICPMS, LA-ICPMS techniques
 Complementary techniques to screen all the critical materials for rare event experiments
Gaseous isotopes (222Rn)
 Dedicated techniques
Medium and short-lived isotopes ≤105 y (226Ra, 228Th, 210Pb/210Po, 60Co…)
low background γ/α/β spectroscopy techniques

## Slide 3: Principle of the ICP-MS technique
Inductively coupled plasma mass spectrometry (ICP-MS)
Sample liquid phase to formate aerosol in nebulizer
Introduction of aerosols into an argon plasma at T>6000 K
Ions extracted from the plasma into mass spectrometer region and detected by an electron multiplier sensor
 Need to have very radiopure acids and vessels to achieve a very low blank

## Slide 4: Principle of Laser Ablation ICP-MS
Concept: to replace the chemical preparation of a sample by using a laser in ablation mode
Main advantages of LA-ICPMS:
small risk of sample contamination
fast access to the spatial distribution of the U/Th radionuclides
A laser to drill, an ICP-MS to count…

## Slide 5: Shaping technique of the fs laser beam
Galvanometric
scanners
HR-fs LASER
Possibility to move in 2D the sample and/or the laser beam in order to fit with the morphology of the sample
Size of the beam ~10 µm
High repetition rate <100 kHz
Translation speed < 2 m/s
Single spot,
(5-200 µm)
1D scan
2D scan

## Slide 6: LA-ICPMS platform @IPREM Pau
IR or UV fs laser: ΔT=360 fs, f=400 Hz, E~15 µJ/pulse
Ablation chamber: rinsing time~500 ms
ICP-MS/MS (Agilent 8900)
ICP-MS/MS
Laser IR fs
Acrylic sample
Ablation cell

## Slide 7: Ablation of acrylic sample

## Slide 8: LA-ICPMS used for acrylic
25 mm
1 mm
Ablation of squares of 1x1 mm2 to study the homogeneity of the U/Th surface contamination depending of the surface treatment
Laser ON
Acrylic sample
Laser OFF
Laser OFF
12C
232Th
238U
Obtained in ~2 mn

## Slide 9: Effect of polishing/no polishing
One order
of magnitude
less 232Th !
Example of two acrylic samples measured:
one with polishing process applied on the top surface
one without polishing
 very powerful tool to measure the U/Th contamination and to choose the best surface treatment of a given material
w/o polishing
w polishing
232Th

## Slide 10: Depth profile of U/Th concentration
Laser
scans
600 µm
Acrylic
Several 2D laser scans
to explore the Z-axis
surface
1 scan
4 µm depth
3D profile of acrylic sample by confocal microscopy
2 scans
8 µm depth
Bulk concentration
Surfacic
contamination
 U and Th contamination of the acrylic is located at the surface in the first micrometers
Good reproducibility of the ablated volume after several scans

## Slide 11: Quantification of U/Th concentration
Quantification of U/Th concentration C in g/g :
A posteriori measurement of the ablated volume by confocal microscopy
Counting the total number of ions during laser ON
Dedicated method to determine the ion transmission from ablation cell to ICPMS

## Slide 12: Sensitivity of LA-ICPMS technique
Assuming 0.2% efficiency, 2 µg/mn of ablated mass of acrylic
Minimum detectable activity in g/g vs time (proportionnal to mass) for different background rates
10 mn
1 h
sub-10-12 g/g level reached within 10 mn @ 1 ion/s bckg rate
238U in acrylic
0.6x10-12 g/g
0.08x10-12 g/g
sub-10-13 g/g level reached within 1 h @ 0.1 ion/s bckg rate
Opportunity to measure several samples per day at the sub-ppt level with LA-ICPMS !

## Slide 13: 2D mapping of a µ-coaxial cable
400 µm
What is the U/Th mapping impurities on the section of the cable ?
µ-coaxial cable produced by the AXON company (France)
13C
65Cu
238U
U contamination observed
in some of the copper wires

## Slide 14: Collaboration between LP2i Bx & IPREM
2h driving
distance
Spain
Nouvelle Aquitaine region

## Slide 15: Round table discussion summary
Christophe Pecheyran is the expert at IPREM Pau on LA-ICPMS in several fields (chemistry, environment, …)
He gave a talk at the 1st GDR DUPhy meeting on Nov. 2021: “Direct ultratrace and isotopic analysis in solids by femtosecond LA-ICPMS”
ESR-EquipEx+ application PLATTO (PLatform for elemental and isotopic ATTO-trace analysis) submitted in 2020 by LP2i Bordeaux/IPREM institutes to develop new instruments for several applications (including rare event experiments)
 unfortunately not successful

## Slide 16: Round table discussion summary
There is clearly an opportunity in France to develop the LA-ICPMS technique for our field (Dark Matter and Neutrino) taking advantage of the expertize at IPREM Pau
WP2 would like to push IN2P3 to have a 2 years CDD position to develop specific protocols at IPREM Pau to quantify U/Th bulk/surface contamination at sub-10-12 g/g for several materials of interest

## Slide 17: Summary
fs laser ablation ICPMS is a complementary technique to ICPMS
  No chemical preparation
  Fast analysis
  Spatial information for micro-analysis
Promising applications of LA-ICPMS technique
  A tool to select the best surface treatment
  A tool to test the surface homogeneity in U/Th of the sample
  A tool to map the U/Th concentration in 2D and 3D
  Achievable sensitivity at the sub-10-12 g/g level with only few tens of µg (precise quantification methods still in development)
Need to request human ressources (CDD) to develop LA-ICPMS quantitative protocols for several materials with experts at IPREM

## Slide 18: SPECTRATOM 2022 conference
Very important conference organized every 2 years in Pau around spectrometry techniques applied to several fields
Next one in 2024

## Slide 19
Thank you for your attention

## Slide 20
Backup slides

## Slide 21: Mass spectrometry vs γ/α/β spectroscopy
Concentration vs Activity
Mass spectrometry: well-suited for long-lived isotopes to reach concentration <10-12 g/g
Gamma spectroscopy: well-suited for short-lived isotopes to reach activity >10 µBq/kg

## Slide 22: Sample preparation and sensitivity
Several equipments are needed for the pre-treatment of the sample without any contamination
Class 10000 clean room
Ultra-pure deionized water, nitric acid, microwave digestion, ultra-sonic cleaner…
it is critical to have ultra-clean vessel and acids for blank test in order to reach a sub-10-12 g/g sensitivity for U/Th in critical materials
Not extensively developed in France for our field but several installation around the world (for example PNNL, USA)

## Slide 23: Advantages of the UV fs LA-ICPMS
Poweful tool for micro-analysis
  Quasi non-destructive technique
  Micro-analysis from few µm to few 1000 µm
  2D mapping to study the homogeneity of a given surface
  High precision depth profile (sub µm)
Alternative method to classical ICPMS
  Direct and fast analysis without chemical preparation
  Smaller risk of sample contamination
  Access to the refractory matrix samples

## Slide 24: Protocole de mesure et bruit de fond
Acrylique (C5H8O2)n  mesure de la masse 52 (12C+40Ar) pour l’estimation de la quantité de matière ablatée
Dwell time : 10 ms (52Cr), 52 ms (238U), 52 ms (232Th)
Laser ON pour la mesure de l’échantillon
Laser OFF pour la mesure du bruit de fond (effet mémoire)
excellents fonds !
238U: ~0,09 ion/s
232Th: ~0,06 ion/s
Laser ON
Laser OFF
Laser OFF