## Slide 1: JUNO sPMT ABC Front-End Board
C. Bordereau, C. Cerna, T. Chabaud, F. Druillole, P. Hellmuth, C. Huss,
L. Labit, P. Perrot, A. Rebii
On behalf of SPMT group
LP2IB, University of Bordeaux/ IN2P3

## Slide 2: JUNO Experiment: International cooperation
694 physicists mainly from China and Europe, 77 institutes including 5 at IN2P3

## Slide 3: JUNO Experience : Goal
Determine the order of neutrino masses using electron antineutrinos emitted from two nuclear power plants in China
Daya Bay
D. Chooz
RENO
JUNO
Oscillations
rapides
Oscillations
lentes
Réacteurs nucléaires
(~1021 ν/s/GW)
Détecteur JUNO
Base de vol ~53 km
Precision measurement (<1%) of 3 neutrino oscillation parameters
Measurement of solar neutrinos, Supernova neutrinos , geo neutrinos, …
Oscillation des neutrinos

## Slide 4: JUNO Experience : Detector

## Slide 5: JUNO Detector Performance
Detect 100,000 neutrinos in 6 years of data collection  20,000 tons of liquid Scintillator (LS) in a sphere with 35 m diameter
Energy resolution of 3%/√E(MeV)
   System of 18000 PM 20’’ and 26000 PM 3’’ (78% angular coverage) to detect ~1500 photoelectrons / MeV

## Slide 6
JUNO – SPMT : (FRANCE, CHINA, CHILE)
  PMTs:
    25 600 photomultipliers 3’’
      CHINA (HZC)
  Electronic
    Front End Electronics for sPMTs
      France: IN2P3
    High Voltages Splitters
      Chile
    Global Control Modules
      China
  Mechanical
    200 UWB (Under Water Box)
      France , Chile, China
JUNO sPMT Project
1st UWB
2nd UWB
Notes: JUNO, est à ce jour le plus grand détecteur neutrinos au monde, et qui sera implanté en chine pour la mesure des paramètres d’oscillation de neutrinos artificiels type électron produit par deux réacteurs EPR.
L’instrument de ce projet, est une sphère acrylique, immergée dans de l’eau pure dans laquelle il y aura un liquide scintillant. Sur cette sphère seront fixé 17000 photomultiplicateurs 20 pouces ou LPMT et 25 600 pm 3 pouces SPMT
On distingue ainsi deux projets sur le même instrument : JUNO LPMT et JUNO SPMT
C’est SPMT sont regroupés en paquet de 128 avec une électronique Front End dédiée, c’est la contribution française à ce projet.

## Slide 7
DDR3
Mem
1Gbytes
SPI
Flash
JTAG
Slow Ctrl
Data Capture
DDR3
Memory
Management
CATIROC 0
Calibration
JTAG Switsh
CATIROC 1
CATIROC 7
GCU
Interface
USB 2.0
FTDI 2232H
In_Calib
Kintex 7
325T-FFG625
Local
Clock
Monitoring
40 Mhz
Slow control signals
Data signals
Ext. Trigger &
Inspection
ABC
Operational
In progress
Trigger
Manager
&
DDS
SAMTEC
SAMTEC
From HVSplitters
GCU
JTAG Interface
JTAG Connector
USB
Interface
GCU
Interface
Temperature & humidity Sensor
ADC
Voltages
Currents
JUNO sPMT : ABC Block Diagram

## Slide 8
12 Layers
  5 Powers plans
compatible with IPC standards
  IPC 7351, IPC 2222, IPC 2221
  IPC 4101
1902 Components
  CMS technology
45 differential pairs
JUNO-SPMT : ABC PCB Overview
Layer | Thickness (µm) | Single impedance | Differential impedance
1 | 45 | 50 | 100
3 | 18 | 40,50 | 
4 | 18 | 40,50 | 
6 | 18 | 40,50 | 
8 | 18 | 40,50 | 80,100
10 | 18 | 40,50 | 80,100
12 | 45 | 40,50 | 

## Slide 9: ABC Production and Tests at FEDD company
PCB Manufacturing
PCB control
review by
Components control review
Boards serialization:
Cleaning & steaming
CMS and feedThrough
Components soldering
Visual inspection
BGA X Ray Inspection
Placement Inspection
Automatic Optical Inspection
Final Quality Control
Burning Test
Takaya test
Functional tests
Tropicalization
Coating
Packaging shipping
Components
provisioning
ISO9001
ISO13485
EN9100
IPCA600: printed circuits
IPCA610: card wiring
IPCA620: wired cabling
Test bench from LP2I
FPGA Programming
CATIROC ship slow control
All interfaces verifications
Functional performances
DAQ test
  Pedestal noise
  Charge injection
  HG/LG verification
  Trigger verif.
  Coarse/Fine Time verif.
220 ABC boards produced
June’21  Jan’22

## Slide 10: LP2IB Combined test bench
128 3’’ JUNO PMTs
Control Power and readout of ABC through GCU (V1.1)
Control and power of HV Splitter boards (V0.3.2) through GCU
JUNO grounding scheme policy

## Slide 11: LP2IB Combined test bench
128 JUNO 3’’ PMTs in a dark room

## Slide 12: Acceptance and calibration test sequence
Sequence step by step and deliverable
Installation and ABC powered
Pedestal measurement in High Gain
HV OFF/ON
S-curve and
threshold tuning
PMT measurement
In Dark Rate mode
@ nominal HV and nominal thresholds
 HG Meanchannel
 HG Sigmachannel
 LG Meanchannel
 LG Sigmachannel
 Charge Resolution
 Dark Rate
 Nominal trigger thresholdsAsic
 Fine charge calibration parameters in HG with SPE
LED ON
@ nominal HV
 HG/LG verification and rough charge calibration
Pedestal measurement in Low Gain
HV OFF/ON
HV variation from 1100 to 1400 V by step of 50 V
All the sequence lasted ~25’ and was fully automatized
 Control of the current (~0.5-0.6 A)

## Slide 13: ABC readout front-end board performances from production
Low noise achieved
  ～4% PE / <1% PMT resolution
Mean and sigma pedestal distributions
Calibration parameters in High Gain
High Gain range [1/3 - 10] PE @Gain 20
Calibration slope in Low Gain
Low Gain range [10 - 150] PE @Gain 20
Front-End boards parameters delivered for integration
SPE + LED
LED only

## Slide 14: ABC mass testing: summary
220 ABC front-end boards have been produced, delivered and tested
212 have been accepted for the SPMT system (96.4%)

## Slide 15: JUNO-SPMT ABC Firmware Progress
CENBG
ABC Firmware 1.0 (2021-2022): deployed for PMT testing and UWB integration
ABC Firmware 2.0: Q/T data stream
  perfectly fits CATIROC performances
ABC Firmware 3.0: Q/T + Discriminator Data Stream (DDS)
  Trigger width gives information on the charge
  (allows to recover charge of hits triggered during
  charge integrator dead-time)
  First measurement: DDS principle is working !
  To be finalised and delivered in 2023

## Slide 16: ABC Firmware 1.0
GCU
Interface
Synchro-nisation
Deserializor
Triggers
Handler
« Event Header » generator
HitRegister
Evt TS
Catiroc data streaming
Catiroc Triggers
128 channels Registers
Event Builder
72
RD
EMPTY
Top_Data_Capture
GCU
IPBus
Master
SPI Master
Ethernet link
SCLK[7..0]
DOUT[7..0]
SELECT
RST
SPI Write
Memory
SPI Slave
RD
DOUT
WR
@
DIN
Catiroc Slow
Control
Signals
IPBus
FIFO
16 bits
Custum
protocol
ABC
SPI Bus

## Slide 17: ABC Firmware 2.x
Event Builder
GCU
Interface
Catiroc data streaming
64
RD
EMPTY
GCU
IPBus
Master
SPI Master
Ethernet link
SCLK[7..0]
DOUT[7..0]
DIN[7..0]
SELECT
RST
SPI Write
SPI Read
Memory
SPI Slave
RD
DOUT
WR
@
DIN
8
Catiroc Slow
Control
Signals
IPBus
FIFO
16 bits
Custum
protocol
ABC
SPI Bus
block_data_Capture #0
block_data_Capture #0
block_data_Capture #0
Data capture is done in a parallel mode
Blocks Readout

## Slide 18: Catiroc trig. deadtime vs Firmware 2.x deadtime
Catiroc trigger deadtime : due to trigger processing in analog part of circuit
Δt > 90 ns
2 signals
Δt < 60 ns
1 signal
Using ABC board and firmware 2.0 with pulse generator in burst mode in the same channel
Δt > 90 ns
2 signals
Δt < 60 ns
1 signal
the curve of the trigger efficiency is not affected by the Firmware 2.0
No additionnal deadtime added by the firmware 2.0

## Slide 19
JUNO-SPMT ABC Remote Programming
SPARTAN 6
ETH SWITCH
KINTEX 7
JTAG MASTER
KINTEX 7
FLASH
TCK
TMS
TDO
TDI
IPBUS
IPBUS
FLASH
GCU
ABC
JTAG MASTER
TCK
TDO
TMS
TDI
Firmware structure
Install Master JTAG ON GCU Kintex7
  Dedicated firmware on GCU
2. Configure ABC Kintex7 as Gateway for Flash programmer
  Dedicated firmware on ABC
3. Install ABC Flash Programmer on GCU Kintex
  Dedicated firmware on GCU
Flash Programmer
4. Download a Bitstream to ABC Flash
  FPGA_Prog Software
The whole procedure is automated with Shell scripts, and can be applied to the 200 ABC boards of the experiment
JTAG Sp6 – K7
JTAG K7 – K7

## Slide 20
The End