## Slide 1: Time Synchronization Scheme for the Hyper-Kamiokande experiment IN2P3-IRFU-INFN
25/10/2022

## Slide 2: Context
LPNHE, CEA and INFN have joined forces to design a full time distribution scheme that has been submitted to the international collaboration in April 2022. The group will be in charge of its construction, commissioning and maintenance.
The proposal envisions to provide a system that fulfills the needs of the HK far detector. Possible synergies could be found with the time generation system at JPARC, the Intermediate Water Cherenkov detector and the future ND280 upgrades (when integrated in HK).
To conceive this proposal, an intense and fruitful R&D campaign has been carried out. The designed time distribution system meets all the far detector requirements.
This program takes advantage of the experience that group members have acquired in designing similar systems for other experiments and of a very fruitful collaboration with the SYRTE lab (Observatoire de Paris) responsible for the French UTC implementation.

## Slide 3: HK Clock Distribution Scheme
WR switch
WR switch
GNSS
Atomic Clock
1st distribution stage
Cavern
Time data and UTC (ethernet)
5MHz
GNSS
Redundancy
Time Ref/
fanout
Time Ref/
fanout
2nd distribution stage
FE
1st distribution stage
2nd distribution stage
FE
…
FE
FE
…
…
Clk and data
(redounded link)
Clk and data
(redounded link)
DAQ
network
125MHz &
data
Sync commands
Slow control
Slow
control
Calibration
Commands and data
/
10
GNSS
GNSS
PPS/125MHz
PPS/10MHz
Time/
Freq comp.
Atomic Clock
5MHz
PPS/10MHz
PPS/125MHz
5MHz
/
/
2
5MHz
125MHz &
data
WR switch
WR switch
10MHz, PPS & Data
WR
Network
WR
Network
10MHz, PPS & Data

## Slide 4: HK Clock Distribution Scheme
WR switch
WR switch
GNSS
Atomic Clock
1st distribution stage
Cavern
Time data and UTC (ethernet)
5MHz
GNSS
Redundancy
Time Ref/
fanout
Time Ref/
fanout
2nd distribution stage
FE
1st distribution stage
2nd distribution stage
FE
…
FE
FE
…
…
Clk and data
(redounded link)
Clk and data
(redounded link)
DAQ
network
125MHz &
data
Sync commands
Slow control
Slow
control
Calibration
Commands and data
/
10
GNSS
GNSS
PPS/125MHz
PPS/10MHz
Time/
Freq comp.
Atomic Clock
5MHz
PPS/10MHz
PPS/125MHz
5MHz
/
/
2
5MHz
125MHz &
data
WR switch
WR switch
10MHz, PPS & Data
WR
Network
WR
Network
10MHz, PPS & Data
LPNHE
CEA
LPNHE
INFN

## Slide 5: The Clock Generation & UTC Conceptual Scheme
The local time is generated (by a precise atomic clock) in the form of 5MHz (HK local time).
The time reference point custom board receives this frequency, generates a local PPS and distributes them to the GNSSs and the first distribution stage. It has extra ports to compare the main and spare base cadence.
The first distribution stage produces the clock which is sent via the second stage to the FE.
The events will be tagged with the HK local time and converted in UTC by the DAQ. A correction algorithm will be applied to compensate the atomic clock drift.

## Slide 6: Time Base and UTC Characterization
All the main chain components have been purchased and an extensive test campaign has been performed to precisely characterize the proposed scheme
At LPNHE we have:
2 SRS SF725 Rubidium atomic clocks
2 Septentrio PolaRx5TR receivers + antennas
2 clock and Frequency counter Keysight 53220
1 PH1008 Passive Hydrogen Maser Atomic clock
We have also a copy of the UTC(OP) PPS and 10MHz clock sent via white Rabbit link from the SYRTE laboratory (where it is produced). This is very important time reference for comparative analysis.
Work performed by M. Guigue, L. Mellet and V. Voisin

## Slide 7: Time Base and UTC Characterization Results and correction studies
Work performed by M. Guigue, L. Mellet and V. Voisin
Measured performances of the atomic
clocks and GNSS receivers
Simulated stability of the local time base after
UTC conversion and correction
The work on the correction algorithm is progressing and we are planning to test it on real data in the next 6 months.
Tests at HK and JPARC site are under discussion. We hope to perform them by March 2023.
The local time base to UTC and JPARC site expected uncertainty is about 10 ns (100 ns required).
More detail in the L. Mellet talk https://hkdbweb.in2p3.fr/doc/meetingfiles/A00004681/Clockstudies_HKpremeeting.pdf

## Slide 8: First Clock Distribution Stage Board Demonstrator
  Principles
  Generic board for R&D on clock distribution based on Trenz SoC with Xilinx UltraScale+ FPGA
  Command interpreter on SoC (bare metal, Linux in final system) + Gbit Ethernet link to control PC
  Mezzanine card interfaces to external reference master clock
  Optical out – 1 Gbps data mode
  Rj = 11 ps and Dj < 13 ps
  Current results
  48 TX ports tested in pure clock distribution mode, modulated clock mode and classical serial data transmission. Runs up to 1 Gpbs and shows low jitter Rj typ. ~10 ps Dj typ. <20 ps
  Stability of input to output latency at level of 10 ps after board power cycling, and comparable stability of relative phase alignment between different optical outputs
  48 RX ports tested for 125 MHz clock feed back reception. Logic based on DDMTD technique being implemented to measure stability of distributed clock via round-trip latency monitoring
Work performed by CEA-IRFU D. Calvet and E. Molina

## Slide 9: First Clock Distribution Stage Board Next Steps
Work performed by CEA-IRFU D. Calvet and E. Molina
  48-port demonstrator – exploitation
  Continue firmware/software functionality developments and device characterization
  Incorporate clock mezzanine card – production suffered delays; board now expected mid-nov
  Validate specific blocks and various concepts as much as possible
  1st stage clock distributor prototype – in design
  Same concept architecture as 48-port prototype and re-use of many of its validated parts.
  Lower density of 24-ports allowable by the use of passive optical splitter in the TX direction and because only a small fraction of 2nd stage distributors will have a return optical link to the 1st stage

## Slide 10: Second Clock Distribution Stage Board Concept
Work performed by LPNHE (s. russo and E. Pierre)
The board receives the clock from the first stage and transmits the it to the time distribution endpoint in the front-end embedded into data via optical fiber.
The main element is an FPGA with embedded processors mounted on Trenz module TE0808.
It runs a custom firmware and software working under a Linux OS (Debian).

## Slide 11: Second Clock Distribution Stage Board Tests and Next Steps
Work performed by LPNHE and INFN (s. russo and F. Ameli)
Clock and data distribution have been tested with EVBs and already available boards. The jitter at the endpoint has been measured at 2.4 ps (requirement 100 ps) after a PLL jitter cleaner.
The firmware and software development continues on EVB and it will be ported on the first prototype as it arrives. The porting time will be very short due to the high compatibility of the two systems.
Tests will be then performed against the DPB prototypes already in our possession (thanks to C. Toledo and M. Ziembicki)
To absorb long delay on procurement a tender for the final production mezzanine purchase has been done. The process will be finalized by mid November.
Jitter measured at the
Endpoint = 2.4ps

## Slide 12: Time Distribution Endpoint
Work performed by LPNHE and INFN (s. russo and F. Ameli)
The time distribution endpoint will be embedded in the Data Processing Board.
We support the DPB development providing firmware/hardware IPs and components.
A full chain composed from the second distribution stage, the DPB and digitizer will be installed at CERN as soon as a suitable space will be available (negotiations started).
To absorb long delay on procurement a tender for the final production SFP transceivers purchase has been done. The process will be finalized by mid November.

## Slide 13: Time Distribution Schedule
The work is on schedule. The production is expected to be concluded in time for the assembly and commissioning in Japan

## Slide 14: Summary
A great synergy and collaboration has been established between IN2P3, CEA, INFN.
The full time distribution scheme has been designed.
Test and characterization campaigns on different subsystems have been carried out.
First prototypes have been designed and are ready/fabricated now.
The system will be tested together with the DPB and digitizer soon

## Slide 16: Clock Generation & UTC - Connections
The local time is generated (by a precise atomic clock) in the form of 5MHz in synch to each other (HK local time).
The time reference point custom board receives the frequency, generate a local PPS and distributes them to the GNSSs and the first distribution stage. It has extra ports to compare the main and spare base cadence.
GNSSs use the input cadence to sync the outputs and have a reference frequency more stable than the internal oscillator

## Slide 17: Clock Generation & UTC - Connections
2 GNSSs receivers (connected to the same antenna) get information from the satellites and measure the time distance between the local PPS and UTC.
The result of this measure is sent to DAQ and used to transform the local tag to UTC
The DAQ monitors the status of the GNSSs and select which one to use

## Slide 18: Clock Generation & UTC - Redundancy
The full chain (atomic clock, GNSSs (x2), and first distribution stage) is replicated for redundancy
Having a total of 4 GNSS receivers will allow the DAQ to identify if one is malfunctioning . Extra different brands receivers will be also used.
The time distance between the 2 chains will be constantly measured and the result sent to DAQ. This info will be used in case of switch from one chain to the other

## Slide 19: Clock Generation & UTC - Reliability
Scheduled interventions on components will be performed to reduce the risk of failures
A third atomic clock will be available at far detector site and periodically turned on. It will serve as cold spare used if the main or the hot spare needs maintenance
All the atomic clocks will be measured against the so-called travelling station every 3 / 4 years.
The GNSS receivers will be changed before the end of life.
The maintenance cost needs to be included in the budget request

## Slide 20: Septentrio PolaRx5TR calibration
Our receiver (RINEX name: LPN1) was calibrated against one at Observatoire de Paris (OP73) in turn calibrated against a travelling station from Bureau International de Poids et de Mesures (BIPM, responsible of defining UTC). OP73 is then a so-called level 1 receiver and LPN1 has been calibrated as a level 2.
The receiver internal delays were measured for both GPS and Galileo constellations.
The combined uncertainty of the LPN1 station hardware delays calibration is 4 ns. This is the conventional uncertainty applied after a “direct calibration” in the TAI system.

## Slide 21: Septentrio PolaRx5TR calibration
The two tables correspond to GPS (P1 and P2 codes) and Galileo (E1 and E2a) codes.
The values will be configured in the receiver and a correction will be applied to the UTC time. Mixing the two constellations' data will be done to provide the UTC time.

## Slide 22: Time Base and UTC Characterization
The work on the time base characterization continues and an extensive test campaign is ongoing (thanks mainly to Lucille Mellet, Vincent Voisin and Mathieu Guige).
Most of the results are summarized here:
https://indico-sk.icrr.u-tokyo.ac.jp/event/7332/contributions/22011/attachments/20557/26254/LPNHE_timingsystem_R%26D.pdf
I give just a quick recap:

## Slide 23: Kind of Test Performed
We basically perform 2 kind of tests:
Local time base characterization. Test our proposed concept on real hardware
Time transfer via common view: a method to measure two local time bases using the GNSS data. This method is implemented using the files published regularly by UTC(k) labs (Syrte in our case).

## Slide 24: Local Time Base Characterization
For this test we use 1Rb clock sending 1 PPS and 10 MHz clock to the Spetentrio receiver. Thanks to the time transfer algorithm, we can estimate the drift of our local time base with respect to GNSS time (only GPS in this case)
Rb clk
GNSS
PPS
10 MHz
CGGTTS file with time distance
Between local time and GPS time
CGGTTS

## Slide 25: Local Time Base Characterization
The “raw” plot (left), as produced from the receiver, shows a liner drift, as expected for the Rb clock. Once removed we can see that the clock “noise”.
(from Lucille’s talk cited before)

## Slide 26: What Do We Do With It?
Extracting the drift from the left plot will allow us to precisely correlate the local time base (used to time tag the events) to GNSS time. Once done we can calculate the uncertainty on it (right plot). In this case is about 4 ns over 6 days (random walk).
(from Lucille’s talk cited before)

## Slide 27: How Accurate Is This Method?
To evaluate our drift estimation method, we have done the following measure:
- We have created a CGGTTS form the GNSS steered by Rb clock (as before).
Then we have used the time transfer technique to measure the time distance between our Rb clock to SYRTE UTC(OP) via common view method.
Finally, we have compared this time distance with the one measured via the counter.
Rb clk
GNSS
PPS
10 MHz
CGGTTS file with time distance
Between local time and GPS time
Freq/Time cnt
SYRTE
GNSS
CGGTTS
CGGTTS
Time distance between
local Rb and UTC(OP)
WR switch
UTC(OP)
PPS 10 MHz
PPS
10 MHz
PPS
WR switch
Vincent Voisin work

## Slide 28: How Accurate Is This Method?
Time distance between local Rb and UTC(OP): Common view vs White Rabbit
Vincent Voisin work
3 days of data. X axis: local time Y axis: ns
Local time base/ UTC(OP)
Via common mode (blue)
Local time base/ UTC(OP)
Via white rabbit (red)
(same slope as before)
Common view method/WR method difference
+/- 3 ns max
Histogram of the difference
Sigma = 0.48 ns