## Slide 1: Sensor Anomalies Overview
Claire Juramy
presenting work from the Sensor Anomalies Working Group
DESC – August 2022

## Slide 2: Sensor anomalies
Goal: from a numerical value per pixel, to how many photons were captured in each pixel
  Source for each Instrument Signature Removal (ISR) step
  Impact on processing
  Some anomalies have been fixed during optimization phase, need metrics
Latest developments in the Working Group (and in the Camera Verification Testing meeting) – but not an exhaustive list
Full focal plane, large set of data: start seeing anomalies that appear on small fractions of sensors

## Slide 3: Sensor and electronics effects
Intrinsic sensor effects:
Defects: bright, dark
Quantum Efficiency, fringes
Static effects: tree rings, edges, midline
Dynamic effect: brighter-fatter
Sensor management effects:
Clocking scheme: bias stability, bias shape
Charge Transfer Efficiency: traps, pinch
Full well
Glows
CCD output stage: gain stability, noise
Readout speed
Clock-induced: tearing
Persistence
Readout electronics:
Gain stability
Linearity
Bias shift
Crosstalk, memory
Acquisition :
ADU favoritism
DAQ scrambling

## Slide 4: Electronic readout chain
ITL CCDs: second output stage on flex cables
Dedicated readout integrated circuit: ASPIC, 8 channels/chip
Two stages: gain (1.6) and Dual-Slope Integrator (RC, integration time)
External buffer amplifier + 18-bit ADC

## Slide 5: Analog-to-Digital Converters
ADC distribution study: Adrian Shestakov CVT page
ADC favoritism: Zoe Smith CVT slides (DNL, testing conversion times)

## Slide 6: Analog-to-Digital Converters
Signal-dependent noise from ADC: Andy Rasmussen CVT slides
Narrow ADC codes model: Pierre Astier CVT slides – does not expect that DNL correction would be required for science products

## Slide 7: Electronic gain variations
Fine reconstruction of monitoring values
Current consumption increased in ASPIC -> temperature increased -> gain changed (RC increased)
Need to clamp (reset capacitors on first stage of ASPIC) regularly

## Slide 8: Linearity
Correction in EOtest: Jim Chiang CVT slides
Evaluating the DM linearizer and proposing a new method: Craig Lage slides
Known non-linearity structure from ASPIC: need fine sampling mid-range

## Slide 9: Crosstalk non-linearity
Crosstalk data on full focal plane: combined analysis by Andrew Bradshaw and Adam Snyder

## Slide 10: Crosstalk and delayed signal
After a bright pixel: delayed signal in the next pixels in all channels of the sensor (Bruno Longo slides from bright columns)

## Slide 11: Bias shift
Seen from bright defects
Very long relaxing time
Threshold ~ 190 kADU (raw)
Adjustement to ASPIC gain + RC + integration time
Detected again in run 5: Duncan Wood slides

## Slide 12: Clock injection in ITL rafts
Individual clock injections: S1 > S2 > S3
Affects bias level: combination of clock injections
Injects noise

## Slide 13
Bias 2D structure: all sensors
Bias drift: specific to some e2v sensors: time-dependent structure
Stability: dependent on how the sensors are run between frames
Bias stability and bias drift

## Slide 14
Bias correction: many contributions, no single solution:
  PCA correction (Thibault Guillemin CVT slides )
  Using the serial overscan
  Pierre Antilogus ITL and e2v studies
Bias structure

## Slide 15: Parallel transfer efficiency
Parallel clocks in e2v: 4 phases, crossing
Delay rise / fall, propagation of clock wavefront over sensor area
Traps: pocket pumping?
No overlap in parallel clocks
1 µs overlaps

## Slide 16: Serial charge transfer efficiency
Traps: fixed amount, present at low flux
Threshold: ‘pinching’ effect
Voltage range + timing
P. Astier

## Slide 17: Full well and Photon Transfer Curve
Full well definitions:
  Intuitive
  Loss of meaning: PTC turn-off
  ‘PSF full well’ (Aaron CVT recap slides)
Dip in the PTC: ‘pinching’ effect in parallel transfer, voltage range + timing (see also Andy Rasmussen slides)

## Slide 18
Slower readout in latest focal plane run (Aaron Roodman slides)
Readout speed

## Slide 19
Skyline pattern: observed at high deltaP (> 9.5V), builds up from middle of sensor
Thin tearing: excess of charges comes before deficit in readout: not an issue with parallel transfer
Divisadero
19
Original tearing patterns

## Slide 20
Bipolar mode: variety of structures depending on working conditions and sensor
Without inversion (< -7V), same conditions as unipolar mode
Metrics: at segment edges
Testing on full focal plane

## Slide 21: Persistence
Andy Rasmussen CVT slides

## Slide 22: Quantum efficiency
Scan over wavelength range (not absolutely calibrated)
Depth of photon conversion (from the back)

## Slide 23: Fringing
Model and fit, for various sensors + conditions: Zhiyuan Guo slides
Amplitude depends on focal ratio

## Slide 24: Tree rings
Several studies, with 9-raft data: Johny Esteves slides

## Slide 25: Brighter-fatter
Fit from PTC data for whole focal plane (Pierre Astier CVT slides)

## Slide 26: Brighter-fatter
Statistics on coefficients: serial direction

## Slide 27: Brighter-fatter
Statistics on coefficients: parallel direction

## Slide 28: Sensor anomalies: perspectives
Progress on many of the top concerns
More precision, more effects
Upcoming: last chance for testing run at SLAC, Narrow Beam CCOB
Raft testing available (TS8)
AuxTel available
ComCam: to be continued

## Slide 30: Pixel readout: time scan

## Slide 31: Sequencer clocks and readout
Raw OS
Integrated
ASPIC signals
ADC trigger
CCD serial clocks