## Slide 1
ML Gallin-Martel*, A. André, Y. Arnoud, A. Bes, JL Bouly, J. Collot, D. Dauvergne, R. Delorme,
F. Di Franco, P. Everaere, L. Gallin-Martel, O. Guillaudin, C. Hoarau, F. Lafont, A. Lacoste, E. Lagorio, C. Léonhart, S. Marcatili, M. Marton, R. Molle, JF Muraz, N. Ponchant, F. Rarbi, M. Reynaud, O. Rossetto, J. Waquet, M. Yamouni
M. Evin, A. Guertin, F. Haddad, C. Koumeir,
V. Métivier, R. Molle, Q. Mouchard, F. Poirier, N. Servagent
P. Barberet, J. Jouve, S. Sorieul
Ph. Laniece, AM Leite, Y. Zhu
JF Adam, R. Serduc, S. Keshmiri
F. Vianna Legros
L. Abbassi, T. Crozes, F. Donatini, E. Gheeraert, J. F. Motte, J. Pernot
J. Bousquet, E. Corne, J. Letellier, D. Nusimovici.
F. Lafont, TN. Tran – Thi
J. Herault, JP. Hofverberg, D. Maneval, R. Trimaud
*mlgallin@lpsc.in2p3.fr
Design of innovative diamond detectors for beam monitoring in highly radiative environment for applications in nuclear and medical physics

## Slide 2
Development of new generations of ion accelerators:
  Nuclear applications: the production of radical species for radiolysis applications
  Medical applications: hadrontherapy, flash therapies and X-ray or synchrotron radiation therapy
  very precise monitoring of the beam with rapid counting in a highly radiative environment.
The intrinsic qualities of diamond:
  speed, low leakage current, excellent SNR, resistance to radiation
  ⇒ an excellent candidate to meet such monitoring requirements over a wide dynamic range from a fraction of pA (single particle) up to µA.
Context

## Slide 3
Solid state ionization chamber
Read-out electronics
Fast current preamplifier
Charge preamplifier
QDC
Diamond as beam monitors
Band Width: 2 GHz
Gain: 40 dB
Impedance: 50 Ω
Dynamic range: ~ +/- 1 V
Power Supply: 12 V / 100 mA
HT
Diamond Assets :
Intrinsic radiation hardness
Fast signal risetime enables timing precision of a few tens of ps
Low noise
Issues :
Cost
Availability of large area
Diamond metallisation
Detector assembly

## Slide 4
Solid state ionization chamber
Read-out electronics
Diamond as beam monitors
HT
Diamond Assets :
Intrinsic radiation hardness
Fast signal risetime enables timing precision of a few tens of ps
Low noise
Issues :
Cost
Availability of large area
Diamond metallisation
Detector assembly
Read out electronics @ LPSC
Discret
C. Hoarau et al 2021 JINST 16 T04005
https://doi.org/10.1088/1748-0221/16/04/T04005
Fast current preamplifier
Charge preamplifier
QDC

## Slide 5
Solid state ionization chamber
Read-out electronics
Diamond as beam monitors
HT
Diamond Assets :
Intrinsic radiation hardness
Fast signal risetime enables timing precision of a few tens of ps
Low noise
Issues :
Cost
Availability of large area
Diamond metallisation
Detector assembly
Read out electronics @ LPSC
Discret
C. Hoarau et al 2021 JINST 16 T04005
https://doi.org/10.1088/1748-0221/16/04/T04005
Fast current preamplifier
Charge preamplifier
QDC

## Slide 6
@ ESRF (France) : Photons 8.5 keV=> 2D « current maps » with CVD diamonds
https://doi.org/10.1016/j.diamond.2020.108236
XBIC = X rays Beam Induced Current
Diamond detector for beam monitoring: sCVD? pCVD? DOI?
N. Vaissiere PhD thesis https://tel.archives-ouvertes.fr/tel-01022652

## Slide 7
Beam monitors for on-line control of radiotherapies in cancer treatment
Flash therapy vs conventionnal => high doses in a very short time
Short pulses to be monitored at high beam intensity !
High particle couting rate capabilities to be demonstrated
Bunch or train of bunches time stamps
Conventional X-rays
or hadrontherapy  Gy/mn
Flash therapies  100 Gy/s
Vectorized Internal Radiotherapy and Boron Neutron Capture Therapy (BNCT)
X-rays vs hadrontherapy
Radionuclides α : 223Ra, 225Ac, 212/213Bi, 211At…
Energy α : 5 – 9 MeV
Objectives: improved dose predictions and biological effects for targeted radiotherapy
monitor low-energy ion beams for radiobiology experiments

## Slide 8
Pulsed proton beam monitoring – application to Flash Therapy (ARRONAX)
R&T DIAMTECH / ANR – DIAMMONI Coll. LPSC SUBATECH ARRONAX (R. Molle PhD thesis)
Current integration (dynamic=105)
Train Counting => Time stamps ~3 ns
Bunch Counting
Micro-Xrays beam monitoring application to Microbeam Radiation Therapy (ESRF)
R&T DIAMTECH / IDSYNCHRO / PAIR TUMC - Coll. LPSC – STROBE (INSERM) (N. Rosuel PhD thesis)
Current integration (dynamic=106)
Integration from 10 ms up to 100 ms
Micro - ion beam monitoring (LP2I Bordeaux/AIFIRA - IRSN/ MIRCOM)
Coll. LPSC Institut Néel-Grenoble LP2I-Bordeaux IRSN (C. Léonhart PhD thesis)
To be crossed by protons with energy up to 4 MeV, alpha particles up 6 MeV and B, C, O, … ions up to 8 - 10 MeV.
Diamond deep etching => diamond menbrane of ~1 µm (Institut Néel)
Integration on the µ beam line = extraction window
DIAMANT – Beam monitoring @IN2P3 and for medical application : various specification lists!
138 QDC channels : 20 ASIC + 8 Diamonds in 1 row
‘
Monitoring low energy ion beam (ALTO)
Coll. LPSC IP2I Lyon IJC lab via PICTURE (MP Diamant & BioALTO)
Counter with 4 diamonds (2x2x0.1mm3) placed in the halo of the beam for irradiation monitoring and beam alignment.
DIAMMONI
DéFI DiaMs
QDC : Train counting
1 sCVD / 1 pCVD diamond
Fast preamp + 500 MHz ADC : Bunch counting
Proton beam monitoring in hadrontherapy (CAL)
Coll. Clarys UFT / Coll. TIARA / Coll. ClaRyS - LaBeX PRIMES (P. Everaere PhD thesis)
Single particle
XY spatial resolution ~1 mm
Time resolution ~100 ps
40 preamps + DFC + TDC 4 diamond sCVD in mosaic or 1 pCVD

## Slide 9
TDC = Time Digital Converter
FPGA
2020
LPSC Grenoble
CMOS 130 nm
12,5 ps RMS
Beam Monitor
CFD Board
ACQ board
+TDCZZZZZZZZZZZ
10 cm
Preamplifiers:
Diamond beam hodoscope in single proton regime (hadrontherapy monitoring): the 100 ps time resolution challenge
Diamond
25 ps RMS
42 channels
CFD=> time stamp independant of diamond signal amplitude
CFD = Constant Fraction Descriminator
S. Curtoni PhD thesis http://www.theses.fr/2020GRALY045 UGA / CLARYS-UFT INSERM
Time resolution between two orthogonal X and Y strips
P. Everaere PhD thesis, Labex PRIMES
May 2023
C. Hoarau et al 2021 JINST 16 T04005
https://doi.org/10.1088/1748-0221/16/04/T04005

## Slide 10
1st veterinary patients
1rst version: 32 channels (QDC) in discrete elec.
2nd Version: 8 channels with ASIC QDC
Final version : 158 channels 20 QDC ASIC 8 diamonds in a row
N. Rosuel PhD LPSC STROBE, Labex PRIMES
http://www.theses.fr/s211637#
Metallization NanoFab, Institut Néel
DIAMTECH WP1 : imageur portal – avancées techniques
Diamond for X-rays µ-beams: charge integration + linear response over a wide dynamic « Flash therapies »
138 channel response to 1 single µbeam swapping the detector
zoom
  Micro Beam Radiation Therapy (MRT)
Innovative radiotherapies using spatially segmented photon beams
Energy 50-200 keV@ ESRF + very high dose rate 104 Gy/s
R&T DIAMTECH/ IDSYNCHRO UGA / PAIR TUMC MRT-CLINTRA INSERM

## Slide 11
Time ( µs)
A.U
Diamond beam monitor for high proton beam intensities : train couting and time stamps towards « proton FLASH therapies »
I = 10nA, dt = 10µs, dit = 1s
Tension(en V)
Temps (µs)
Time ( µs)
Train integral = DSO post analysis
Train integral = on line measurement with LPSC QDC
Train - Diamond signal
Train Counting Mode
Beam intensity up to 300 nA measured on diamond
DSO Lecroy
2 GHz; 10 or 20 GS/s
68 MeV Proton beam
QDC board
developped @ LPSC
DAQ
developped @ LPSC
Volts-seconds
Time ( µs)
Amplitude (V)
FE elec.
+ DAQ
validation
ANR - DIAMMONI

## Slide 12
Charge in a 50 µs-train (µC)
Study of linearity as a function of beam intensity: sCVD versus pCVD
R. Molle, PhD Thesis LPSC ARRONAX SUBATECH, ANR DIAMMONI
sCVD: Linear response from 1 Gy/s up to 102 kGy/s
pCVD: Linear response from 1 Gy/s up to 4 102 kGy/s
for monitoring proton beams in FLASH conditions
with typical dose rate : ~ 300 Gy/s
Diamond beam monitor : the challenge of a linear response on a wide dynamic with ion beams
Each point = ACQ on a single train to avoid any influence of the fluence
~300 Gy/s
sCVD / pCVD
Diamond

## Slide 13
High proton flux : diamond signal analysis and simulation towards a better understanding of charge collection mechanism
PyDiam simulation package in Python 3 (R. Molle PhD thesis, ANR DIAMMONI)
sCVD and pCVD diamond signals for different beam intensities in ARRONAX
Y. Arnoud simulation of a gas ionization chamber
Time (µs)
Time (µs)
Amplitude (au)
Amplitude (au)
Beam intensity
3,8 µA
0,03 µA
Beam intensity
3,8 µA
0,03 µA
sCVD
pCVD
DIAMMONI needs to be
Thin: 50 µm
Biased > 1V/µm (= « normal » bias): 200 V = 4 V/µm
500 µm diamond thick
X µA = beam intensity measured on diamond

## Slide 14
Targeting objective
Objectives of the “aging” experimental set-up:
each area corresponds to a beam intensity (= flux = 𝜑)
Irradiation of each zone until obtaining an identical cumulative fluence (=𝜙) between the zones at the end of the experiment
XBIC mapping at ESRF
Before
irradiation
sCVD 550µm E6
High fluences : what’s the limit for diamond continuous exposure in a ion beam?
ANR - DIAMMONI

## Slide 15
This diamond was irradiated with 68-MeV alpha particles at ARRONAX with a fluence of approximately 1015 ions/cm2
The absolute values cannot be compared since the beam intensity is not the same
The current scales are in relation to the X-beam intensities in the two experiments (before vs after irradiation)
=> The main observation is that the XBIC is close to the leakage current in the irradiated areas which means that the charges are no longer collected.
To understand if the origin if this degradation, Bragg diffraction imaging was performed.
No irradiation-related structural defects are visible => the vacancies created by the irradiation are too far from one another to induce a significant deformation of the crystalline mesh.
From simple calculations, we estimated the displacements per atom to be of the order of few p.p.m
XBIC mapping at ESRF
Before irradiation
XBIC mapping at ESRF
After irradiation
In blue the areas of less effective charge collection
High fluences : diamond aging studies with 68 MeV alpha particles
Binocular observation
Black traces of irradiation on the PCB
XBIC mapping at ESRF
After annealing
(1000 °C for 6 hours)
Diamond has recovered
The diamond is colored in green typical of GR1 type defects
sCVD 500µm E6
F. Laffont, TN , Tran – Thi XBIC setup BM05 ESRF
ANR - DIAMMONI
Photoluminescence spectra

## Slide 16
Fluence (1015 alpha/cm2)
Charge collected per alpha (au)
Normalized charge collection per alpha particle
Impact of flux on Charge Collection Efficiency (CCE)
unusable
With increasing flux, charge collection decreases.
The fluence creates damage in the detector.
It results in the creation of traps that disrupt the collection of charges.
CCE drops faster with fluence at higher flux.
1013 alphas/cm2/s = 0.1 alpha/nm2/s: interplay between vacancy creation and stable electronic defects (charge carrier traps)?
A
C
D
E
High fluences : diamond aging studies with 68 MeV alpha particles
68-MeV alpha particles
dt = 50µs
dit : variable => <I> = cste
Duration : ~30 mn for each A to E point
ACQ : every ~3 mn between two consecutive points
R. Molle PhD thesis ANR DIAMMONI

## Slide 17
70-MeV protons
**L. Bäni et al, RD42 coll. , « A study of the radiation tolerance of CVD diamond to 70 MeV protons, fast neutrons and 300 MeV pions », Sensors 2020, 20, 6648
sCVD and pCVD diamond response to 70 MeV protons: aging studies
PRELIMINARY
= fluence
, = Cste = fit parameters
RD42 CERN coll. results**:
k is decreasing with proton energy
Question to be answered:
k = k (𝜑) with 𝜑 = flux?
Ongoing studies!
Proton therapy:
~1010 protons/cm2/treatment
~20 patient/day
~1013 proton/cm2/year
Literature * and **:
* A. Bhattacharya et al, « Degradation of single crystal diamond detectors in swift heavy ion beams », Diamond and related Materials 70 (2016) 124-131
sCVD results
dt = 50µs
dit : variable => <I> = cste
R. Molle PhD thesis ANR DIAMMONI
Fluence (p/cm2)
Fluence (1015 p/cm2)
Charge in a 50µs-train (µC)
Charge in a 50µs-train (C)
Charge in a 50µs-train of 70 MeV protons as a function of fluence
Flux in p/cm2/s
log. scales on the both axes

## Slide 18
Conclusion
The development are in connection with collaborations established at IN2P3, CNRS (INP), CAL, IRSN and ESRF
It has been done in the context of interdisciplinary research IN2P3 - INP skills exchanges take place between
  characterization: sources (labs) + accelerator beams @ IN2P3 AIFIRA / GIP - ARRONAX, XBIC @ ESRF
  Instrumentation (IN2P3 labs, Institut Néel, etc.)
Diamond beam monitors + FE electronics and DAQ have been developed
  Various specifications (time resolution, continuous or pulsed beams, single particle counting up to 106 particle in a bunch)
  Diamond beam monitoring satisfies the FLASH therapy condition (linear response in the range 1 to 4 105 Gy/s with X-rays and proton beams)
  => the proposed detection systems will bring significant added value to the transfer of high dose rate flash radiotherapy to clinical trials
Simulation + aging studies are on going for a better understanding of the signal formation and the impact of very high fluences on charge collection

## Slide 19
THANKS FOR YOUR ATTENTION !
Diamond Jokes Quotes
©