## Slide 1: The performance of Visible Telescope --- in the view of scientists
Yulei QIU Liping XIN
NAOC

## Slide 2: Outline
General specifications and functions
Detection capability
Two major modifications for VT FM
Calibrations
Data Processing Pipelines

## Slide 3: General Specifications of VT
Diameter : 44cm ; Focal Length: 360cm
FOV: 26‘x26’ Sensitivity : 22.5 Mv in 300s
Encircled energy : 80% and 70% within the radius of 0.77aresec in blue and red channels, respectively.
Two channels: 400 – 650 nm(blue); 650 – 1000 nm(red)
CCDs for two channels :
    2kx2k, 13.5 μm for both CCDs
    Cooling temperatures: -65C （blue） and -75C （red ）
    blue : thinned CCD, QE ~90% at 600nm cooling
    Red : deep depletion CCD, QE > 50% at 900nm)
    Readout noise for both CCDs:
    < 8e-/pix （for 15s exp.) ; <6e-/pix （for >=100s exp.）

## Slide 4: General functions of VT
Supports all of the SVOM tasks (CP ,GP , ToO)
Calibrations : photometry, astrometry and alignments
Data processing : onboard and on-ground
Data link (VHF, Beidou, X-band, S-band)
Data budgets : maximum of 30Gbit per day.

## Slide 5: Detection capability
SNR=
R* : derived from throughput of the telescope:
    CCDs : calibration of FM
    Optics: calirbration of FM
    whole system: direct measurements and indirect measurements
    Rsky : derived from brightness of sky background the throughput
    Npix : dirived from Encircled energy performace(measurements)
    RN and Dark : noise of the detector (measurements)
Estimated detection capability : For a GRB with magnitude of 22.5 and spectrum of index of -1 for a 3*100s exposure, SNR:
    Blue band: 4.56
    Red band: 6.20
The performance of the detection capability meets the requirement in MRR (>3sigma)

## Slide 6: New thermo design one of two major modification of FM
Old thermo design
new thermo design
E
E
(pix)

## Slide 7: New design of VT to avoid ghost images: one of two major modifications of FM
  New baffle design can effectively remove the ghost-images and have no significant impact on the detection capability.

## Slide 8: Calibration :throughput
  Before launch:
  Two different method
  : Directly
  CCD and optics were measured separately
In orbit :
validation and zero point Correction by observing standard stars

## Slide 9
Calibration: SVOM AB magnitudes to Johnson System

## Slide 10: Photometric accuracy
Short term calibration
  Psf correction
  psf is measured for each frame.
  Longterm calibration
  Before launch : Throughput test
  In orbit: validation and re-calibration with standard stars
  Monitor GAIN of detectors
    Internal calibration (LED lamps)
Systematic photometric calibration accuracy can be
better than 1%--meets the requirements in MRR

## Slide 11
Astrometric accuracy
Simulations show that
Systematic calibration accuracy can be
better than 0.1 pixel

## Slide 12: Alignment calibrations
  VT and Star-trackers
  VT and Eclairs
  VT and MXT

## Slide 13: Data Processing Pipelines
VOPP (VT Onbarod Processing Pipeline)
  NAOC and SCEM teams are testing software
  VT Beidou packets will be optimized
VVPP(VT VHF data Processing Pipeline)
  CALDB: astrometric and photometric calibrations
  Participate in the NMS test in April 2023
  Qsourcelist_VT, Qsky_VT, 1 bit images with WCS
VTAC(VT VHF Afterglow Candidate)
  GRB candidate search by crossmatch of VT source lists and archived data
Science products:
  Qsourcelist_VT
  Qsky_VT
  QCANDI_VT
  QIM1B
  VXPP(VT X-band data Processing Pipeline)
    CP:
    CP01: in the case that the optical counterpart has not been identified.
      Optimize the strategy of image-combination
      Find variable objects in error box of MXT
    CP02: photometry for identified GRBs.
    ToO-MM
    find transients in the nearby galaxies with the algorithm of image-subtraction .
  Science products : CANDI_VT
    Candidates list or counterpart
    Positions
    Light curves,
    Finding charts

## Slide 14: VXPP-p01: photometry for optical counterparts of GRB
For faint GRBs

## Slide 15: VXPP-p02: automatically search for candidates of optical afterglows in VT images

## Slide 16
Merci beaucoup!
谢谢！