## Slide 1
SHI Haoli,HUANG Yue,HE Jiang
2023.09.18
GRM performance
Requirements & Validation

## Slide 2
GRD (Gamma Ray Detector)：3GRDs (NaI + PMT)，used for GRB trigger, spectrum measurement and rough localization.
GCD (GRM Calibration Detector)：241Am+PS+SiPM，used for on-orbit calibration.
GPM (GRM Particle Monitor)：PS+PMT，monitor the electrons and judge the SAR.
GEB (GRM Electronic Box)：to generate GRM packets.
GRD(GCD)
GPM
GEB
GRM introduction

## Slide 3
FM calibration experiment introduction
items | GRD used
Energy response, detection efficiency, angular response calibrations and HV effect with radioactive sources | GRD01-05
Energy response, detection efficiency calibrations with x-ray beam (10-160keV) | GRD01-05
GPM threshold and x-ray efficiency calibrations (electron: 1MeV, proton: 20MeV, x-ray beam for efficiency ranging from 10keV-130keV) | GPM01-02
GRD threshold and HV research (threshold:0x21, 15keV detectable, 0x0d to check the lowest detectable energy) (HV: E-C, resolution, ratio of high gain and low gain) | GRD01-03
Temperature response (5-35℃, 241Am, 57Co, 133Ba…, full-energy peak, energy resolution) | GRD05
Dead time test | GRD01-05
Performance variation test for a long term (Energy response, detection efficiency, energy resolution) | GRD01-05，GCD01-05，GPM01-02
GRD01-05:
GRD04 & 05 as spares
GPM01-02:
GPM02 as a spare

## Slide 4
Equipment assembly
for X-ray beam
for sources

## Slide 5
No. | Source type | energy(keV) | 活度(Bq)
1 | 137Cs | 661.657 | 1.993E+06
2 |  | 32.194 | 
3 | 113Sn | 391.698 | 2.457E+06
4 | 60Co | 1173.228 | 3.504E+06
5 |  | 1332.492 | 
6 | 88Y | 14.958 | 1.518E+06
7 |  | 898.042 | 
8 |  | 1836.063 | 
9 | 57Co | 14.41 | /
10 |  | 122.06 | 
11 | 22Na | 511.0 | /
12 |  | 1274.537 | 
13 | 241Am | 59.5 | /
Energy(keV)
(8、9)10
11-12
13-15
16-20、22、24
26、28、30-31
32-40
41-44
45-51
55/65/75/80/
110、120、130、140、150、160
Non-linearity of NaI crystal in low energy
The absorption edge of Iodine element on 33keV
We do the test with more energy points under 20keV and near 33keV
Energy test points
Sources list
X-ray beam setting and sources selection

## Slide 6
Item | Specification | Verification method
Energy range | 15 - 5000 keV | test & analysis
Field of View | ± 60° (for one GRD) | Simulation
Sensitive area | > 200 cm2 (for one GRD) | test
Dead time | < 8 μs | test
Temporal resolution | < 20 μs | review of Design
Energy resolution | ≤ 19%@60 keV | test
Burst observation rate | >90/year | simulation & analysis
GRB localization error | <5 degrees (fluence > 1e-6 ergcm-2@1-1000 keV, 1s) | simulation & analysis
Performance requirements

## Slide 7
E-C relation calibrations
E-C relation of GRD with hard X-ray beam and sources
E-C relation of GRD with radioactive sources
Energy range

## Slide 8
8kev spectrum measurement
with x-ray calibration (threshold:0x0d)
Energy range
channel | detector |  | 
 | GRD1 | GRD2 | GRD3
3200 | 6109keV | 5861keV | 5763keV
Energy range calculation of low gain
According to the EC for low gain, one can get:

## Slide 9
The FOV of GRD is defined as the solid angle with the incident polar angle where the effective area of GRD drops half.
Field of View
simulation (angular response)
Theta=60°: effective area is more than half of the sensitive area.
Meet the requirement of FOV (±60°)

## Slide 10
113Sn的24keV全能峰相对效率变化
113Sn的392keV全能峰相对效率变化
Field of View

## Slide 11
Sensitive area of GRD
The diameter of NaI crystal is 16 cm, thus the sensitive area of each GRD is 201.06 cm2
Dead time
The test result on GRD FM is about 4us.
Sensitive area and dead time
60Co time interval spectrum

## Slide 12
Temporal resolution
The temporal resolution is the accuracy of arrival time of each event, which is designed to be 1μs. Crystal frequency 30MHz, for events, its time precision is 1us in the raw data.
Sensitive area and dead time
Byte No. | contents |  |  |  |  |  |  | 
 | D7 | D6 | D5 | D4 | D3 | D2 | D1 | D0
byte 1 | Packet type (‘1’) | CH(1 to 0): “00”invalid,“01”GRD1,“10”GRD2, “11”GRD3 |  | ‘0’single gain event | Coincide flag： 0 non-coincide event 1 coincide event | Gain falg： 0 low gain，1 high gain | Packet sequence number | 
 |  |  |  | ‘1’double gain event |  |  |  | 
byte 2 | ptime(19 to 12)：microsecond counts(19 to 12) |  |  |  |  |  |  | 
byte 3 | ptime(11 to 4) ：microsecond counts(11 to 4) |  |  |  |  |  |  | 
byte 4 | ptime(3 to 0) & energy(11to 8)：microsecond counts(3 to 0) &channel(11to 8) |  |  |  |  |  |  | 
byte 5 | energy(7 to 0)：channel(7to 0) |  |  |  |  |  |  | 
GRM x-band event packet definition

## Slide 13
energy resolution
Energy resolution of GRD FM.
Measured with 241AM.
 | GRD1 | GRD2 | GRD3
Energy Resolution @60keV | 15.85% | 16.2% | 16.5%

## Slide 14
Fig.1 B1 law operational strategy
Fig.2 The exposure time in one year(FOV=).
We assume that there are ~550 gamma-ray bursts with peak flux greater than 0.3 ph/cm2/s within one year. Therefore, within three years of operation, there will be approximately 1650 gamma-ray bursts. These bursts will be uniformly sampled in time and space. The energy spectrum and peak flux of GRBs are derived from the results of BATSE. if at least two detectors are required for triggering(>5), then ~330 GRBs can be triggered within three years, with an average of approximately ~110 triggers per year.
Burst observation rate

## Slide 15
Incident direction(30,120)
68% confidence interval（1）：
68% confidence interval（1）：
Local error ~5°
Local error ~10°
localization error

## Slide 16
All the performances of GRM meet the requirements, and validated by tests, simulation, calculation or analysis for GRM FM.
Merci!
Thank you!
conclusion