## Slide 1: DIAMMONI: diamond-detector for beam monitoring at ultra-high dose rate
Molle Robin
molle@lpsc.in2p3.fr

## Slide 2: Why beam monitoring?
Ion beam
50 MeV – 150 MeV for protons
Beam monitoring
Time tagging for particles
Secondary particle emissions
Source : Experimental determination of absorbed dose to water in a scanned proton beam using a water calorimeter and an ionization chamber, S. Gagnebin,2010. URL : https://www.researchgate.net/publication/272623275_Experimental_determination_of_absorbed_dose_to_water_in_a_scanned_proton_beam_using_a_water_calorimeter_and_an_ionization_chamber
Source: The pulsing chopper-based system of the ARRONAX C70XP Cyclotron, Poirier et al., International Particle Accelerator Conference, 2019. URL https://doi.org/10.18429/JACoW-IPAC2019-TUPTS008
Time stamps for trains
Number of particles
Beam structure:
Microstructure : ~30 ns bunch
(4 ns ON – 29 ns OFF for ARRONAX cyclotron)
Macrostructure : few µs to 1 ms train
train made of bunches
tunable duration dt
tunable duration between trains dit
Diamond detector

## Slide 3
d = 150-500 µm
Ionizing particle
Diamond compared to other semiconductors
Diamond Assets :
Intrinsic radiation hardness
Fast signal risetime enables timing precision of a few tens of ps
Low noise
Issues :
Cost
Availability of large area
Flash therapy vs conventional?
 high doses in a very short time: diamond beam monitoring!
Short pulses to be monitored at high beam intensity !
High particle couting rate capabilities to be demonstrated
Bunch or train of bunches time stamps
Conventional X-rays
or hadrontherapy  Gy/mn
Flash therapies  100 Gy/s

## Slide 4
68 MeV proton / 70 MeV alpha particles cyclotron
Beam intensity up to 20 µA
Continuous or pulse mode beam delivery with tunable dt and dit time duration
Bunch : 4ns ON, 29ns OFF (f = 30,45 MHz)
Source: The pulsing chopper-based system of the ARRONAX C70XP Cyclotron, Poirier et al., International Particle Accelerator Conference, 2019. URL https://doi.org/10.18429/JACoW-IPAC2019-TUPTS008
ARRONAX cyclotron (France)

## Slide 5
Experimental set-up
Voltage bias
Oscilloscope
Diamants
Collimateur 3 mm
Collimateur 1 mm

## Slide 6
Diamond response with proton flux
For each beam intensity:
1 train of 50 µs of 68 MeV protons
Saturation effect for higher fluxes :
at high flux, an amount of e-h pairs created
can disturb the electric field applied
lifetime of e-h in diamond longer
 more risk of trapping and recombination
 saturation effect
Linearity of sCVD diamond until a flux of ~2.5E13 p/cm²
Flash : >100 Gy/s
for beam monitoring in FLASH conditions
What about lifetime of the detector?
   towards aging studies in extreme conditions in terms of flux and fluence
Simulation work on doing

## Slide 7
Single-crystal CVD diamond chosen for the experiment
Irradiated diamond by 68-MeV alpha particles at ARRONAX. Total fluence: 1015 ions/cm2
Bragg diffraction imaging was performed
No irradiation-related structural defects are visible
 Vacancy density is to low to have an impact on the crystal mesh
From simple calculations, we estimated the displacements per atom to be of the order of few p.p.m
XBIC mapping at ESRF
Before irradiation
XBIC mapping at ESRF
After irradiation
In blue the areas of less effective charge collection
High fluences : diamond aging studies with 68 MeV alpha particles
XBIC mapping at ESRF
After annealing
(1000 °C for 6 hours)
Diamond has recovered
The diamond is colored in green typical of GR1 type defects
sCVD 500µm E6
F. Laffont, TN , Tran – Thi XBIC setup BM05 ESRF
ANR - DIAMMONI
Fluence experiment
Homogeneous response of the active surface before irradiation

## Slide 8
Aging studies (flux and fluence)
70-MeV protons
**L. Bäni et al, RD42 coll. , « A study of the radiation tolerance of CVD diamond to 70 MeV protons, fast neutrons and 300 MeV pions », Sensors 2020, 20, 6648
PRELIMINARY
= fluence
, = Cste = fit parameters
RD42 CERN coll. results**:
k is decreasing with proton energy
Question to be answered:
k = k (𝜑) with 𝜑 = flux?
Ongoing studies!
Remind proton therapy:
~1010 protons/cm2/treatment
~20 patient/day
~1013 proton/cm2/year
Literature * and **:
* A. Bhattacharya et al, « Degradation of single crystal diamond detectors in swift heavy ion beams, Diamond and related Materials 70 (2016) 124-131
sCVD results
dt = 50µs
dit : variable => <I> = cste
(lin,lin)
(log,log)

## Slide 9
Time ( µs)
A.U
I = 10nA, dt = 10µs, dit = 1s
Time ( µs)
Train integral = DSO post analysis
Train integral = on line measurement with LPSC QDC
DSO Lecroy
2 GHz; 10 or 20 GS/s
68 MeV Proton beam
QDC board
developped @ LPSC
DAQ
developped @ LPSC
V.s
Tension(en V)
Temps (µs)
Signal of 10µs train of protons
Time ( µs)
Amplitude (V)
DIAMMONI detector development

## Slide 10
Conclusion
DIAMMONI: Diamond detector for pulsed beam monitoring in FLASH conditions
diamond used as solid-state ionizing chamber
  studies function of
    flux (up to very high dose rate beyond FLASH conditions)
    fluence (lifetime of the detector under extreme radiation conditions)
front-end and back-end electronic developments (current integration
in proton train)  instrumentation of diamond as beam monitor
Further studies function of
ionizing particle type
signal formation (simulation)
Ultimate goal : integration on the beam line (summer 2024)

## Slide 11: Thank you for your attention !
R. Molle*, Y. Arnoud, A. Bes,
JL Bouly, J. Collot, D. Dauvergne,
P. Everaere, L. Gallin-Martel, ML Gallin-Martel,
O. Guillaudin, C. Hoarau, F. Lafont, A. Lacoste,
E. Lagorio, S. Marcatili, M. Marton, JF Muraz, N. Ponchant, F. Rarbi, M. Reynaud, O. Rossetto, J. Waquet
A. Guertin, F. Haddad, C. Koumeir,
V. Métivier, R. Molle*, Q. Mouchard, F. Poirier, N. Servagent
F. Lafont, TN. Tran – Thi
* molle@lpsc.in2p3.fr

## Slide 12: BACK-UP slides

## Slide 13
@ ESRF (France) : Photons 8.5 keV
2D « current maps » with CVD diamonds
XBIC = X rays Beam Induced Current
Diamond detector for beam monitoring:
quality of the diamond?
N. Vaissiere PhD thesis https://tel.archives-ouvertes.fr/tel-01022652