## Slide 1
LOREM IPSUM DOLOR
SIT AMET
Beam monitoring with silicon detectors
Prof. R. Sacchi – University of Torino and INFN
… follow-up of the workshop held in November 2021
CNAO – IN2P3 collaboration meeting 24 October 2023

## Slide 2
Motivation for Solid State sensors
IONIZATION CHAMBERS
SOLID STATE
Main issues
Counting particles: signal pile-up
→ fast sensors & readout
→ segmentation
→ difficult above 1010 p/cm2s
High dose rates (FLASH)
→ 103 × dose rates
→ plasma effects in silicon
Radiation tolerance
→ manufacturing strategies
→ damage compensation
Increased complexity
Collection times | ~ 100 µs | ~ ns
Sensitivity | ~ 104 protons ~ 103 C ions | single particle
Time resolution | ~ no/poor | < 100 ps
Deviation from linearity @ high dose rates |  | Less recombination @ high dose rates 102 × E field 102 × charge mobility 10-1 × thickness
Not suitable for fast scanning modalities timing applications high dose rates (FLASH) |  | New applications direct counting # particles timing applications high dose rates (FLASH)

## Slide 3
Thin Low Gain Avalanche Detectors (LGADs)
thickness of sensitive volume < 50 µm
internal charge multiplication ~ 10
+
Enhanced signal of very small duration
Time resolution of tens of ps
Strip detectors (strip area ~ 3 mm2, active thickness 45µm)
Detectors for particle counting
Large area (2.7×2,7 cm2)
144 strips
Detectors for timing applications
Smaller size, 11 strips
Si- substrate removed to reduce total thickness to 70 µm
Beam spot
1 cm FWHM
Pads for large ionization rate studies
4 active thicknesses (15/20/30/45 µm)
5 pad sizes (0,125/0,25/0,56/1/2 mm2)
Internal gain | 
yes | no
 | 
 | 
 | 

## Slide 4
INFN MoveIT project (2017-2021)
Proton beam particle counter (ESA ABACUS)
Six ABACUS front-end discriminators -> 3 FPGA boards
2.7×2.7 cm2 active area (144 strips)
Counting rate up to 100 MHz with < 2% pileup inefficiency
For larger rates, inefficiency measurement implemented in FPGA
Beam energy detector
High precision mechanical system
XYZ axes remotely controlled
8 channel FE board, sensor active area 20 mm2
accuracy on ToF measurement < 10 ps
Self-calibration method developed and tested
A. Vignati, et. al., Phys. Med. Biol. 65 (2020) 215030
A. Vignati, et. al., Med. Phys. 50 (2023) 5817-5827
Mohammadian-Behbahani M, et. al., NIM A 1040 (2022) 167195

## Slide 5
INFN SIG project (2022-2025)
R&D towards an advanced Superconducting Ion Gantry
Multi-ion (He → O)
Lightweight (based on 4-5 T SC curved dipole)
Integrated novel Dose Delivery and in-vivo Range Verification Systems for ions
Prompt Gamma Timing (PGT)
measurement of prompt gamma emission time to get insights into the range of ions
Integration with beam monitoring for time synchronization
PGT distributions measured @ CNAO with protons and C-ions (Merlino INFN project)
    Non-optimized acquisition system
    Low efficiency (large deadtimes)
    Sub-clinical beam intensities
Develop new acq. system based on TDC
beam monitoring system
secondary radiation detectors

## Slide 6
Amplitude (mV)
Time (ns)
C-ions signal in thin sensors
11-strips sensor with gain=1
8-channels amplifier board
Acquisition with CAEN DT5742 digitizer
  16+1 channels, 12 bit ADC
  5 GS/s sampling rate
Ebeam = 398.8 MeV/u
Vbias=149V

## Slide 7
C-ions signal amplitude
Peak Amplitude (mV)
Number of peaks
Charge sharing
between neighbouring channels
399 MeV/u
269 MeV/u
166 MeV/u
115 MeV/u
Electronic noise
& Cross talk
Vbias=149V
Peak amplitude (mV)
Released charge (fC)
0
20
40
60
80
100
MPV from Landau fit

## Slide 8
C-ions time resolution
115 MeV/u
399 MeV/u
2 identical sensors 30 cm apart
8 + 8 channels acquired by CAEN 5 Gs/s digitizer
CFD algorithm to correct for time walk
Coincidences
Histogram of arrival time difference
Coincidences
Combinatorial background
σt = σ()/
Single Hit time resolution (ps)

## Slide 9
100%
50%
Beam Intensitiy 20%
C-ion beam counter
Beam Profile
3% difference between low-high energy
Tails of beam profile
3%
1% difference between low-high intensity
Pile-up inefficiency reduced to 0.5% after corrections
1%
Spots of 5*107 C ions
Comparison of beam FWHM with GAFCHROMIC films

## Slide 10
Time measurements with CERN picoTDC
3ps or 12ps binning
very low jitter (<1ps)
High rate capability
Readout through FPGA
CERN picoTDC evaluation board (64 input channels)
Successfully integrated with 1 channel of ESA-ABACUS board
Conversion efficiency 100%
Tested up to 150 MHz freq.
15 ps time resolution
Qin = 48 fC
f = 1 MHz
1 channel
CFD
PG detector
LaBr3
+
PMT/SIPM
8 chn.
1-4 chn.
Preparation for December integration test @ CNAO

## Slide 11
Monitoring of FLASH UHDR electron beams
13mm solid water slab
FlashDiamond and silicon sensor in same conditions
ElectronFlash accelerator (CFR - Pisa)
9 MeV electrons pulsed beam
Beam current: 1-100 mA
Pulse duration: 4 𝜇s
Pulse frequency: 5 Hz
Uniform fields using 3 cm PMMA plastic applicator
f
45/ 30𝜇m thickness
2/1/0.25 mm2 area
Bias voltage: 10V ÷ 200V
Dose/Pulse 0 ÷ 10Gy
Sensors tested

## Slide 12
Monitoring of FLASH UHDR electron beams
Collected charge/pulse scales with pad area and sensor thickness
Ratios between different area/thickness independent from dose/pulse
Charge collected (nC/pulse)
Measured dose (Gy/pulse)
Measured dose (Gy/pulse)
4 𝜇s pulse duration
4 𝜇s pulse duration
FLASH regime
FLASH regime

## Slide 13
Monitoring of FLASH UHDR: front-end readout
Readout with TERA08 readout ASIC
64-channels front-end used @ CNAO
deadtime free
RC input circuit to prevent from saturation
UNPUBLISHED RESULTS!
45 𝜇m thickness, 2mm2 area pad
200 V bias voltage
Charge collected (nC/pulse)
Measured dose (Gy/pulse)
Detector interface board for TERA09 front-end
144 strip sensor
4 × dynamic range compared to TERA08
Large area sensor (2.7×2.7 cm2) to cover proton pencil beam cross section
TERA09
64 ch.
TERA09
64 ch.
Tests foreseen at TIFPA beginning 2024

## Slide 14
Summary and Outlook
Silicon detectors offer interesting features for new developments in beam monitoring in PT
Integrating counting and timing in the same device seem possible with state-of-the-art TDCs
Covering larger (20×20 cm2) area and reaching clinical proton beam fluxes require pixel segmentation
  quantum leap in system complexity (and costs !!!)
Good linearity with dose per pulse was demonstrated in FLASH e- beams
    Interesting for possible combined Si – IC technology
    Results need to be confirmed with p-beams

## Slide 15
Thank you