## Slide 1: Commissioning of the DESIR HRS
A High Resolution Separator for low energy physics at GANIL
Speaker : Julien Michaud
Authors : J. Michaud, P. Alfaurt, A. Balana, B. Blank, L. Daudin,T. Kurtukian, B. Lachacinski,
F. Méot, L. Serani, F. Varenne

## Slide 2: Contents
The DESIR HRS
Optical aberrations : detection and correction
Spectrometer characterisation

## Slide 3: Beam preparation for DESIR – LP2I Bordeaux contribution
RFQ-Cooler Buncher
GPIB
Double Penning Trap
High Resolution Separator
HRS
Mass separation/beam purification:
Up to isobars level
HRS
Piperade
trap
Beam preparation:
Cooling and bunching
Accumulation trap
GPIB
Mass measurements:
Mass precision:
Piperade
2nd trap
Piperade
2nd trap
Optimal performances can be reached only if a 1st good cleaning is done with the HRS
Piperade
2nd trap
Next talk
By A. Husson

## Slide 4: The HRS
HRS: High Resolution mass Separator
Lattice [1] and elements:
D : Two 90° magnetic dipoles (36° entrance/exit angles)
MQ : Matching quadrupoles
FQ : Focusing quadrupoles
FS : Focusing sextupoles
M : A multipole (up to 5th order)
Configuration: MQ-MQ-FS-FQ-D-M-D-FQ-FS-MQ-MQ
Mirror symetry is imposed to minimize aberrations
Electrostatic
[1] T. Kurtukian Nieto, R. Baartman, B. Blank, T. Chiron, C. Davids, et al.. SPIRAL2/DESIR high resolution mass separator. Nuclear Instruments and Methods in Physics Research Section B: Beam Interactions with Materials and Atoms, Elsevier, 2013, 317, pp.284-289.

## Slide 5: The HRS at LP2I Bordeaux

## Slide 6: HRS synoptics: a compact beamline
ions source
Pepperpot-type
emittance-meter
Electrostatic multipole
: Magnetic
dipole
: Electrostatic
deflector
: Electrostatic
quadrupole
: Electrostatic
sextupole
: Beam profiler
: Faraday cup
: Adjustable slits
: Vaccum valve
DC11-HO
DC11-VE
Q11
Q12
Q13
PR21
F21
CF21
Q21
Q22
DC21-HO
DC21-VE
F22
CF22
PR22
Q23
H21
F31
D31
D32
F32
Q41
H41
DC41-VE
DC41-HO
PR41
F41
CF41
Q43
Q42
F42
CF42
PR42
PR31
PR32

## Slide 7: Contents
The DESIR HRS
Optical aberrations : detection and correction
Spectrometer characterisation

## Slide 8: Different levels of aberrations (simulations)
Hexapolar aberration ( order)
is dominant
Octupolar aberration ( order)
Hexapolar corrected
Higher order aberration ( order)
Octupolar corrected
Typically « C-shaped »
Typically « S-shaped »
Optical aberrations tend to increase the beam size and need to be corrected

## Slide 9: Effect of optical aberrations on beam separation (simulations)
Corrected beam (few aberrations)
Beam position spectra
Front view of the beam
Emittance profile (XX’)
Hexapolar aberration dominant
(HRS-like)

## Slide 10: Measurement of aberrations
23/02/2024
Front view: tantal mask
Side view:
MCP + Phosphore screen
Front view:
CCD camera
Pepperpot
Emittance-meter

## Slide 11: Recent upgrade of the emittance-meter
1D motorized mask

2D motorized mask(!)
New : motorized 2D mask allows for finer and faster sampling.
Allows for « live » emittance measurements.
Cons : fixed size between the holes. The emitttance-meter must be taylored to the measured beam (beam size and divergence must already be known).
10 steps = 10 points measured
10 steps = 400 points measured
Emit vertical : 40 points
Emit horizontal : 400 points

## Slide 12: Correct the aberrations with the multipole
The 48-pole electrostatic multipole

## Slide 13: Hexapolar correction (2nd order)
Not corrected
Overcorrected
Corrected
Emittance
measurements
Projection on x axis gives beam profiles
-> Shifted left -0,2G
-> Shifted right +0,2G
Beam profiles obtained by scanning the beam with the dipoles (more precise than wire profilers)

## Slide 14: Higher order correction (up to 3rd order)
Hard to see a change, but a computer should

## Slide 15: Contents
The DESIR HRS
Optical aberrations : detection and correction
Spectrometer characterisation

## Slide 16: Measuring the resolution of a spectrometer/separator
HRS source: no direct mass measurement/separation possible
But :
Dipoles work the same for mass or energy shift (momentum separation) at first order
It can also be observed in the HRS transfer matrices where: {x , } = {x , } = -31 cm/%
Resolution can be measured by measuring position/size of two beams with close energies

## Slide 17: Populate multiple energies of a beam
Acceleration of ions depends on the potential between the source and the beamline
High voltage supplies can’t handle fast and small voltage variations (less than 1V on many kV)
A pulse generator can supply such variations

## Slide 18: Signal generator : square signal
Duplicates the beam into 2 close energies beams
Square signal
Populates ONLY
two energies
Energy separation can be arbitrary set
V : +3V
Energy 1 = 25003eV
V : -3V
Energy 2 = 24997eV

## Slide 19: Square signal with noise
V : +3V
Energy 1 = 25003eV
A square signal populates two and only two energies.
Adding noise to the signal increases the energy dispersion of the beam.
V : -3V
Energy 2 = 24997eV

## Slide 20: Square signal with noise
V : +3V
Energy 1 = 25003eV
A square signal populates two and only two energies.
Adding noise to the signal increases the energy dispersion of the beam.
V : -3V
Energy 2 = 24997eV
The HRS resolution can be caracterized as a function of the beam energy dispersion

## Slide 21: To go further…
A signal with multiple steps with adjustable amplitude and length can create (almost real) beam contaminants
Create custom beam contaminants
The HRS can be commissioned in almost real operating conditions, with no radioactive beam and (relatively) high intensities
0V
Period = 1
Period = 5
Period = 0.2

## Slide 22: A concrete example : separation of 100Sn and 100In
Simulation (COSY infinity)
Experimental separation
Prod. Ratio
1:605
Highly produced contaminants with close masses are still difficult to separate, in our case :
The major quantity of the contaminant can be separated.
Beam can be almost totally purified by sacrifying a part of the beam of interest.
Send the beam to a higher-level purification device (Penning trap : PIPERADE).

## Slide 23: Conclusion
Optical aberrations fully corrected at 2nd order and partially at 3rd order
Best resolution at FWHM : R = 23400 for 25keV Cs beam with 1-2 mm.mrad emittance and 1mm beam
Contaminants creation technique to test HRS in real conditions
To do:
  Correction coils to fix dipoles magnetic lenght
  Auto-tuning CorrAb software
  New emittance-meter under development at LP2i Bordeaux (HRS specific)
  Poles re-shaped at 6,9m to naturally correct 2nd order => ongoing
 HRS should be sent fully operational to DESIR by 2024-2025

## Slide 24: Conclusion
Optical aberrations fully corrected at 2nd order and partially at 3rd order
Best resolution at FWHM : R = 23400 for 25keV Cs beam with 1-2 mm.mrad emittance and 1mm beam
Contaminants creation technique to test HRS in real conditions
To do:
  Correction coils to fix dipoles magnetic lenght
  Auto-tuning CorrAb software
  New emittance-meter under development at LP2i Bordeaux (HRS specific)
  Poles re-shaped at 6,9m to naturally correct 2nd order => ongoing
Bonus : the bat-beam
THANK YOU !