## Slide 1
CAT’S UP
GANIL/CENBG
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
TRACKER FOR GANIL EXPERIMENT

## Slide 2
What is CATS and why an upgrade?
  Detector
  Electronic and DAQ system
  Current use and result
Detector Upgrade
Electronics Upgrade (Renocats)
  History
  New Front-end
  SAM backend
Status and raodmap
Summary
GANIL/PHY/DELTA/Sebastien Herlant
07/10/2019

## Slide 3
What is CATS
CATS is a Low presure multiwire proportionnal chamber for secondary beam tracking
Goal:
  Providing event by event particle tracking until 5.10^5 pps
  Time of flight measurement (accuracy between 440ps to 1,2ns)
  Position of ion (spatial resolution 400-700µm)
  Efficiency close to 100% with ion from He up to Ni at 40-50 MeV/u
Particularity:
  Very transparent detector (4,4µm of Mylar foil)
  Fast response to accept high flux
  Aluminium strips for low budget material
  Low pressure gas (6-8 mbar)
Typical use:
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
Silicon detectors
And/or Gamma detector around the target
beam
Two detectors CATS
Separated by around 50cm
S.Ottini-Hustache et al, CATS…,NIMA431(1999).

## Slide 4
The detector
Design
  One wire anode (70 wires of 20µm every 1mm)
  Two cathodes aluminium strip X and Y (28 strips/pitch 2,54mm, 0,9µm thick Mylar foil with 200nm Aluminium strips evaporated on Mylar)
  Area of 70*70mm^2 and 6,4mm gap
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
Difficulties:
  Very transparent and fragile detector (4,4µm of Mylar foil)
  Fast response to accept high flux
  Aluminium strips for low budget material
  Low pressure gas: 6/8 mbar
Cathode fragile
need regular, difficult maintenance

## Slide 5
The electronic
Front-end
  CPLear: Differential integrator with very low capacitance
  Mainly electron signal over 40ns (ion currents last about 1.5 s)
  Very low dead time
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
Back end
  Delay line 500ns (500m cables) by channel for coupling with master detector
  Digitization by QDC on VXI system
  Acquisition coupling with new system (VME//Numexo/µTca) by centrum module
Very hard to manage
precision on low capacitance
Obsolete system / No more production and difficulties to couple DAQ with the others system
Cplear design

## Slide 6
Actual result
Used currently on Mugast Campaign at Lise
  3.10^5pps during 300 hours with no or one detectors changing
  Normalization on beam intensity, checking beam purity
  Incidence ion on target for reconstruction
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
Multiplicity around 5 channel
Picture from Lise Mugast campaign 2023

## Slide 7
Detector Upgrade
Easier maintenance detector
Independence of striped mylar cathode from PCB signal reading
Same pieces for X and Y plane and possibility to put wires plan in X or Y position.
Mechanical Connection stripped without silver paste
Mechanical connection for High tension on the anode (no soldering wire)
Modification of wire plan for capacitance isolation of strips planes from HT line in calibration mode
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
New design: M.Riallot (IRFU) followed by P.Gangant (Ganil)

## Slide 8
Electronic Upgrade
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
A.Garzon Camacho et al, SEDA electronics (2012).
History
  Upgrade idea: Use current preamplifier for fast response signal
  KBTD: first design at IRFU on 2008: Pb of noise, project stopped
  Review on 2013 by University of Sevilla KBTD for SED tracker without changing DAQ system
  Prototype and publication made but no finalization of the project and loss of the details studies and simulation.
  Opening Cat’s up project on 2018 at Ganil with CENBG collaboration to review the SEDA design and coupled it with SAM acquisition system (Get system like)

## Slide 9
Electronic Upgrade
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
SEDA design:
Transimpedance wideband amplifier and
triple amplifier/shaper
Front end design by current preamplifier
  Response signal is independent of integration time and just proportional to the charge.
  Wide bandwith : so very high sensibility to noise and impedance variation
Characteristic
Spark protection
TIA: AD8015 IC
  80dB/10KΩ
  240MHz BW
  3pA/√Hz noise
Shaper block :LMH 6733
  Rise and fall time adjustment
  HPF and LPF cut-off biquad
  40dB/decade reduction noise
  Note: BW and noise decrease with input capacitance
Goal:
dynamic: 20fC to 400fC

## Slide 10
Electronic Upgrade
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
Very Compact design
Delay adjustable to 5 µs
Back end system
  SAM design from CENBG (Get light system)
    Squeeze Get premaplifier.
    Use shapping Get (Sallen-Key Filter)
    Digitization with one GET chip (64-channels)
    Acquisition by Zedboard (Commercial card): No CoBo Card
    Coupling with dedicated card (collaboration with GTA Ganil)
Old system
New system
Electronic integration
SAM+Zedboard

## Slide 11
Results
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
Pulse 50mV
pulse 25mV
Global response to 3 différents signal input
Detector response alpha source

## Slide 12
Status and roadmap
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
Detector Upgrade
  Prototype validated
  Production made (4 detectors)
  Need final tests with definitive electronic and beam commissioning expected in 2025
Electronics Upgrade
  Prototype characterized and lot of behaviours understood
  New design and production of mother front end board to be tested for S/N improvement by :
    Reduction of strips length before preamplifier
    Stripped embedeed into PCB (more parasitic capacitance but less CEM sensibility)
  Actual test and optimization of centrum card adaptation for zedboard
  Mechanical electronics integration design to be adjusted

## Slide 13
Thanks for your attention
GANIL/PHY/DELPH/Sebastien Herlant
19/03/2024
GANIL team
  Sebastien Herlant
    Julien Pancin
  Charly Nicolle
  Thomas Roger
  Johan Goupil
  Frédéric Saillant
  Gilles Wittwer
  Sebastien Coudert
CENBG Team
  Patrick Hellmuth
    Abdel Rebii
  Fréderic Druillole