## Slide 1: Matterhorn, a high flux, single photon counting detector for 4th generation synchrotrons
Roberto Dinapoli, Aldo Mozzanica, Anna Bergamaschi, Bernd Schmitt, Carlos Lopez Cuenca, Christian Rueder,Davide Mezza, Dhanya Thattil, Dominic Greiffenberg, Erik Fröjdh, Filippo Baruffaldi, Jiaguo Zhang, Jonathan Mulvey, Julian Heymes, Khalil Daniel Ferjaoui, Kirsty Paton, Konstantinos Moustakas, Mar Carulla Areste, Martin Brückner, Martin Mueller, Partick Sieberer, Saverio Silletta, Shuqui Li, Simon Ebner, Thomas King, Vadym Kedych, Viktoria Hinger, Viveka Gautam, Xiangyu Xie.
11th PIXEL conference, Strasbourg France, 18-22 November 2024

## Slide 2: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021

## Slide 3: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
For the detector this means:
Faster count rate
Higher frame rate
Wider energy range
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021

## Slide 4: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
For the detector this means:
Faster count rate
Higher frame rate
Wider energy range
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021
- 15-20 M photons/pixel/second
- 20 kHz continuous frame rate
- 250 eV – 80 keV

## Slide 5: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
For the detector this means:
Faster count rate
Higher frame rate
Wider energy range
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021
- 20 kHz continuous frame rate
- 250 eV – 80 keV
LGADs
High Z
- 15-20 M photons/pixel/second

## Slide 6: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
For the detector this means:
Faster count rate
Higher frame rate
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021
- 20 kHz continuous frame rate
Charge integration
Single Photon Counting
- 15-20 M photons/pixel/second

## Slide 7: The “speed challenge”: high photon flux
Single photon counting works extremely well if the incoming photon rate is reasonable low

## Slide 8: The “speed challenge”: high photon flux
Single photon counting works extremely well if the incoming photon rate is reasonable low
But as the rate increases we lose efficiency due to pulse pileup

## Slide 9: The “speed challenge”: high photon flux
Single photon counting works extremely well if the incoming photon rate is reasonable low
But as the rate increases we lose efficiency due to pulse pileup
Pile-up mitigation techniques for SPC:
Time over threshold
Retriggering
…

## Slide 10: Pile-up counting with multiple thresholds
Paralyzable counter:
Probablility of two and three events pile-up:
Th as fraction of photon energy
[1] Glenn F. Knoll, Radiation Detector and Measurement 4th Edition, Wiley
[2] M. Andrae The MYTHEN III Detector System - A single photon counting microstrip detector for powder diffraction experiments ETHZ Doctoral Thesis
Counter 0
Counter 1
Counter 2
50%
130%
170%
100%
100%: Energy of the monochromatic incoming photons

## Slide 11: Multi-thresholding : proof of concept with Mythen3.02 (1D)
Settings: standard
Energy: 15 keV
90% efficiency at:
Single counter : 1.03Mcps
MS Beamline @ SLS

## Slide 12: Multi-thresholding : proof of concept with Mythen3.02 (1D)
Settings: standard
Energy: 15 keV
90% efficiency at:
Single counter : 1.03Mcps
Three counters : 6 Mcps
Noise: 175e- RMS
MS Beamline @ SLS

## Slide 13: Multi-thresholding : proof of concept with Mythen3.02 (1D)
Settings: standard
Energy: 15 keV
90% efficiency at:
Single counter : 1.03Mcps
Three counters : 6 Mcps
Noise: 175e- RMS
MS Beamline @ SLS
Settings: Fast [1]
Single counter 3.52 Mcps
Three counters 20.87 Mcps
Noise:~250e-
[1] M Andrä The MYTHEN III Detector System - A single photon-counting microstrip detector for powder diffraction experiments ETHZ PhD Thesis 2021
x6
x6

## Slide 14: Matterhorn – a new single photon counter for SLS2.0
 | EIGER | MATTERHORN
Technology | UMC 250 nm | UMC 110 nm
Pixel size | 75 x 75 um2 | 75 x 75 um2
Thresholds | 1 | 4
Counter depth | 12 bit (3x4) | 4x16 bit
90% efficiency | 350k counts/pixel/s* | 15-20M counts/pixel/s
Module size | 4x8cm2 | 4x8cm2
HighZ/LGAD | No | Yes
*DECTRIS EIGER2 90% efficiency at ~2 Mcps (UMC 110nm)
T. Donath et. al. EIGER2 hybrid-photon-counting X-ray detectors for advanced synchrotron diffraction experiments J. Synchrotron Rad. (2023). 30, 723-738

## Slide 15: Matterhorn pixel architecture
Pixel
1)
1) Both e- and h+ readout possible
1)
1)
(Vrf voltage)
(Vrfsh voltage)
Static power/pixel (4 comparators) | 9.6μW

## Slide 16: Matterhorn pixel architecture
Pixel
2)
2) One comparator can be connected to two counters (pump-probe)

## Slide 17: Matterhorn pixel architecture
Pixel
3)
3) Each comparator can individually be ‘masked’ (disconnected from the counter)
 reduce switching for noisy channels

## Slide 18: Matterhorn pixel architecture
Pixel
4)
4) Each comparator can individually be powered-off when noisy or not used

## Slide 19: Periphery blocks
PLL
  Locking range: 930 MHz ÷ 2600 MHz
  Power: 18 mW
Pixel & counter selector
CML driver
  Main driver and pre-emphasis current tunable (3 bits): 3 mA ÷ 15 mA
  Pre-emphasis delay tunable (2 bits)
Pixel/counter addressing logic
Pixel/counter readout logic (includes 16 or 8 bit mode)
8b/10b encoding
High speed serializers
All periphery logic synthesized with digital tools for the first time!

## Slide 20: Matterhorn 0.1, 12.2022
ASICs: ~5x5 mm2 (3.6x3.6 active=48x48=2304 pixels)
MPW chips bump-bonded to sensor
Matterhorn 0.2,10.2023
One pixel version only
2x 3.125 GHz serializers
3.125 GHz CML driver
Internal DACs for biasing
Some “features” fixed
Matterhorn prototypes
2 Pre/shaper variations
2 Comparator variations
2x 1.6 GHz serializers
1.6 GHz CML driver (with pre-emphasis)

## Slide 21: Matterhorn 0.1, 12.2022
ASICs: ~5x5 mm2 (3.6x3.6 active=48x48=2304 pixels)
MPW chips bump-bonded to sensor
Matterhorn 0.2,10.2023
One pixel version only
2x 3.125 GHz serializers
3.125 GHz CML driver
Internal DACs for biasing
Some “features” fixed
Matterhorn prototypes
2 Pre/shaper variations
2 Comparator variations
2x 1.6 GHz serializers
1.6 GHz CML driver (with pre-emphasis)
PLL working
  Jitter < 5 ps
  (limited by measurement setup)
HS logic/data transmission working
  BER< 10-15
(limited by measurement time)
DACs working
  8 bits

## Slide 22: Basic characterization of MH02
Single pixel data
Metrologie beamline at SOLEIL with help from: Marie Andrä and Arkadiusz Dawiec
Calibration at: 7, 9, 12, 15 keV
Threshold scans fitted with analytical model:

## Slide 23: Four counters scan
Scanned at the same time

## Slide 24: Noise performance of MH02
fast - slow
Preamplifier feedback voltage
(gain)

## Slide 25: Threshold dispersion at 9 keV
vrf | Untrimmed [eV] | Trimmed [eV]
200 | 998 | 59
300 | 905 | 43
400 | 668 | 33
500 | 477 | 29
Small chip (48x48 pixels)
Non optimized trimming procedure
8e-
:16 to 20

## Slide 26: Single counter rate tests at MS beamline - MH01
Predicted rate capability using analytical model
Measurement

## Slide 27: Single counter rate tests at MS beamline – MH01
Note: slightly over estimated due to kB peak

## Slide 28: Single counter rate tests at MS beamline – MH01
Noise vs. rate capability

## Slide 29: Multi threshold rate tests at SOLEIL with MH02
Sum counter 0-3
Counter 0
x6

## Slide 30: Multi threshold rate tests at SOLEIL with MH02
Sum counter 0-3
Counter 0
Note: fitted with simple paralyzable function
Sum counter 0-3
Counter 0
x6
x3

## Slide 31: Multi threshold rate tests at SOLEIL with MH02
Sum counter 0-3
Counter 0
Sum counter 0-3
Counter 0
x6
x3
Comparator biased in “low power” mode
Vrf can be set even lower (100-150)
Preamp/shaper biases can be further tuned

## Slide 32: The “speed challenge”: high frame rate
 | EIGER | MATTERHORN 1
Technology | 250 nm | 110 nm
Chip outputs | 32xCMOS, 200 Mb/s | 4xCML, 3.125 Gb/s
Single chip | 6.4 Gb/s | 12.5 Gb/s
Module (8 chips) | 51.2 Gb/s | 100 Gb/s
Readout channels/module | 2x10Gbit/s | 1x100Gbit/s
Frame rate | ~10 kHz 8bit (burst) | ~20 kHz 8bit (continuous)
Continuous R/W | Yes | Yes

## Slide 33: Timeline
First prototype back!
First prototype submitted (MPW)
Second prototype (MPW) Submitted
SLS2.0
Matterhorn
Second prototype back!
Single module detectors?
Full ASIC submission
Multi module
detectors
Planned submission
of third prototype

## Slide 34: Timeline
First prototype back!
First prototype submitted (MPW)
Second prototype (MPW) Submitted
SLS2.0
Matterhorn
Second prototype back!
Single module detectors?
Full ASIC submission
Multi module
detectors
Planned submission
of third prototype

## Slide 35: Matterhorn 1, the full-size chip
Pixel data transfer:
  Current mode signal distribution (to pixel periphery)
New periphery, additional features:
  Global signals distribution (RES,STO,EN<0:3>)
  Counter summing mode
  4 bit mode
  64b/66b encoder for serializers
  SPI interface for configuration/slow control
Gating (4 independent EN) <20ns
No active circuitry on the chip sides (less dead area)
 | MATTERHORN 1
Technology | 110 nm
Chip active area | 1.92x1.92 cm2 256x256 pixels
Pixel size | 75 x 75 um2
Thresholds | 4
Counter depth | 4x16 bit
Trim bits/compar. | 6
Static power/pixel (4 comparators) | 9.6μW
HighZ/LGAD | Yes
Same pixel as Matterhorn 0.2
New features
MH0.2
MH 1.0
~2 cm

## Slide 36: Conclusion and outlook
Matterhorn is a new SPC detector for SLS2.0 targeting:
  20Mcps count rate with 10% loss of efficiency
  250 eV – 80 keV energy range (with LGADs and HighZ)
  20kHz frame rate; <20 ns gating
Results from MH01 and MH02:
Noise down to <~80e- in slow settings
Demonstrated ~7.5 Mcps at 90% efficiency using multiple thresholds
Tau of 42ns in single counter mode indicates 15Mcps should be possible with pileup counting
Tuning ASIC parameters
Submission of full ASIC spring 2025…

## Slide 37
Left to right: Jiaguo Zhang, Bechir Braham, Konstantinos Moustakas, Viktoria Hinger, Christian Ruder, Shqipe Hasanaj, Dominic Greiffenberg, Julian Heymes, Anna Bergamaschi, Khalil Ferajoui, Carlos Lopez-Cuenca, Xiangyu Xie, Pawel Kozłowski, Martin Brückner, Roberto Dinapoli, Kirsty A. Paton, Filippo Baruffaldi, Mar Carulla, Bernd Schmitt, Thomas King, Patrick Sieberer
Not in the picture: Simon Ebner, Erik Fröjdh, Davide Mezza, Aldo Mozzanica, Dhanya Thattil, Viveka Gautam, Vadym Kedych, Shuqui Li, Martin Mueller, Saverio Silletta
THANKS!

## Slide 38
Left to right: Jiaguo Zhang, Bechir Braham, Konstantinos Moustakas, Viktoria Hinger, Christian Ruder, Shqipe Hasanaj, Dominic Greiffenberg, Julian Heymes, Anna Bergamaschi, Khalil Ferajoui, Carlos Lopez-Cuenca, Xiangyu Xie, Pawel Kozłowski, Martin Brückner, Roberto Dinapoli, Kirsty A. Paton, Filippo Baruffaldi, Mar Carulla, Bernd Schmitt, Thomas King, Patrick Sieberer
Not in the picture: Simon Ebner, Erik Fröjdh, Davide Mezza, Aldo Mozzanica, Dhanya Thattil, Viveka Gautam, Vadym Kedych, Shuqui Li, Martin Mueller, Saverio Silletta
THANKS!

## Slide 39
Backup slides

## Slide 40: Pile-up mitigation techniques
Use charge integration with dynamic gain switching (Jungfrau, Jungfrau 2)
Time over threshold/retriggering
Credit: R. Ballabriga

## Slide 41: Operating modes of a single photon “counter”

## Slide 42: Comparison of pileup mitigation
[1] A. Bergamaschi et. al. Time-over-threshold readout to enhance the high flux capabilities of single-photon-counting detectors J. Synchrotron Rad. 18, 923-929. 2011
[2] P. Zambon, Dead time model for X-ray photon counting detectors with retrigger capability, NIMA 2021

## Slide 43: Pixel testability
Pixel
5)
5) The preamp can be pulsed with a programmable voltage step
Preamp and shaper outputs of pixel 0,0 can be monitored
6)
6)
6)
6)
6) Each counter can individually be pulsed
Comparator and “count” outputs of pixel 0,47 can be monitored
Pre
Shaper
Comp2
Count2

## Slide 44: The «energy challenge»: low photon energies
  Low quantum efficiency (QE):
  Development of “thin entrance window” sensor
  Low signal-to-noise ratio (SNR):
        Increase the signal with inverse Low Gain Avalanche Diodes (iLGADs) with a multiplication factor 5-20
  Results: 80 % QE for a planar pin diode @250eV
  Results: nominal noise of 35 e- with standard sensor at RT down to 3.7 e- at -22 oC with Moench
Soft X-ray applications at SwissFEL and SLS (250 eV – 2 keV)

## Slide 45: The «energy challenge»: high photon energies
CdZnTe is the material of choice above 15 keV
4×4 cm² sensor commercially available
Dynamic behavior (signal stability, afterglow) close to silicon sensors
Higher photon energies need to be studied
Crystal defects influence sensor uniformity
Fluorescence photons with long absorption range
GaAs:Cr is promising, but doesn‘t perform as well as CdZnTe
Two European suppliers, 8×4 cm² sensors available
 Afterglow acceptable up to photon fluxes
 Continuous: <10(8-9) ph/mm²·s  Pulsed: <10(6-7) ph/mm² ·pulse
Some signal loss at photon fluxes above > 3×1010 ph/mm²·s
Crystal defects influence sensor uniformity, variation of effective pixel size + Short hole lifetime  Craters
High dark current  For long exposure times: ideally SPC
      Hard x-ray applications at Synchrotrons, 15-80keV

## Slide 46: Timeline
First prototype back!
First prototype submitted (MPW)
Second prototype (MPW) Submitted
SLS2.0
Matterhorn
Second prototype back!
Single module detectors?
Full ASIC submission
Multi module
detectors
Planned submission
of third prototype
Move to new building