## Slide 1
LOREM IPSUM DOLOR
SIT AMET
Accelerators
M. Pullia

## Slide 2: Synchrotron
Dipoles (deflection)
Quadrupoles (focusing, 80 000 km!)
RF Cavity (acceleration)
R fixed E varies F varies B varies
Sorgenti

## Slide 3: The CNAO synchrotron

## Slide 4: ECR ion sources
Itot = 250 mA
I180p = 85% Itot
ECR, always on

## Slide 5
Sources and LEBT
LEBT Chopper
0.008 MeV/u H3+
0.008 MeV/u C4+
I ~ 0.5 mA (H3+)
I ~ 0.2 mA (C4+)
Continuous beam
Two ECR sources

## Slide 6
Linac system (RFQ + Linac)
IH
0.4-7 MeV/u H3+
0.4-7 MeV/u C4+
RFQ
0.008-0.4 MeV/u H3+
0.008-0.4 MeV/u C4+
217 MHz
RFQ
LINAC

## Slide 7: Stripping foils

## Slide 8
MEBT
7 MeV p
7 MeV/u C6+
I ~ 0.75 mA (p)
I ~ 0.12 mA (C6+)
(x,x’)Inj
(Current selection)
Debuncher
Emittance dilution
Stripping foil
Match betas

## Slide 9: Debuncher
TRACE 3D computation of the beam (P 0.75 mA) in the longitudinal phase space after the debuncher
TRACE 3D computation of the beam (P 0.75 mA) in the longitudinal phase space before the debuncher
PATH computation of the beam (P 0.75 mA) in the longitudinal phase space before the debuncher
PATH computation of the beam (P 0.75 mA) in the longitudinal phase space after the debuncher
2 gaps
20 (40) kV/gap

## Slide 10
Synchrotron
7-250 MeV p
7-400 MeV/u C
I ~ 0.1-5 mA (p)
I ~ 0.03-1.5 mA (C)
Multi turn injection
RF cavity
Betatron
RF-KO
kicker
Slow extraction
Betatron core and RFKO
Pushes the beam in resonance
(~1V acceleration)
Blows up emittance transversally
(~1 mrad kicks)

## Slide 11: Multi-turn injection
Bumper
Bumper
Septum
Closed orbit
Incoming beam
The bumper creates an orbit bump collapsing in tens of turns
First injection – 08-10
First multiturn injection – 08-10

## Slide 12
RF cavity
Betatron
Resonance
sextupole
D = D’ = 0
D = D’ ≠ 0
 | p inj | p – 60 MeV | p – 250 MeV | C6+ inj | C6+ – 120 MeV | C6+- 400 MeV
Br (T m) | 0.4 | 1.1 | 2.4 | 0.8 | 3.3 | 6.4
El Septum
Mag Septum
RFKO
2 Superperiods
2 Closed dispersion bumps
1 Dipole Family
3 Quadrupole Families
3 Sextupole Families

## Slide 13: Magnet cycle
Fixed ramps connected by linear ramps
Bmin
Binj
Bext
Bmax
Hysteresis cycle for repeatability

## Slide 14: Machine cycle
Inj bumper on
Multiturn
Rf trapping
Acceleration
Rf jump
Rf off
Start betatron
Stop betatron
Start hysterestis
End cycle
12/04/12
C. Biscari - LNF, INFN
14
BPM signal
BETATRON
# PARTICLES
DIPOLE field
Bmin
Bmax
Bextr
Binj
Slow extraction: ~1E6 turns

## Slide 15
Treatment execution
Sync
1 -3 s
Extracted beam
B ~ p~ E
acceleration
extraction
hysteresis ("washing ")

## Slide 16: Treatment execution with Multi Energy Extraction
Sync
1 s
Saved
time
Saved
time
(Courtesy K. Noda – CAS Medical)
Use remaining beam after slice completion

## Slide 17: Implementation of RFKO in addition to betatron
Dp/p
Amplitude
(b)
Amplitude-momentum
selection
X’
XES
Dp/p
Amplitude
(c)
RF-KO
X’
XES
Standard method
at CNAO
Being implemented
for multi energy
extraction studies
direct comparison
of the two methods

## Slide 18
HEBT
60-250 MeV p
120-400 MeV/u C
1010 p/spill (~2nA)
4 108 C/spill (~0.4nA)
3 Treatment rooms
2 H
1 H+V

## Slide 19
Chopper
Fast turn on/off for the beam
Intrinsically safe
Used for
- Turn off the beam in case of interlocks/problems
- Irradiation of separated parts of slices (if two voxels are far from each other >20mm)
- synchronization with breathing.
The chopper stops (and starts) the beam within 200 ms.

## Slide 20
Experimental room
+1 Experimental Room
3 Treatment rooms
2 H
1 H+V

## Slide 21
2 Strip chambers (X and Y):
Beam position measure every 100 s, with 100 m of precision
1 Pixel chamber:
Beam position and dimension measure every 100 s/1 ms, with 200 m of precision
1 Integral chamber:
Beam Intensity measure every 1 s
1 Integral chamber:
Beam Intensity measure every 1 s
Integral 1
STRIP X
STRIP Y
Integral 2
PIXEL
Box 1
Box 2
Dose Delivery System: in house development
Dose driven
Real time measurement
Feedback on scanning magnets

## Slide 22
Scanning control
< Δt > = 35.1 ± 3.5 µs between 20% to 80%
ΔI /Δt ~ 170 kA/s or ~ 85 T/sec
BEAM POSITION PRECISION
beam position precision specification of
0.1mm

## Slide 23: Patient positioning and position verification

## Slide 24: Motion in magnetic field

## Slide 25: Magnetic rigidity
F = q (E + v × B)
B
v
F
r
Force is orthogonal to velocity: change in direction, not in energy!
(for static B field)
Centripetal force
Centrifugal force

## Slide 26: Practical units
0.2998 B [T] r [m] = p [GeV/c]/q [e]
“Zero comma c”

## Slide 27
Small B, large r
Large B, small r

## Slide 28: Dipoles to deflect the beam
r
a
L = ar
Given a bending radius,
to deflect the beam by an angle a
a magnet of length L = ra
is needed

## Slide 29: Need for focusing

## Slide 30: Need for focusing

## Slide 31: Weak focusing
The first machines were built with weak focusing and a guide field which does not depend on the azimuthal angle.
If B is independent of the azimuth angle the equilibrium orbit is a circle of radius p = p/qB
A particle starting with an initial offset (but with the nominal energy) follows a circular path with the same radius. It remains near the nominal particle when they move (geometrical focusing)
A particle with the nominal energy starting with an initial vertical velocity drifts away

## Slide 32: Weak focusinsg
The motion is stable if for small deviations of the particles from the nominal (or better equilibrium) orbit restoring forces arise which lead to oscillations around the orbit. These are called betatron oscillations.
Stability requires that the Lorentz force is smaller than the centrifugal force for r < r and larger than it for r > r
n is called field index
Stability in the horizontal plane requires n < 1
(B non deve diminuire troppo)

## Slide 33: Weak focusing
An horizontal component of B is needed for vertical focusing
Altogether 0 < n < 1 is needed for simultaneous focusing

## Slide 34: Strong focusing
An alternating series of focusing and defocusing lenses leads to
overall focusing (if the distances between the lenses are not too large)
It’s like getting a discount the same percentage as the price increase
-20% and +20%
=> -4%

## Slide 35: Coordinates
s
s
s0
s1
x0
x1
Dx
Ds
x’0
Similarly fo z and z’

## Slide 36: Phase space
s
s
s0
s1
x0
x1
Dx
Ds
x’0
x’
x
x’
x
Dx
x0
x1
Ds
x’1=x’0
x’0

## Slide 37: Motion equations
After many simplifications and linearization,
using
Reference frame attached to ref. particle
s is the coordinate along the beam path
We get
Where

## Slide 38: Locally constant, hard edge model
Dipole (k(s)=0, 1/r2 ≠ 0)
Drift (k(s)=0, 1/r2 = 0)
Quadrupole
(k(s) ≠ 0, 1/r2 = 0)
Combined function
(k(s) ≠ 0, 1/r2 ≠ 0)
Harmonic Oscillator!
H)
I)
Sin(a s)’ = a Cos(a s)
Sin(a s)’’=[a Cos(a s)]’=-a2 Sin[a s]
=>
Sin(a s)’’+ a2 Sin(a s) = 0

## Slide 39: Useful transfer matrices

## Slide 41
Rectangular magnet

## Slide 43: Periodic solutions
K(s) is a periodic function. The period L may be as large as the circumference C of the accelerator, but the accelerator may be composed of N identical sections or "cells" with C = N·L.
In that case K(s) has the cell length L as a period

## Slide 45: The beta function
Phase advance
tune

## Slide 46
Parametric equation of ellipse

## Slide 48: Phase space ellipse

## Slide 49: Beam emittance
Statistical Twiss parameters

## Slide 50
Particle trajectory
Particle
Particle ellipse maximum size
Particle ellipse
Phase-space trajectory
Phase-space periodic looks
Foc. Quad. middle
Foc. Drift. middle
Defoc. Drift. middle
Defoc. Quad. middle

## Slide 51: Periodic dispersion
A particle with different energy moves along a different path
Dx
Dx = D Dp/p
Dx' = D' Dp/p

## Slide 52: Motion of particles with Dp/p ≠ 0
Design orbit
Design orbit
On-momentum particle trajectory
Off-momentum particle trajectory
Chromatic close orbit

## Slide 54: Phase slip factor

## Slide 55: Phase stability
Chart: RF voltage
t
TRF
TRF
2 phases have the
right energy gain

## Slide 56: Phase stability
Chart: RF voltage
t
TRF
TRF
What happens to particle near the nominal one?

## Slide 58: https://cds.cern.ch/record/235242/files/

## Slide 59
“If you think research is expensive, try disease!”
Mary Lasker