## Slide 1
Development of a self-calibration technique for gamma-ray tracking arrays
Sidong Chen and Stefanos Paschalis
University of York
4th AGATA-GRETINA/GRETA Workshop, 20-22 Nov 2024

## Slide 2
Gamma-ray tracking array
Gamma-ray tracking array: highly segmented HPGe (36 segments)
Pulse Shape Analysis (PSA) => interaction position
PSA performed by comparing with signal basis for every detector
PSA:
Energy (keV)

## Slide 3
Current challenges
signal basis generation
Experimental (scanning)
  long acquisition times
  different conditions between scanning and experiment, e.g. noise, radiation damage
  mechanical alignment
Analytical (calculated)
  intrinsic space-charge density
  the electron/hole mobility
  crystal temperature and
  crystal orientation
  passivated and contact thickness
  shape of charge cloud

## Slide 4
Self-calibration concept
Group interaction points from different gamma-rays into hit collections
Use Compton formula to order interaction points
Define tracks between interaction points that also link the hit collections with each other
Optimise coordinates of hit collection using the tracks that link their constituent points and Compton formula
S. Heil, S. Paschalis, and M. Petri, Eur. Phys. J. A (2018) 54: 172
Produce pulse shape basis for all detectors
Strong gamma source illuminate the whole array
Compton formula optimize scattering events
Generate signal basis experimentally & in situ
source
HC1
HC2
HC3
path2
path1
Det1
Det2
Det3

## Slide 5
Method validation timeline
simulation of realistic geometry of AGATA (even with 1Pi)
end-to-end simulation including pulse shape analysis
validation using experimental data (source and in-beam)
forthcoming, dedicated in-beam LNL experiment Dec 2024
2018
2019
2020
2021
2022
2023
2024
2025
2018: proof of principle with a simple simulation
2022: test with AGATA simulation with pulse shape
2023: test with real gamma source data at LNL
2024: produce optimized basis and validating with in-beam experiments

## Slide 6
Simulation with Pulse Shape
Geant4 simulation:
Geant4 simulate AGATA-1Pi array, save Compton scattering events
Pulse shape basis linear interpolation → simulation pulse data
Group pulse shape according to similarity
Normalised Charge
segments
Group pulse shape:

## Slide 7
Self-calibration using a full Simulation (with PSA)
initial
1st iter.
2nd iter.
5th iter.
HC position Y:X
dist: calib pos – real pos
source
HC1
HC2
HC3
path2
path1
Det1
Det2
Det3
X / mm
Y / mm
Y / mm
Y / mm
Y / mm
distance / mm
Simulate 2MeV gamma 2e10 events
Position optimization

## Slide 8
Experiment: source data
22Na 130kBq source at center of array
Large signals, CoreE>200keV (CoreE>300keV data is used for self-calibration)
Compton scattering events (fold 2 trigger)
source at target position
22Na source, 130kBq
13 ATC at LNL

## Slide 9
Pulse Shape Analysis (PSA)
ADL (AGATA Detector Library): theoretical calculated basis on 2 mm grid
Chi2: the difference between experimental signal and the signal basis fitting
The PSA final chi2 with self-calibration basis is smaller than that with ADL basis
⇒ self-calibration basis better describe experimental signal
PSA:

## Slide 10
Compton Scattering Angle
Compton scattering of 1274keV gamma
Interaction position from PSA with ADL basis and SelfCalib basis
Comparing scattering angle from PSA and the Compton angle from energy deposit
1274keV
E1
E2
theta
E1 + E2 = 1274keV

## Slide 11
GANIL benchmark data
(in beam)
A006
11A (DetId=33)
00A (DetId=0)
E680 fission data of 98Zr from 2015
Crystal A006 commonly used in GANIL setup and LNL setup
Electronics following the crystal can be different between two setups

## Slide 12
GANIL benchmark data
A006 spectra with ADL basis, self-calib basis and segment center

## Slide 13
A006 Signal basis
ADL basis
Self-calib basis

## Slide 14
GANIL benchmark data
Self-calib basis
slice<1
slice<2
slice<3

## Slide 15
PRISMA
Forthcoming, dedicated experiment at LNL (Dec 2024)
AGATA
28Si@175MeV
197Au
1mg/cm2
28Si@160MeV
PRISMA at 45°
197Au target 1mg/cm2 (0.52um)
28Si scattered to 45°, 160MeV (β=0.11)
28Si 2+ t1/2 = 0.475ps
Scheduled in December 2024

## Slide 16
Summary
2018
2019
2020
2021
2022
2023
2024
2025
2018: proof of principle with a simple simulation
2022: test with AGATA simulation with pulse shape
2023: test with real gamma source data at LNL
2024: produce optimized basis and validating with in-beam experiments
Technique validated with simulation data using realistic AGATA 1PI geometry and PSA
Applied to experimental data (22Na source at LNL) obtaining results consistent with simulations
Self-calibration basis (from LNL) applied to the GANIL in-beam data yielding reasonable results despite different setups
PSA performance using self-calibration basis and original basis shows improvement
A dedicated in-beam experiment at LNL (28Si Coulex) accompanied by a dedicated 88Y source calibration are scheduled for December 2024, stay tuned!

## Slide 17
Collaboration of LNL experiment Q24.001 scheduled for Dec 2024
S. Chen1, S. Paschalis1, M. Bentley1, M. Petri1, M. Labiche2, F. Holloway3, L. Harkness-Brennan3, A.J. Boston3, J.J. Valiente Dobón4, M. Rocchini5, A. Goasduff4,6,7, R.M. Pérez-Vidal4, D. Kalaydjieva8, K. Stoychev8, N. Marchini9, M. Zielińska10, A.O. Macchiavelli11, I.Y. Lee12, J. Dudouet13
1 University of York, York, UK. 2 STFC Daresbury Laboratory, Daresbury, UK. 3 University of Liverpool, Liverpool, UK. 4 INFN, Laboratori Nazionali di Legnaro, Legnaro, Italy. 5 Università degli Studi and INFN Sezione di Firenze, Florence, Italy. 6 Dipartimento di Fisica dell’Università di Padova, Padova, Italy. 7 INFN, Sezione di Padova, Padova, Italy. 8 University of Guelph, Canada. 9 INFN Firenze, Italy. 10 Irfu, CEA, Université Paris-Saclay, France. 11 Oak Ridge National Laboratory, USA. 12 Lawrence Berkeley National Laboratory, USA. 13 Université de Lyon, CNRS/IN2P3, Villeurbanne Cedex, France
Thank you for your attention

## Slide 18
Backup

## Slide 19
TAC comments
The PRISMA MWPPAC focal plane detector is known to have quite low efficiency for Z<20 ions. Please consider about 50-60% intrinsic efficiency for Z=14 ions.
If the statistics are lower than expect, we can combine data from several detectors.
Could the target (Au) excitation be a problem for you? You will have many gamma-ray transitions around 200-500 keV which may be problematic for your event selection. Have you considered the possibility to use instead Nb, Pt, Pb targets?
We don’t think the Au excitation will make serious problem. If it does, we can use Pb target instead.
Beam: The required energy from Tandem is high: in this configuration (double stripper in the Tandem) the required beam current is at the very limit: it could be that the maximum current achievable is lower, around 1pnA.
We will use 1-mg/cm2 Au target and 1pnA 28Si beam.

## Slide 20
Experimental Estimates
Beam intensity: 28Si 1-pnA
Target thickness: 197Au 1-mg/cm2
GOSIA estimated yield 400,000 gamma-particle coincidences per day
PRISMA rate 3kHz
PRISMA energy resolution 1/1000
MCP entrance detector position resolution 1mm

## Slide 21
Source calibration
Strong 88Y source, ~500kBq x2
source at some positions (target position, close to some detectors, ...)
Flod 2 trigger with core energy threshold 300keV or higher
Save trace data, validation rates below 1kHz per crystal
Data taking for 7 days
AGATA
source
Det 00
Det 39
Selfcalib basis from simulation data, 88Y source at target position, 1e11 decays

## Slide 22
Simulation: Position fidelity
select HCs
Select HC linked with large number of paths
Converged HC position reach ~1mm (RMS) fidelity
Slightly worse resolution in phi direction
distance / mm
distance / mm
Number of paths
Simulate 2MeV gamma 2e10 events

## Slide 23
Self-calibration basis (simulation)
Chi2 difference: self-calib pulse vs. real pulse
Large Chi2 observed around segment boundary
Simulate 2MeV gamma 2e10 events

## Slide 24
PSA position resolution (simulation)
Using the self-calibrated
basis
Using the calculated
Basis on a 2mm grid
difference / mm
difference / mm

## Slide 25
GANIL benchmark data
A006 spectra with GANIL cross-talk parameters and LNL cross-talk parameters

## Slide 26
Position resolution
Simulation data size influence to position resolution
Number of paths
distance / mm

## Slide 27
Self-calibration result
Identify incoming gamma energy by OFTtracking
consistent results between simulation and source calibration
Simulation data (same size as source data)
source data
Real HitCollection Position
Calib HitCollection Position
Calib HitCollection Position

## Slide 28
Simulation: Position fidelity
Distance between self-calib position and “real” position
Better than 1mm
1-2 mm
without electrical noise
with electrical noise
distance / mm
distance / mm
Number of paths
2MeV gamma 2e11 events

## Slide 29
Simulation with Pulse Shape
Geant4:
Signal Basis:
Y : X
Z : X
2mm
grid
Pulse shape signal basis
AGATA Detector Library (ADL)
AGATA-1Pi array: 45 detectors
Geant4: Compton events information
Linear interpolation ADL pulse shape basis

## Slide 30
Group Pulse Shape
Chi2s[0]: seg20, core
neighbour sectors:
Chi2s[1]: seg14, seg26
neighbour slices:
Chi2s[2]: seg19, seg21

## Slide 31
Group Pulse Shape

## Slide 32
Simulation: Position resolution
without noise
with noise
Simulate 2MeV gamma 1e11 events

## Slide 33
Gamma source
correctness 91%
correctness 73%
correctness 99%
correctness 96%
Simulate 1173keV
Simulate 1332keV
Simulate 898keV
Simulate 1836keV
Identify incoming energy according to figure of merit from OFT
60Co source
88Y source

## Slide 34
60Co source
Select events with total Energy of 1173±10keV and 1332±10keV
correctness 99%
Identify incoming energy according to total energy deposit in array

## Slide 35
60Co source
Simulate 60Co source 2e10 evts (energy gate ±10keV, ~2e8 good evts)
accurate source position
10mm mistake

## Slide 36
Source position
(0,0,0)
(0,0,-400)
(0,0,-160)
Source at (0,0,-400)
Source at (0,0,-160)

## Slide 37
Calibration data
multi segments hit
small signal
small Compton signal
accepted