## Slide 1
A Digital Twin of the Biograph Vision Quadra Long Axial Field of View PET/CT:
Monte Carlo Simulation, Image Reconstruction Framework and Clinical Applications
Wenhong Lan1, Ezzat Elmoujarkach2, Christian Pommranz3,4, Hong Phuc Vo2, Jorge Cabello5, Pia Linder1,
Julia Mannheim4, Andrea Santangelo3, Maurizio Conti5, Christian la Fougère1,6, Magdalena Rafecas2, Fabian Schmidt1,4
1 Department of Nuclear Medicine and Clinical Molecular Imaging, University Hospital Tübingen, Germany; 2 Institute of Medical Engineering, University of Lübeck, Germany; 3 Institute for Astronomy and Astrophysics, University of Tübingen; 4 Werner Siemens Imaging Center, Department of Preclinical Imaging and Radiopharmacy, University of Tübingen; 5 Siemens Medical Solutions USA, Inc., Knoxville, TN, USA; 6 Cluster of Excellence iFIT (EXC 2180) “Image Guided and Functionally Instructed Tumor Therapies”, University of Tübingen;

## Slide 2: Motivation
Simulation scenario A) Development of new systems (predict behaviour)
Pommranz, PMB 2023
More Details, Please visit us (H. P. Vo) 
Session: M-03, MIC Poster Session
Date & Time: Wednesday, 30.10.2024 - 14:00
Poster Number: M-03-058
Location: East / Central Hall
Notes: Add wsic + UZL logo + Hong Poster number and presentation time

## Slide 3: Motivation
B) Accurate Model of existing system (mimic behaviour)
axial FOV: 106 cm
Sensitivity: 176 kcps/MBq
Timing Resolution: 230 ps
Biograph Vision Quadra LAFOV PET/CT
at University Hospital Tuebingen

## Slide 4: Motivation
B) Accurate model of existing system (mimic behaviour)
Biograph Vision Quadra LAFOV PET/CT
at University Hospital Tuebingen
Overcome limitation of real patient examinations (e.g. repetitions, dose, examination time)
Provide "non-measurable" ground truth information (e.g. event characteristic, motion free reference, dose deposition)

## Slide 5
Digital Twin + Validation

## Slide 6: Concept Digital Twin Biograph Vision Quadra
* Vendor specific image reconstruction software
GATE model
Patient Simulation Model
Patient
PET/CT scanner
Image Data
e7 tools*
Image Data
Listmode data
e7 tools*
Image reconstruction
Listmode data
(proprietary)
Root2lm
Notes: Please add pictogram fro listmode, e7 tools, rootfile, patient

## Slide 7: Simulation + Image Reconstruction Workflow
GATE model
Image Data
Patient Simulation Model
root2lm (Siemens Healthineers)
Listmode data
(proprietary)
e7 tools*
Root2lm
Notes: This slide will be a zoomed in region of the

## Slide 8: Quadra Digital Twin - Collaboration
In submission
Julia Mannheim
Christian Pommranz
Ezzat Elmoujarkach
Wenhong Lan
Jorge Cabello
Pia Linder
H. P. Vo
Andrea Santangelo
Maurizio Conti
Christian la Fougère
Magdalena Rafecas
Fabian Schmidt
University
Hospital
Industry
Notes: Add picture of all people + logo WSIC, Uniklinik, Astrophysics, Siemens Healthineers, UZL
Add pictograms for university hosptiel and industry

## Slide 9: Validation: NEMA NU 2-2018
Replicate NEMA NU 2- 2018 experiments in silico
 Comparison of (real) experimental and simulation derived metrics

## Slide 10: Image Quality, Noise & Contrast Recovery
Comparable contrast recovery (e.g. 13 mm: 67.0% and 67.9%)
Comparable IQ + image noise - CV 7.5% (Sim) 7.7% (Exp)

## Slide 11: Spatial Resolution
0.3±0.3 mm (radial), 0.4±0.1 mm (tangential), 0.4±0.3 mm (axial)
Small difference in spatial resolution 0.4±0.3 mm

## Slide 12: Sensitivity
MRD322: 174.6 cps/kBq (Sim) and 175.3 cps/kBq (Exp)
MRD85: 82.7 cps/kBq (Sim) and 82.6 cps/kBq (Exp)

## Slide 13: Count Rate & Scatter Fraction
Scatter fraction 36% matches experiment
NECR: max deviation of 8%/9% (MRD85/MRD322) <= 10 kBq/mL range

## Slide 14
Motion Artefacts and Correction

## Slide 15: Motivation
Uncorrected
Motion Corrected
Image Courtesy: L. Sundar
W. Lan
H. P. Vo
F. Schmidt
C. Pommranz
E. Elmoujarkach
M. Rafecas
Notes: Nicht the best sondern comparing motion derived vs data driven

## Slide 16: Validation and Comparison of Motion Correction Methods
Subjective improvement + relative comparison of correction methods

## Slide 17: Validation and Comparison of Motion Correction Methods
Subjective improvement + relative comparison of correction methods
Lack of ground truth (motion free) information for activity distribution
 | Uncorrected | OncoFreeze AI | FALCON
SUVmean | 24.2 ± 11.4 | 35.7 ± 16.7 (+ 51.5 % ± 36.9) | 34.9 ± 28.8 (+ 41.5 % ± 58.6)
SUVmax | 36.6 ± 16.8 | 53.6 ± 24.7 (+ 50.0 % ± 35.2) | 52.5 ± 43.7 (+ 39.9 % ± 57.2)
Lesion Volume [cm³] | 0.84 ± 0.71 | 0.42 ± 0.48 (- 48.1 % ± 25.5) | 0.68 ± 0.70 (- 17.2 % ± 28.5)

## Slide 18: Digital Twin for Simulation Derived Motion Free Ground Truth
Simulation without (static motion free reference) and with different motion scenarios modelled
Comparison w/wo motion to evaluate motion induced error
Apply motion correction methods and compare with motion free ground truth
Digital Twin
Motion free (ground truth)
Realistic Motion Model
Motion correction
Motion corrected

## Slide 19: Digital Twin Motion + Correction: Results IQ
Model respiratory motion
Visual comparison w/wo motion
Apply motion correction:

## Slide 20: Digital Twin Motion + Correction: Results Lesion Quantification
Assess quantification changes/recovery for defined lesion size (Ø = 5, 10, 20mm) /uptake
Table 1. Mean activity concentration (kBq/ml) of lesions
Lesion Quantification
Motion correction → increase recovery of static activity concentration
Residual error lower for larger lesions (-12%) than for the smallest one (-28%) in the lung

## Slide 21: Digital Twin Motion + Correction: Results Lesion Profiles
Assess quantification changes/recovery for defined lesion size (Ø = 5, 10, 20mm) /uptake.
Table 2. FWHM [mm] of line profiles (SI direction) for lesions
Excellent recovery by MoCo: Ø=10 mm, FWHM: 8.0 (static), 8.6 (MoCo) vs 13.4 mm (no MoCo)
Lesion Line Profiles
Notes: Picture without caption and create bulletpoint for result

## Slide 22: Digital Twin Motion + Correction: Results Respiration Curve
Comparison of data derived respiratory curve (Oncofreeze AI) with input to simulation
Recovery of respiration curve with high precision (±0.5Hz) by Oncofreeze AI

## Slide 23: Quadra Digital Twin – Motion Presentation
More Details, Please visit us (W. Lan) 
Session: M-08, MIC Poster Session
Date & Time: Thursday, 31.10.2024 - 14:00
Poster Number: M-08-068
Location: East / Central Hall
Suitability of our framework to provide ground truth information on patient motion in a total-body PET/CT scanner
Next:
Extension towards variations in scan duration, activity concentration, lesion location/size and modelling whole-body motion
Evaluation of various motion correction algorithms to identify most effective application-specific solution and optimize the corrections based on the ground truth data of our framework.

## Slide 24
Y-90 Imaging and Radioembolization Dosimetry
Notes: Delete Nils, Christian, Ezzat, add la Fougère and elisa

## Slide 25: SIRT + Y-90 PET Imaging
Selective internal radiation therapy (SIRT) for liver tumor treatment with Y-90 microspheres
Y-90 beta emitter (99.98%)  Bremsstrahlung SPECT
0.0032% positron fraction  PET (extreme low event statistics)
 compensate with high sensitivity TB-PET
Sensitivity: 176 kcps/MBq
Timing Resolution: 230 ps
~ ×20 Sensitivity
W. Lan
F. Schmidt
P. Linder
J. Brosch-Lenz
C. la Fougère
E. Körner

## Slide 26: Optimization of Y-90 LAFOV PET Imaging
2i5s
3i5s
4i5s
1i5s
Objective metrics for phantom based imaging optimization

## Slide 27: Optimization of Radioembolization Dosimetry?
Patient image/dosimetry subjective or relative changes
Problem no ground truth on absorbed dose

## Slide 28: Dosimetry: Simulation Workflow
Emission Image
Quadra Digital Twin
MC Dose Simulation
MC Dose Simulation
optimum image reconstruction + residual error for dose estimation
Potential for NN trained on n>20 patients
Dose map
Dose map

## Slide 29: Dosimetry: Simulation Ground Truth
Emission Image
Quadra Digital Twin
MC Dose Simulation
MC Dose Simulation
optimum image reconstruction + residual error for dose estimation
Potential for NN trained on n>20 patients
Ground Truth
Dose map

## Slide 30: Dosimetry: Simulation Ground Truth - Application
optimum image reconstruction + residual error for dose estimation
Potential for NN trained on n>20 patients
Emission Image
Quadra Digital Twin
MC Dose Simulation
Dose map
Determine error introduced by scanner/imaging
Evaluate impact of different image reconstruction and correction methods
Scatter, Motion …

## Slide 31: SIRT Dosimetry : Dosimetry: Simulation Ground Truth - Application
optimum image reconstruction + residual error for dose estimation
Potential for NN trained on n>20 patients
Emission Image
Quadra Digital Twin
MC Dose Simulation
Dose map
Determine error introduced by scanner/imaging
Evaluate impact of different image reconstruction and correction methods
Benchmark different dose estimation methods
Scatter, Motion …
Kernel
…

## Slide 32: SIRT Dosimetry : Dosimetry: Simulation Ground Truth - Application
optimum image reconstruction + residual error for dose estimation
Potential for NN trained on n>20 patients
Determine error introduced by scanner/imaging
Evaluate impact of different image reconstruction and correction methods
Benchmark different dose estimation methods
Training dataset for neural network  enhance accuracy dosimetry

## Slide 33
Neural Network for Low-Dose PET Denoising

## Slide 34: Low Dose PET Imaging and Denoising
High sensitivity of TB-PET allows for dose reduction (e.g. 3MBq/kg  0.5MBq/kg)
Neural network for denoising: further dose reduction / recover full dose images
Requires pairs of low dose and high dose scans  typically gained from reframing high dose scans
Problem: simplified, not taking into account scanner characteristics as random rate, dead time, pile up effects, ….
 Use digital twin to assess the impact of reframing vs. actual real low dose input for a denosing NN
C. Pommranz
E. Elmoujarkach
F. Schmidt
M. Rafecas
J. Cabello

0
300s
PET Listmode data: (e.g. 300s scan)
Subset of event data to mimic low dose scan

## Slide 35: Digital Twin and Denoising Network
Digital Twin Biograph Vision Quadra
AI-denoising
XCAT phantom with different dose reduction factors (DRF) : real „injected“ and „reframed“
E. Elmoujarkach, PETAL-ED: Progressive elimination of noise towards accurate ultra low dose PET images using 3D U Net, PSMR, 2024

## Slide 36: Quadra Digital Twin – Training NN for Low-Dose PET Denoising
3 MBq/kg (ref.)
0.5 MBq/kg
0.25 MBq/kg
0.125 MBq/kg
De-noised data
Image quality comparable for reframed and real low dose data
Performance Metrics (PSNR, SSIM, NRMSE) showed similar values for both methods
Large difference in random fraction, e.g. DRF6 46% (true low dose) vs 9% (reframed)
Negligible difference in IQ and performance metrics for denoised data
PSNR: peak signal-to-noise ratio
SSIM: structural similarity index measure
NRMSE: normalized root mean squared error
CoR: coefficient of range
CV: coefficient of variation

## Slide 37: Quadra Digital Twin – Training NN for Low-Dose PET Denoising
More Details, Please visit us (E. Elmoujarkach) 
Session: M-03, MIC Poster Session
Date & Time: Wednesday, 30 October 2024, 14:00 - 15:48
Poster Number: M-03-061
Location: East / Central Hall
Re-framing high-dose PET scans provides an alternative to true low-dose data for training AI-based de-nosing.
Next:
Explore diverse activity concentrations and configurations, as well as potential differences in AI-based denoising at projection level

## Slide 38
I
Acknowledgements
Nuclear Medicine and Clinical Molecular Imaging
University Hospital Tuebingen
Christian la Fougère
Wenhong Lan
Pia Linder
Stefan Weigel
Thank you for your attention
Werner Siemens Imaging Center, Eberhard Karls University Tuebingen
Fabian Schmidt
Julia Mannheim
Jorge Cabello
Maurizio Conti
Noah Birge
Paul Schleyer
Antonis Kalemis
Sven Zhuelsdorff
Institute of Medical Engineering, University of Lübeck
Magdalena Rafecas
Ezzat Elmoujarkach
Hong Phuc Vo
Institute for Astronomy and Astrophysics,
Eberhard Karls University Tuebingen Andrea Santangelo
Christian Pommranz
Notes: Some clinical applications with Simulation.