## Slide 1: GATE Simulation for Sub-0.5 mm Resolution PET
Han Gyu Kang1, Hideaki Tashima1, and Taiga Yamaya1
            1. National Institutes for Quantum Science and Technology (QST), Japan
kang.hangyu@qst.go.jp
2024 IEEE NSS/MIC in Tampa
GATE User Meeting
Date: Thursday, 31-Oct-2024
Time: 12:15-12:30 (12 min presentation + 3 min Q&A)
Place: 105&106
Session: GATE User Meeting
Notes: Good afternoon. I’m Han Gyu Kang, a researcher at QST in Japan.
It is my great pleasure to present my research works in this session.
The title of my presentation is “GATE Simulation for sub-0.5 mm resolution PET”.

## Slide 2: Outline
Factors affecting PET resolution
Optimal design of a sub-0.5 mm resolution PET
Preliminary experimental results
  Front-end electronics
  PET gantry design
Conclusions
Notes: Here is the outline.
First, I’ll explain the factors affecting PET resolution.
Second, I’ll present the optimal design of a sub-0.5 mm resolution PET.
Third, I’ll be presenting the preliminary experiment results such as front-end electronics and PET gantry design.
Lastly, I’ll summarize my presentation and discuss the future plan

## Slide 3: Factors affecting the PET resolution[1]
Multiplexing
Photo sensor
Crystal
Scintillation detector
④ Crystal pitch
e-
Crystal
Photo sensor
β+
① Positron range
② Acollinearity
③ Parallax error
β+
Positron
e-
Electron
Annihilation（511 keV）
Ring diameter
Small ring diameter
(e.g. <50 mm)
Fine crystal pitch
(e.g. <1 mm)
Depth encoding
(e.g. 3-layer)
Opaque reflector
(e.g. BaSO4)
Solution>
⑤ Block effect
ESR
BaSO4
A PET scanner that considered these factors
[1] Moses 2010 NIMA
0.1 mm FWHM
18F
Notes: This slide illustrates the factors affecting the PET resolution.
PET resolution is defined as a convolution of several different factors as shown in this equation.
The first factor is the positron range which has about 0.1 mm in full width at half maximum for F-18.
The second factor is acollinearity which can be reduced by using small ring diameter of less than 50 mm.
The third factor is parallax error, in which the radial resolution is blurred at the periphery of a PET FOV.
The parallax error can be reduced by using depth encoding detectors such as staggered 3-layer configuration.
The most important factor is the crystal pitch, which dominates the overall PET resolution.
The PET resolution can be substantially enhanced by using fine crystal pitch of 1 mm or even smaller.
To encode a large amount of finely pixelated crystal array, multiplexing readout is commonly used.
However, for multiplexing readout, the block effect arises thereby degrading the crystal decoding accuracy.
The block effect can be effectively reduced by using optically opaque reflector such as barium sulfate powder rather than using the ESR film.
Last year, we developed a PET scanner that addressed these issues to achieve submillimeter-resolution.

## Slide 4: SR-PET achieved 0.55 mm resolution (2022 MIC)
Axial FOV
=51.5 mm
[2] Kang et al., 2023 Journal of Nuclear Medicine
0.85
0.8
0.75
0.55
0.5
0.45
PET image
SR-PET (Submillimeter-Resolution PET)[2]
SR-PET achieved the 0.55 mm resolution
Sub-0.5 mm resolution was NOT achieved due to
Large crystal pitch
Unoptimized crystal layer
Crystal pitch = 1 mm
Crystal layer = 4+4+7 mm (1st+2nd+3rd)-> DOI
1st layer
2nd layer
3rd layer
Pitch=1 mm
Goal: Design of a Sub-0.5 mm Resolution PET
Notes: The submillimeter-resolution PET, which is named as SR-PET has a 1 mm crystal pitch.
The 3-layer LYSO crystal arrays were stacked with staggered configuration to encode DOI information.
With the SR-PET, a resolution approaching 0.5 mm was achieved in a phantom study.
However, sub-0.5 mm resolution was not achieved due to the large crystal pitch and unoptimized crystal layer.
Therefore, the goal of this study is to design a sub-0.5 mm resolution PET by optimizing the crystal pitch and crystal layer configuration.

## Slide 5
SPECThead
SPECThead
crystal
pixel
ECAT
ecat
block
crystal
CylindricalPET
cylindricalPET
rsector
module
crystal
layer0
layer1
Optical System
crystal
pixel
OpticalSystem
CTscanner
CTscanner
module
cluster
pixel
PETscanner
PETscanner
level1
level2
level3
level4
level5
Scanner
scanner
level1
level2
level3
level4
level5
Predefined PET systems in GATE
No DOI
2-layer DOI
3-layer DOI
Notes: GATE provides various types of predefined PET systems such as ECAT, CylindricalPET and PET scanner.

## Slide 6
SPECThead
SPECThead
crystal
pixel
ECAT
ecat
block
crystal
CylindricalPET
cylindricalPET
rsector
module
crystal
layer0
layer1
Optical System
crystal
pixel
OpticalSystem
CTscanner
CTscanner
module
cluster
pixel
PETscanner
PETscanner
level1
level2
level3
level4
level5
Scanner
scanner
level1
level2
level3
level4
level5
PETscanner was used for 3-layer DOI PET
Predefined PET systems in GATE
Notes: Here, we used the “PETscanner” for the simulation of the 3-layer DOI PET scanner.

## Slide 7
GATE PET simulation
From GATE simulation to image reconstruction
ROOT output
Custom-written code
E.W=440-560 keV
List-mode data
Block ID
Crystal ID
DOI ID
Image Reconstruction
(In house software*)
*Developed by Dr. Hideaki Tashima (QST)
1st (level2ID)
2nd (level3ID)
3rd (level4ID)
Block (level1ID)
Reconstructed PET image
0.85
0.8
0.75
0.55
0.5
0.45 mm
10 mm
GATEv9.1
Goal: Optimization of a DOI PET geometry
Notes: This slide shows the procedure from the GATE simulation to image reconstruction.
First, we conducted the GATE PET simulation with 3-layer DOI detector configuration.
Second, the ROOT output file was imported by using a custom-written ROOT code to generate list-mode data.
An energy window of 440 to 560 keV was used.
Lastly, the list-mode data were reconstructed by using an in-house software developed in our lab.

## Slide 8: Sub-0.5 mm resolution PET design
Notes: First, let me describe PET design.

## Slide 9: Sub-0.5 mm resolution PET
ID=48 mm
2nd layer
1st layer
3rd layer
SiPM (5 × 5 array)
Phantom
<3D view of the PET>
<Staggered 3-layer DOI detector>
14×14
13×13
13×12
Crystal pitch
=0.8 mm
LYSO array
2.4 mm
23.4 mm
Lightguide
Notes: The PET scanner has an inner diameter of 48 mm and axial coverage of 23.4 mm.
The PET scanner has two rings, each of which has 16 DOI detectors.
Each DOI detector consists of 3-layer LYSO crystal array to encode DOI information.
The crystal pitch was 0.8 mm to achieve sub-0.5 mm resolution.
The 3-layer LYSO array was optically coupled to an SiPM 5x5 array with a pitch of 2.4 mm.
A 1 mm thick lightguide was inserted between the crystal and SiPM.

## Slide 10: Crystal layer designs
3 mm
5 mm
3 mm
4 mm
3 mm
4 mm
7 mm
4 mm
4 mm
1st layer
2nd layer
3rd layer
(A) 4+4+7 mm
SiPM 5×5 array
1st layer
2nd layer
3rd layer
(B) 3+4+4 mm
SiPM 5×5 array
1st layer
2nd layer
3rd layer
(C) 3+3+5 mm
SiPM 5×5 array
15 mm
11 mm
11 mm
Notes: We used three different crystal layer designs.
The first design-A has the crystal thickness of 4, 4 and 7 mm for the 1st, 2nd, and 3rd layers thereby resulting a total thickness of 15 mm.
The second design-B has the crystal thickness of 3, 4 and 4 mm, resulting in a total thickness of 11 mm.
The design-C has the crystal thickness of 3, 3 and 5 mm for the 1st, 2nd, and 3rd layers.
The relatively short crystal thickness was used for the design B and C to improve the DOI resolution at the expense of sensitivity.
The crystal thickness of each layer was chosen to make a uniform detection efficiency for each layer.

## Slide 11: Spatial resolution
Notes: Let me move on to spatial resolution

## Slide 12: Spatial resolution with the NEMA 22Na sources
52.5 mm
(A) 4+4+7 mm
Sensitivity
15 mm
5.34%
22Na
15 mm
48 mm
(B) 3+4+4 mm
11 mm
3.99%
22Na source
15 mm
48 mm
(C) 3+3+5 mm
11 mm
4.07%
22Na source
15 mm
5 mm
E.W=440-560 keV
Notes: The spatial resolution was evaluated by using a NEMA Na-22 source with a diameter was 0.25 mm.
The Na-22 sources were placed from the center to radial offset position of 15 mm with a spacing of 5 mm.
The spatial resolution was evaluated for three different crystal layer designs of A, B, and C.

## Slide 13: Spatial resolution (FBP)
(A) 4+4+7 mm
(a) NEMA 22Na sources images
Rad. Offset [mm]
0
5
10
15
1D Line profiles
Chart: Radial offset [mm]
(B) 3+4+4 mm
(C) 3+3+5 mm
FWHM
The design (C) resulted in the best spatial resolution of 0.72 mm at radial offset of 5 mm
(b) Radial resolution with FBP
Chart: Radial offset [mm]
| 0 | 5 | 10 | 15
4+4+7 mm | 0.67 | 0.95 | 0.97 | 1.23
3+4+4 mm | 0.64 | 0.76 | 0.84 | 1.02
3+3+5 mm | 0.71 | 0.72 | 0.79 | 1.01
(A)
(B)
(C)
Voxel size = 0.2×0.2×0.2 mm3
Notes: The PET images were reconstructed with FBP algorithm for three different crystal layer designs.
A line profile was extracted from the FBP images. Then, full width at half maximum was calculated from the line profiles.
The right figure shows the radial resolution as a function of radial offset.
The design-C resulted in the best spatial resolution over the design-A and –B.
The spatial resolution was 0.72 mm at the radial offset of 5 mm.
Therefore, we chose the design-C as the optimal crystal layer design.

## Slide 14: Resolution Phantom with Design (C) 3+3+5 mm
ID=48 mm
Phantom
<3D view of the PET>
23.4 mm
11.9 mm
11.9 mm
Phantom dimension
Acquisition time = 10 min
Total activity of 18F= 0.95 MBq
10 mm
0.85
0.8
0.75
0.55
0.5
0.45 mm
Notes: To evaluate imaging performance with the design-C, a resolution phantom was placed at the center of the FOV.
The phantom has an outer diameter of 11.9 mm, and axial length of 10 mm.
The resolution phantom has 6 different sections, and the smallest rod diameter was 0.45 mm.
All the rods of the phantom were filled with 18F of 0.95 MBq.
The acquisition time was 10 min.

## Slide 15: Resolution Phantom with Design (C) 3+3+5 mm
Line profiles (D=0.45 mm)
0.85
0.8
0.75
0.55
0.5
0.45
Reconstructed PET image
10 mm
Subset=8
Iterations=50
Valley-to-peak ratio (VPR) = 0.543
Valley
Peak
Well below the 0.735 (Rayleigh criterion[3])
[3] Hellen et al., 2020 PMB
VPR histogram
0
12
Resolvability = 100%
Reyleigh criterion (i.e. VPR=0.735)
VPR=0.573±0.062
Notes: PET images were reconstructed by using OSEM algorithm with 8 subsets and 50 iterations.
In the reconstructed PET image, all the rods can be clearly identified.
For quantitative evaluation of spatial resolution, a line profile was extracted for the 0.45 mm rod diameter.
The valley-to-peak ratio was 0.543, which was well below the Rayleigh criterion.
If the valley-to-peak ratio is under the Rayleigh criterion, the line profile is resolvable.
The line profiles were obtained for all rods in three different orientations and valley-to-peak ratio histogram was obtained.
As you can see, all the 12 line profiles can be resolved with an average valley-to-peak ratio of 0.573 thereby resulting in a 100% resolvability.
Iteration = 50
3+3+5 mm (P=0.8 mm)
PET Mouse Phantom (F-18), 10 min

## Slide 16: Preliminary experimental results
Notes: Based on the simulation results, we recently started to develop a prototype PET scanner.

## Slide 17: Front-end electronics
Front-end board
Top
DAQ
Power
2.4 mm
SiPM5×5
A
B
C
D
Anode 25 Ch
SiPM
SiPM biasing
Cathode (+HV)
Anode (Signal)
SiPM5×5
Position
4×4=16 ch
Timing
4×1=4 ch
Rpz
Rf
-
+
Output
AD8000
Cpz
Rg
50Ω
Input
200 ns
200 mV
Timing signal
Pole-zero-cancellation circuit
Resistive network
Bottom
Amplifier board
Flexible Flat Cable
miniHDMI
14 mm
<150ns
Notes: The front-end board has SiPM5x5 array with a pitch of 2.4 mm.
The SiPM bias was applied to a common cathode and signals were extracted from anode.
The 25 anode signals were multiplexed into 4 positional signals by using a resistive network.
The multiplexed signals were transferred to an amplifier board via flexible flat cable.
The position signals were amplified with quad channel amplifiers.
The timing signals were amplified by using a fast amplifier followed by a pole-zero-cancellation circuit to obtain fast decay time of less than 150 ns.
The amplified signals were transferred to a DAQ via miniHDMI cables.
The power lines were supplied by using the same cables.

## Slide 18: Staggered 3-layer DOI detector
<Staggered 3-layer DOI detector>
Pitch=
2.4 mm
SiPM
5×5
<SiPM with front-end board>
Pitch=0.8 mm
LYSO 3-layer
Acrylic lightguide
(12.2×12.2×1 mm3)
14 mm
12.5 mm
1st layer
2nd layer
3rd layer
Notes: Two LYSO 3-layer arrays were mounted on top of the SiPM.
A 1 mm thick lightguide was inserted between the crystal and SiPM.

## Slide 19: Crystal map (Experiment)
Pitch=0.8 mm
Crystal map
The crystal map of a 3-layer LYSO crystal array with 0.8 mm pitch was well identified.
22Na
Front-end
1st layer
2nd layer
3rd layer
511 keV
1st layer
2nd layer
3rd layer
Notes: Currently, the front-end electronics are under development.
We plan to develop a full-ring PET system within this year.

## Slide 20: Conclusions
The optimal design of sub-0.5 mm resolution PET was investigated
  The design (C) 3+3+5 mm resulted in the best spatial resolution
  The 0.45 mm rod structure can be resolved clearly
1st layer
2nd layer
3rd layer
Design (C)
0.85
0.8
0.75
0.55
0.5
0.45
Phantom image
3 mm
5 mm
3 mm
Notes: Let me summarize my presentation.
The optimal design of sub-0.5 mm resolution PET was investigated.
The design-C resulted in the best spatial resolution.
The 0.45 mm rod structure can be resolved clearly with the optimized design-C.

## Slide 21: Progress of the PET development
ID=48 mm
Phantom
23.4 mm
GATE simulation
Optimize the crystal layer
Front-end development
Optimize the PET detector
Front-end
Amplifier
3D CAD gantry design
Notes: Based on the GATE simulation, a PET detector was developed and gantry design was finished.
Next year, I would like to show you the imaging performance of the proposed PET scanner.

## Slide 22: Thank you for your kind attention!
2024 Imaging Physics Group
April-2024 in QST
HanGyu
Tashima
Yamaya
17/17
Notes: This is end of my presentation.
Thank you for your kind attention.

## Slide 23: Design of the HR-PET
OD=92 mm
44 mm
23.4 mm
Signal (18 pin)
HV&TMP(18 pin)
(a) Front view
FOV
(b) Cutaway side views
Front-end board
Amplifier board
LYSO
Mini-HDMI
OP-AMP
3D printed jig
1
2
48 mm
Notes: The SR-PET scanner has an inner diameter of 52.5 mm, and the crystal thickness of 15 mm.
The number of blocks is 16, each of which consists of 3-layer LYSO array.
The number of rings is 4, thereby resulting in an axial FOV of 51.5 mm.
The bore diameter is 48 mm, and the outer diameter is 107 mm.
The PET scanner consists of front-end board and amplifier board.
The SiPM signals and power lines were transferred by using flexible flat cables.