## Slide 1
Cherenkov imaging dosimetry for heterogeneous phantom in the context of a MR-Linac
Audran Poher1, 2, 3, Janelle Morrier3, Tigran Galstian1, Louis Archambault1, 2, 3 and Luc Beaulieu1, 2, 3
1Département de physique, de génie physique et d’optique, Université Laval
2Centre de recherche sur le cancer, Université Laval
3Service de physique médicale et de radioprotection, Centre Intégré de Cancérologie et CRCHU de Québec, CHU de Québec

## Slide 2: Introduction
Past studies showed that polarized Cherenkov imaging dosimetry provides accurate continuous 2D dosimetry
MR-Linac dosimetry:
Using visible light represents a major advantage
Presence of the strong external magnetic field affects the electron trajectories

## Slide 3: The electron return effect
Main challenge :
Heterogeneities + = electron return effect (Lorentz force affect secondary trajectories)
Affects ion chamber readouts
Produces unconventional overdosages at interfaces
Electron return effect (ERE)
Roberts et al. (2021)
Malkov and Rogers (2017)
Use luminescence-based dosimetry
Notes: Main challenge comes from the combination of hereto + B_0, resulting in the electron return effect

## Slide 4: Cherenkov light emission
Cherenkov is partly polarized due to conical emission
Cherenkov production :
Charged particle faster than light in medium
Constructive interferences of the electromagnetic wave = photon emission
Cherenkov production threshold in water = 0.264 MeV
Average energy from a MR-Linac = 2 MeV
Cherenkov light can be used for luminescence-based dosimetry
Ashraf et al. (2020)
Peshkov (2019)

## Slide 5: Our study
Objective: characterize polarized Cherenkov imaging dosimetry in a heterogeneous phantom irradiated by a combination of fields from an MR-Linac
Malus law:
Cherenkov polarimetry:
Field combination
Fluence energy spectrum averaging with AP-PA (Glaser et al. 2014)
- 1% / 1 mm agreement
Deposited dose [Gy]
(Cloutier et al. 2022, Poher et al. 2026)
Heterogeneities:
Characterize the electron return effect with Cherenkov
Glaser et al. (2014)

## Slide 6: Predator filter
Predator V2 system (PATQER Photonique Inc.) (Pusenkova et al., 2025):
Polarizer-free electro-optic device
Nematic liquid crystal windows:
  positive dielectric anisotropy (electronically tunable)
  doped by a mixture of dichroic azobenzene molecules (Disperse Red 1)
Advantage: perform polarimetry and keep spatial resolution (no pixels in the filter)
Polarization Imaging User manual, PATQER Photonique Inc (2026).
Notes: - nematic: radomly disposed in the medium
-Doping allows for : dichroic macroscopic light absorption (parallel and perpendicular components are attenuated differently)
- Positive dielectric anisotropy allows for the filter to be electronically tunable

## Slide 7: Predator filter: working principles
40%
Layer 1:
Dichroic attenuation
Layer 2:
Adiabatic rotation 45°
Layer 1 and 2 excited
80%
Layer 1 (and 2) excited
Layer 1 ground state
Layer 1 and 2 ground state
Notes: Layer 1: straight liquid crystals
Layer 2: twisted liquid cristals, in ground state: adiabatic rotation at 45 deg
Combination of ground and excited states applied on layer 1 and 2

## Slide 8: Experimental setup
Camera’s point of view
Camera + Predator V2
MR-Linac
Phantom
Acrylic walls
Water
Beam 0°
Beam 45°
Beam 115°
Air cavity
Isocenter
17 cm
20 cm
6.7 cm
2.5 cm
3 cm
1.7 cm
ROI 1
ROI 2
ROI 3
15 cm
2 cm
ROI 1: upper horizontal (UH) interface
ROI 2: right vertical (RV) interface
ROI 3: isocenter region
Beam characteristics:
7 MV
600 MU at 425 MU/min
Phantom limitations prevented directly symmetrical opposed beams

## Slide 9: Polarimetric and dosimetric analysis
Polarimetric analysis (using the Predator software):
Image intensities at 0°, 45°, 90° and 135° polarization angles
AoP + DoLP map
Dosimetric analysis:
Projected percent depth dose
Gamma index maps with a 3% / 3 mm criteria
ROI 1: UH interface
ROI 2: RV interface
ROI 3: isocenter region
ROI 1: UH interface
ROI 2: RV interface
ROI 3: isocenter region
Notes: Image ressampling with RegularGridInterpolator from scipy.interpolate

## Slide 10: Reference Monte Carlo dose map
108 histories per beam
Beam simulated using phase space provided by Elekta (scored 129 cm from the source)
2 mm × 2 mm × 2 mm voxels
Modeled phantom geometry
TOPAS Monte Carlo toolkit version 3.9
  Geant4 version 10.7
Perl et al. (2012), Faddegon et al. (2020)
Agostinelli et al. (2003)
Gantry at 0°
ROI 1: UH interface
ROI 2: RV interface
ROI 3: isocenter region

## Slide 11: Angle and degree of linear polarization
AoP + DoLP
Variation of the AoP (5° → 81° → 164°) with the gantry angle (0° → 45° → 115°)
Angle of polarization relatively stable above the air cavity:
  Increase in Cherenkov signal at UH ROI not due to reflections
UH ROI
Noisy image
Increase irradiation time
Notes: Approx 5 deg +/- 5 deg angle variation in the UH ROI (ERE region) , relatively stable compared to other gantry orientations resulting in vastly different angle of polarization

## Slide 12: Projected percent depth dose
Averaging a 0.5 cm wide strip
Monte Carlo: 120% ERE peak
Cherenkov I0: 117% ERE peak
No ERE peak!
Noisy I0 signal (av. 8% rel diff to MC)
Non-pol. Che. av. rel diff to MC = 1.4%
ROI 1: UH interface
ROI 2: RV interface
ROI 3: isocenter region
ROI 1: UH interface
ROI 3: isocenter region
Air cavity

## Slide 13: Gamma index maps
ROI 2: RV interface
ROI 3: isocenter region
ROI 1: UH interface
ROI 2: RV interface
Criteria: 3% / 3 mm
94% of pix,
61% of pix,
ERE minimal
ERE minimal
Better agreement to MC
ROI 1: UH interface
ROI 3: isocenter region
AoP + DoLP
Notes: Greater non-polarized intensity vs polarized intensity in isocenter region
Greater non-polarized intensity vs 𝐼_0 polarized intensity in isocenter region

## Slide 14: Conclusion
Summary:
A heterogenous phantom was irradiated at 3 gantry orientations by the Unity MR-Linac producing Cherenkov emissions.
Using the Predator filter, polarization characteristics of the Cherenkov emissions were extracted.
Polarized and non-polarized images of the Cherenkov emissions were compared to Monte Carlo dose distributions
Key findings:
From the AoP+DoLP map : variation of AoP with gantry orientation
From the Gamma index maps:
  UH interface ROI: better agreement to MC with Cherenkov
  Isocenter ROI: better agreement to MC with non-polarized Cherenkov
Notes: No ERE peak is present in the UH ROI in the non-polarized Cherenkov map
Greater non-polarized intensity vs polarized intensity in isocenter region
No ERE peak is present in the UH ROI in the non-polarized Cherenkov map
Greater non-polarized intensity vs 𝐼_0 polarized intensity in isocenter region

## Slide 15: Future work
Encountered limitations:
Noisy AoP + DoLP image
1 polarization angle used
Gamma index: 3% / 3 mm criteria
No directly opposed beams
Proposed solutions:
Increase MU to 1200 (doubled)
Combination of polarization angles (Malus regression and extraction)
Gamma index: 1% / 1 mm criteria
Dose volume histograms (DVH)
Acrylic phantom

## Slide 16
Medical Physics Research Group, Université Laval, 2026
Acknowledgments
Cherenkov imaging dosimetry for heterogeneous phantom in the context of a MR-Linac
Audran Poher1, 2, 3, Janelle Morrier3, Tigran Galstian1, Louis Archambault1, 2, 3 and Luc Beaulieu1, 2, 3
1Département de physique, de génie physique et d’optique, Université Laval
2Centre de recherche sur le cancer, Université Laval
3Service de physique médicale et de radioprotection, Centre Intégré de Cancérologie et CRCHU de Québec, CHU de Québec

## Slide 17: Past study using Cherenkov
Past study (Andreozzi et al., 2020):
Used the MRIdian (= 0.345 T)
Used larger fields (8.3 cm × 8.3 cm and 24 cm × 4.15 cm)
Top-down irradiation
Our study:
Using the Unity (= 1.5 T)
Using smaller fields (2 cm × 2 cm)
Using combination of fields (gantry at 0°, 45 and 115)
Investigate polarization
Andreozzi et al. (2020)
Criteria: 3% / 3 mm
Phantom B:
Notes: I would like to start here by mentionning the past study by Andreozzi et al.
This study investigated similar factors

## Slide 18: Discussion

## Slide 19: Acrylic phantom: AoP + DoLP
Reflections