## Slide 1
Uncertainty-Aware Deep Learning Reconstruction for 3D Scintillation Dosimetry
Alexis Horik1,2,3, Raphaël Brodeur1,3, Nicolas Drouin1,3, Simon Thibault1,2, and Louis Archambault1,2,3
1Département de physique, de génie physique et d’optique, Université Laval, Québec, Canada
2Centre d’optique, photonique et laser, Université Laval, Québec, Canada
3Centre de recherche du CHU de Québec-Université Laval et Centre de recherche sur le cancer de l’Université Laval, Université Laval, Québec, Canada
International Conference on 3D Dosimetry 2026

## Slide 2: Presentation outline
Motivation and proposed approach
Overview of the 3D dosimeter
  Experimental setup;
  Deep learning reconstruction model.
Bayesian formulation
  Motivation and introduction;
  Bayesian neural network for tomographic reconstruction.
Results
  Uncertainty analysis;
  Bayesian vs deterministic models.
Discussion and conclusion
Notes: The presentation outline is as follows.
I will begin by motivating briefly the proposed approach and giving an overview of the 3D dosimeter that we are proposing.
After laying this ground work, I will dive into the bayesian formulation that allows us to have uncertainty information on the predictions.
Finally I will obviously present the results, and discuss limitations and conclusions.

## Slide 3: Toward real-time, uncertainty-aware 3D dosimetry
3D dose measurements are needed for satisfactory linac QA.
Real-time dosimetry is desirable
The reliability of dose measurements is crucial.
Dynamic treatments benefit from temporal information
Technicians’ and physicists’ time is valuable
Develop a novel dosimeter that can (i) measure the dose in 3D and in real-time, and (ii) estimate the reliability of its measurement.
Objective
Motivation and Context
Notes: So I don’t think it will be shocking to you that 3D dose measurements are needed or at least would be beneficial for linac QAs.
Another thing that is desirable is the real-time aspect of the measurements and this is for two main reasons:
  first, dynamic treatments benefit from temporal information so we can identify when an error has occured
  and secondly, the time of the people making those measurement is very valuable, so we want to make them as quick as possible.
The reliability of dose measurements is also crucial to make consistent QA decisions.
For these reasons the objective of the project is to develop a novel dosimeter that can:
  Measure the dose in 3D and in real-time
  And estimate the realibility of it’s measurement.

## Slide 4: Proposed approach
Dose measurement
Tomographic reconstruction
3D dose distribution estimation
QA metrics and decision-making
Dose deposition in scintillator
1
2
3
4
Scintillation light emitted
Optical system records projections
Input: Projection data
Tomographic reconstruction algorithm
Output: 3D distribution
Development of acceptance criteria.
Spatial information on the errors.
Flagging of reconstruction errors.
3D voxelized dose map.
Uncertainty estimation of the reconstruction.
Uncertainty map
Reconstructed dose
Example use:
Thresholding on uncertainty to determine if:
  The reconstruction is reliable.
If the reconstruction is reliable:
  Flag high-error regions;
  Guide targeted review of profiles or PDDs;
  Global acceptance criteria.
Motivation and Context
Notes: Here is the outline of the proposed approach:
The first step is the dose measurement, or more specifically the measurement of the scintillation light that is proportional to the dose deposited in the volume.
The second step involves the tomographic reconstruction of the scintillation, which we do using a deep learning model.
After that, we obtain the reconstruction dose, and with the bayesian forumation that I will present later, we also obtain the uncertainty map associated with the reconstructed distribution.
Lastly, the goal would be to implement QA metrics and decision making based on the reconstruction and uncertainty map.
  An example use would be to use thresholding on the uncertainty to determine if the reconstruction is reliable and if so, flag high-error regions or implement global acceptance criteria.
  Note that these are example use, and are not implemented yet.

## Slide 5: Dose Measurement
Notes: Going sequentially through these steps, let’s begin by talking about the dose measurement itself.

## Slide 6: Volumetric scintillation dosimetry
Source: Eljen Technology
Scintillation Absorbed dose;
Fast mechanism ( s);
Water-equivalent;
Dose rate independent;
Mostly energy independent.
Dosimetric Characteristics
The problem becomes…
Imaging of the scintillation light;
Tomographic reconstruction.
Motivation and Context
Notes: To do the physical measurements, we use volumetric scintillation dosimetry.
The characteristics of these devices are ideal from a dosimetric standpoint:
  The scintillation is proportional to absorbed dose and this mecanism is extremely fast, which allows for real-time measurement.
  They are water-equivalent, dose-rate independent and mostly energy independent.
The problem however becomes the imaging of the light field, since we are in free space.
We also need to reconstruct this light in 3D to retrieve the underlying dose distribution, which poses challenges for this kind of dosimetry.

## Slide 7: Experimental constraints
Source: Siemens Healthineers
Geometric constraints – Restricted space around the treatment isocenter.
Camera placement considerations – Radiation can damage cameras and gantry motion can obstruct optical paths.
Time – Physicists have limited time to perform measurements.
Setup complexity – Intricate setups can suffer from reproducibility problems.
Cost – Multiple cameras or recent technology (ex. plenoptic cameras) can bump up the cost of the device.
Motivation and Context
Notes: The problem of imaging the scintillation light is also enhanced by various experimental constraints.
For example, geometric contraints like the restricted space around the isocenter limit the kind of imaging setup that we can implement.
Moreover, camera placement is a challenge since the radiation from the limit can damage cameras and gantry motion can obstruct optical paths.
Time is also an important factor because we don’t want to waste precious time from the people performing those measurements.
Intricate setups can also suffer from reproducibility problems, so ideally we want a simple setup not requiring lots of alignements.
And lastly, we would want the cost to be as low as possible, so ideally no system with tons of cameras or recent technology that bump up the cost of the device.

## Slide 8: Experimental setup
Scintillating
Cube
Polyvinyltoluene
Plane
Mirrors
CCD
Camera
Proposed 3D dosimeter – Experimental data
Notes: Having that all in mind, we build this experimental setup.
It is composed of a 10x10x10 scintillator, coupled with two plane mirrors placed at 45 degrees angle so that we can aquire three orthogonal views from a single image.
The CCD camera is placed such that the field of view includes the three views in the frame
Having that all in mind, we build this experimental setup.
It is composed of a 10x10x10 𝑐𝑚^3 scintillator, coupled with two plane mirrors placed at 45 degrees angle so that we can aquire three orthogonal views from a single image.
The CCD camera is placed such that the field of view includes the three views in the frame

## Slide 9: Experimental data collection and curation
Background noise substraction ;
Raw image
Processed image
Intensity (16-bit scale) [ - ]
Pixel [ - ]
Nicolas Drouin
Temporal median filter.
Collect of 249 experimental data of the form:
Proposed 3D dosimeter – Experimental data
Notes: So this gives the kind of image that you see on the left hand side.
However, these images are quite noisy, so we need to post-process them.
This is work mostly done by nicolas drouin during an internship a few years back, he is now a master student in the group and just presented before the coffee break actually.
Anyway, by substracting background noise and applying a temporal median filter on multiple acquisitions, the noise was greatly reduced as you can see on the profile here.
And with the great help of Nicolas, we collected 249 experimental data, with their associated dose distribution computed by the TPS, which we can use as groundtruth.

## Slide 10: Tomographic Reconstruction
Notes: - Now that we have images the scintillation, we need to consider the tomographic reconstruction of this signal.

## Slide 11: Tomographic reconstruction of scintillation light
In our case, the problem is not simple…
High-dimensional inverse problem: voxels reconstructed from only 3 projections.
Underdetermined problem: many 3D dose distributions may be consistent with the same projection data.
Iterative methods are time-consuming.
High representation capacity
Fast reconstruction after training
Reconstruction from raw light signal.
Deep learning reconstruction
Proposed 3D dosimeter – Reconstruction
Beister et al. Physica Medica. 2012
Notes: In most cases, we use iterative reconstruction to estimate the solution of an inverse problem. So basically, at each iteration we update the reconstruction by comparing the forward projections of the current image with the input projections and update the image until we have a satisfactory solution.
==========================================================
However, in our case, even iterative reconstruction can have quite a hard time because of the high-dimensionality of the problem. So we have 1 million voxel to reconstruct with only 3 projections.
This also means that the problem is highly underdetermined, meaning that many 3D dose distributions may be consistent with the same projection data.
Iterative methods are also time-consuming, which is obviously not desirable in a real-time setting.
=============================================================
That is why we are using deep learning reconstruction in our approach. The high representation capacity of neural networks can help learn ill-posed problems.
They are also extremely fast at inference, so a reconstruction takes less than a second.
And they can reconstruct from raw light signal, while when using classical methods, one needs to filter out undesired light signal.

## Slide 12: Deep learning reconstruction model
Transformer
Encoder
3D CNN
Decoder
CNN
Embedding
Latent
Representation
Input
projections
Output
distribution
Encoder-Decoder framework
Proposed 3D dosimeter – Reconstruction
Notes: Here I’m showing the high-level schematic of the deep learning model that was used to make the reconstruction.
So briefly, here we see that the input is the three orthogonal projections presented before.
They then go through a CNN embedding and after a transformer encoder that brings the data to a latent representation of lower dimensionality.
After that, the data is up-sampled using a 3D CNN decoder that outputs the 3D dose distribution.

## Slide 13: Architecture
Proposed 3D dosimeter – Reconstruction
Notes: Going into a bit more details, in the transformer encoder, the data passes through multiple transformer layers so it can learn complex mapping before being down-sampled in the latent representation.
After that, a cascade of transposed and standard convolution layers brings the data to the desired resolution.

## Slide 14: Training data
Proposed 3D dosimeter – Training data
Synthetic data
Experimental data
Synthetic dataset
Validation set
Training set
Experimental dataset
Validation set
Test set
Training set
75%
25%
~70%
~18%
~12%
Data generated by ray-tracing in python.
Gaussian filter to approximate lateral dose falloff.
Empirical equation for depth dose deposition.
Data acquired using the experimental setup.
Notes: Deep learning models are known to be data-hungry, or to need a lot of data to generalize well. That is why we generated a synthetic dataset to pretrain the model before using the experimental data for finetuning.
For the pretraining, 30,000 data were generated using simple ray-tracing in python, and they were processed to approximate lateral and depth dose deposition.
After this pretraining was performed, the model could then be trained using the experimental data described previously.

## Slide 15: Bayesian Formulation
Notes: - Now that I have layed the ground work, we can focus our attention towards the bayesian formulation.

## Slide 16: Limitations of DL – Prediction without confidence
?
Hard to interpret (black box);
Limited view reconstruction remains ill-posed the model can “hallucinate”;
Produces a single dose estimate without confidence level we must “blindly” trust the model.
Limitations
Bayesian neural network
Predicts both dose and uncertainty;
Identifies regions with lower reconstruction confidence.
Bayesian formulation – Motivation and Introduction
Notes: So why am I even talking about that…
Well, classical deep learning has quite a big drawback, where the predictions are not accompanied with a confidence level.
Hence, these models being black boxes, we need to blindly trust their prediction, while they are known to hallucinate in certain context, especially if the problem is ill-posed like the one we have here.
We would then want to have a confidence level on the predictions, and it is exactly where Bayesian neural networks come into play.
=======================================================
These kinds of model can provide both prediction and uncertainty.
And in our context, it can allow for voxel-wise uncertainty prediction, which helps identify regions with lower reconstruction confidence.

## Slide 17: Bayesian Neural Network (BNN)
Deterministic model
is fixed after training
a given input always gives the same output
Blundell et al. Int conf on machine learning. 2015
Bayesian formulation – Motivation and Introduction
Notes: But how does it work ?
So in the case of deterministic models, we have artificial neurons that are connect by weights. The values of those weights are learned during training, but once the training is done, the weights are fixed such that an input x always gives the same output y.

## Slide 18: Bayesian Neural Network (BNN)
Deterministic model
Bayesian model
is fixed after training
a given input always gives the same output
In practice:
Blundell et al. Int conf on machine learning. 2015
Bayesian formulation – Motivation and Introduction
Notes: In the Bayesian model case, the weights are reparameterized by probability distributions.
Hence, during training, we learn the posterior distributions and doing a forward pass of the model means sampling from those probability distributions and then getting an output.
Because of the probabilistic nature of the approach, a given input now does not always give the same output.
So at inference, we would want to do the integral over all weights, but in practice we can’t do that, so we instead sample a finite set of weights from the distributions and perform multiple forward passes using the same input.
In doing so, we can then compute a mean prediction for the final output and a variance over all predictions for an estimation of the uncertainty.

## Slide 19: Bayesian Neural Network (BNN)
Doubles the number of parameters ;
Potentially requires more training data;
Makes inference longer (multiple forward passes necessary);
Gives information on the confidence level of the model !
Generally, distributions on weights are approximated by gaussian distributions
Bayesian formulation – Motivation and Introduction
Notes: In practice we could use any probability distribution to reparameterize our weights, but usually the gaussian distribution is used and we also used this one.
Since it is characterized by only two parameters, we essentially learn these parameters during training, so it doubles the number of parameters.
This can make the training harder which can call for more training data.
Also, it makes the inference longer because the higher number of parameters makes a single pass longer and we also need to do multiple passes if we want the uncertainty information.
However, having said all that, the main selling point of the Bayesian formulation is the information on the confidence level that it gives. So if this information is useful enough, it can compensate for the associated drawbacks.

## Slide 20: BNN implementation
Raphaël Brodeur
https://github.com/torchbayesian/torchbayesian
GitHub repository
Deterministic-to-Bayesian conversion of any existing model.
Compatible with any nn.Module – Each weight is reparameterized as a variational distribution.
Direct KL divergence access using analytic computation when available and Monte Carlo estimation otherwise.
PyTorch-native design – Preserves existing workflows and is fully compatible with other PyTorch-based tools such as MONAI.
Bayesian formulation – Experimentation
Notes: Now for the implementation of the bayesian network, I used the torchbayesian framework that was coded by my collegue Raphael Brodeur.
If what I’m talking about remotly interests you, I highly recommend you check out the repository by scanning the qr code or by searching torchbayesian.
To do a very quick selling pitch, with this framework you can convert any model coded using pytorch to bayesian by using one line of code.
It basically reparameterize every nn.parameter object to a probability distribution and handles automatically all other considerations such as regularization.
It’s great, I recommend it, so now back to what I was saying.

## Slide 21: Deterministic vs Bayesian Training
Training
Training
Pre-trained model
Synthetic dataset
Training set
Validation set
Pre-training
Fine-tuning
Experimental dataset
Training set
Validation set
Pre-trained model
Test set
Training
Training
Deterministic formulation
Bayesian formulation
Evaluation
Evaluation
Model analysis and comparison
Uncertainty analysis
Bayesian formulation – Experimentation
Notes: So using the torchbayesian module, we could train exactly the same way the model I presented before and this same model but using the Bayesian formulation.
Once the training were performed, we could compare the performance of the two models, and in the case of the BNN, we could do an uncertainty analysis to see if the uncertainty correlates with the errors in the reconstruction.

## Slide 22: Results

## Slide 23: Qualitative correlation
Results – Uncertainty Analysis
Notes: First looking at the uncertainty analysis, I show here qualitative results, so some slices of the reconstructions along with their associated TPS calculation.
If we zoom in on the uncertainty maps and gamma maps…

## Slide 24: Qualitative correlation
Results – Uncertainty Analysis
Notes: - We see from the regions that I highlighted that very high uncertainty seems to correlate with error.

## Slide 25: Transverse profiles
?
?
Results – Uncertainty Analysis
Notes: Now looking at transverse profiles, we see that on these three profiles, most notably on this one, a large error also means way larger uncertainty bands, which suggest some correlation between error and uncertainty.
However, on those two for example, we see that error do not systematically correlate with uncertainty, at least not qualitatively.

## Slide 26: Depth dose curves
Results – Uncertainty Analysis
Notes: Looking at PDDs, the top one is a typical example with quite small uncertainty band and error.
But on the bottom, we can see this region where the error is large and the uncertainty becomes larger exactly at that point in the profile, which is reassuring to see.

## Slide 27: Quantitative correlation
Pearson and Spearman correlation coefficients on the complete test dataset
Results – Uncertainty Analysis
Notes: For the quantitative correlation, I’ve binned by uncertainty percentile in each reconstruction in the test dataset.
That way I could calculate the mean absolute error and the mean gamma index in those bins, and plot them to see if the error seems to increase with uncertainty, which is the case, at least for the three cases presented previously.
Pearson and Spearman correlation coefficients were also computed on the whole test dataset, which revealed a very high correlation of above 0.9 for mean absolute error, and a pretty good correlation of above 0.6 for mean gamma index.
This proves that uncertainty indeed correlates with error, which is what we wanted from this Bayesian formulation.

## Slide 28: Deterministic model vs BNN performance
Performance difference not statistically significant !
Results – Deterministic vs Bayesian models
Notes: Now for comparison between the deterministic model and the BNN, three metrics were used, namely the gamma pass rate, the structural similarity index measure and the root-mean-square error.
We observe in those results that the values are very similar between the models, where on some metrics the deterministic model while one others, the BNN is better.
================================================================
By calculating a p-value over the test dataset, a p-value of about 0.8 was obtain, which means that the performance difference was shown to be not statistically significant.
This then means that the BNN gives information on the uncertainty without deteriorating the model
Info sur l’incertitude sans déteriorer le modèle

## Slide 29: Discussion and Conclusion

## Slide 30: Reconstruction time
Deterministic formulation
Bayesian formulation
x
. . .
(single pass)
(single pass)
(multiple passes)
(in our case)
Discussion and limitations
Notes: The first limitation of the bayesian formulation is about reconstruction time.
As previously said, in the deterministic case, a single pass is needed, which in our case takes about 15 ms.
Because the number of parameters is doubled in the bayesian case, the forward pass time was almost doubled, which was expected.
====================================================================
However, we also need to take into account that if we want the uncertainty information, we need to make multiple forward passes, and the time needs to be multiplied by this number of forward passes to obtain the final inference time.
This then amounts to just less than a second, which is still extremely fast.
Using other iterative methods, we still couldn’t even dream about inference that fast, so I would say that we are still in the near real-time regime.

## Slide 31: Uncertainty calibration
Gawlikowski et al. Artif Intell Rev 56. 2013
Kuleshov et al. Int conf on machine learning. 2018
Discussion and limitations
Calibration is needed to ensure that the derived predictive confidence represents a good approximation of the actual probability of correctness.
Notes: I also want to address the uncertainty calibration.
Right now the uncertainty gives us relative confidence information, but if we want to take QA decision based on this uncertainty, which we do, we would want the predicted confidence to represent well the actual probability of correctness.
There are a few ways to do this, but we plan on using conformal prediction, to calibrate the uncertainty in the future.
I still want o emphasize this is not done yet, and it will be done in future work.

## Slide 32: Conclusion
Deep learning reconstruction of a scintillation volume was shown to be a promising approach for real-time 3D dosimetry.
However, classical DL gives prediction without confidence level.
Bayesian neural networks can help provide uncertainty information on the predictions.
Take-home message
The BNN reconstruction framework showed that uncertainty correlates with error.
BNNs allow for confidence prediction without trade-off on performance.
Conclusion
Notes: In conclusion, the deep learning reconstruction of a scintillation volume was shown to be a promising approach for real-time 3D dosimetry.
However, classical deep learning models do not give prediction on their confidence level and can be prone to hallucinations, and BNNs can help mitigate this limitation of neural networks.
But the real take-home message of this presentation is that the BNN framework proposed here showed that uncertainty correlates with error.
We’ve also shown that the BNN formulation allowed for confidence prediction without trade-off on performance, at least in our specific context.
This is why I would encourage any of you to try this approach to see if these results can generalize to other contexts.

## Slide 33: Next steps
Uncertainty calibration and use:
  How can the uncertainty be used ?
Optimize learning methods:
  Test other architectures;
  Add physics priors in the training (gradients, spatial continuity).
Improve training data:
  Implement a new experimental setup providing more information;
  Use of optical design software for synthetic data generation.
Conclusion
Notes: For the next steps, the first thing to do will be to calibrate the uncertainty and to determine how can we use this uncertainty in real scenarios. Can we actually use this information to make decisions ?
We also want to optimize learning methods by testing other architecture, as this discipline advances so fast it’s hard to keep up. So it would be nice to test other models and see how they compare.
We would also want to add physics priors in the training to penalize reconstruction that are not physically consistent. For examples, dose gradients steeper than the theoretical limit.
Finally, deep learning models are as good as their training data, so we would want to implement a new experimental setups that could give us more information on the deposited dose, which should help greatly the reconstruction.
We could also use optical design software to simulate the physics of optical photons to generate synthetic data that better ressembles the experimental data.

## Slide 34
Medical Physics Research Group, Université Laval, 2026
Uncertainty-Aware Deep Learning Reconstruction for 3D Scintillation Dosimetry
Alexis Horik1,2,3, Raphaël Brodeur1,3, Nicolas Drouin1,3, Simon Thibault1,2, and Louis Archambault1,2,3
1Département de physique, de génie physique et d’optique, Université Laval, Québec, Canada
2Centre d’optique, photonique et laser, Université Laval, Québec, Canada
3Centre de recherche du CHU de Québec-Université Laval et Centre de recherche sur le cancer de l’Université Laval, Université Laval, Québec, Canada
alexis.horik.1@ulaval.ca