## Slide 1
Uncertainty-Aware Deep Learning Reconstruction for 3D Scintillation Dosimetry
Alexis Horik1,2,3, Raphaël Brodeur1,3, Nicolas Drouin1,3, Simon Thibault1,2, and Louis Archambault1,2,3
1Département de physique, de génie physique et d’optique, Université Laval, Québec, Canada
2Centre d’optique, photonique et laser, Université Laval, Québec, Canada
3Centre de recherche du CHU de Québec-Université Laval et Centre de recherche sur le cancer de l’Université Laval, Université Laval, Québec, Canada
International Conference on 3D Dosimetry 2026

## Slide 2: Presentation outline
Motivation and proposed approach
Overview of the 3D dosimeter
  Experimental setup;
  Deep learning reconstruction model.
Bayesian formulation
  Motivation and introduction;
  Bayesian neural network for tomographic reconstruction.
Results
  Uncertainty analysis;
  Bayesian vs deterministic models.
Discussion and conclusion

## Slide 3: Toward real-time, uncertainty-aware 3D dosimetry
3D dose measurements are needed for satisfactory linac QA.
Real-time dosimetry is desirable
The reliability of dose measurements is crucial.
Dynamic treatments benefit from temporal information
Technicians’ and physicists’ time is valuable
Develop a novel dosimeter that can (i) measure the dose in 3D and in real-time, and (ii) estimate the reliability of its measurement.
Objective
Motivation and Context

## Slide 4: Proposed approach
Dose measurement
Tomographic reconstruction
3D dose distribution estimation
QA metrics and decision-making
Dose deposition in scintillator
1
2
3
4
Scintillation light emitted
Optical system records projections
Input: Projection data
Tomographic reconstruction algorithm
Output: 3D distribution
Development of acceptance criteria.
Spatial information on the errors.
Flagging of reconstruction errors.
3D voxelized dose map.
Uncertainty estimation of the reconstruction.
Uncertainty map
Reconstructed dose
Example use:
Thresholding on uncertainty to determine if:
  The reconstruction is reliable.
If the reconstruction is reliable:
  Flag high-error regions;
  Guide targeted review of profiles or PDDs;
  Global acceptance criteria.
Motivation and Context

## Slide 5: Dose Measurement

## Slide 6: Volumetric scintillation dosimetry
Source: Eljen Technology
Scintillation Absorbed dose;
Fast mechanism ( s);
Water-equivalent;
Dose rate independent;
Mostly energy independent.
Dosimetric Characteristics
The problem becomes…
Imaging of the scintillation light;
Tomographic reconstruction.
Motivation and Context

## Slide 7: Experimental constraints
Source: Siemens Healthineers
Geometric constraints – Restricted space around the treatment isocenter.
Camera placement considerations – Radiation can damage cameras and gantry motion can obstruct optical paths.
Time – Physicists have limited time to perform measurements.
Setup complexity – Intricate setups can suffer from reproducibility problems.
Cost – Multiple cameras or recent technology (ex. plenoptic cameras) can bump up the cost of the device.
Motivation and Context

## Slide 8: Experimental setup
Scintillating
Cube
Polyvinyltoluene
Plane
Mirrors
CCD
Camera
Proposed 3D dosimeter – Experimental data
Notes: Having that all in mind, we build this experimental setup.
It is composed of a 10x10x10 𝑐𝑚^3 scintillator, coupled with two plane mirrors placed at 45 degrees angle so that we can aquire three orthogonal views from a single image.
The CCD camera is placed such that the field of view includes the three views in the frame

## Slide 9: Experimental data collection and curation
Background noise substraction ;
Raw image
Processed image
Intensity (16-bit scale) [ - ]
Pixel [ - ]
Nicolas Drouin
Temporal median filter.
Collect of 249 experimental data of the form:
Proposed 3D dosimeter – Experimental data

## Slide 10: Tomographic Reconstruction

## Slide 11: Tomographic reconstruction of scintillation light
In our case, the problem is not simple…
High-dimensional inverse problem: voxels reconstructed from only 3 projections.
Underdetermined problem: many 3D dose distributions may be consistent with the same projection data.
Iterative methods are time-consuming.
High representation capacity
Fast reconstruction after training
Reconstruction from raw light signal.
Deep learning reconstruction
Proposed 3D dosimeter – Reconstruction
Beister et al. Physica Medica. 2012

## Slide 12: Deep learning reconstruction model
Transformer
Encoder
3D CNN
Decoder
CNN
Embedding
Latent
Representation
Input
projections
Output
distribution
Encoder-Decoder framework
Proposed 3D dosimeter – Reconstruction

## Slide 13: Architecture
Proposed 3D dosimeter – Reconstruction

## Slide 14: Training data
Proposed 3D dosimeter – Training data
Synthetic data
Experimental data
Synthetic dataset
Validation set
Training set
Experimental dataset
Validation set
Test set
Training set
75%
25%
~70%
~18%
~12%
Data generated by ray-tracing in python.
Gaussian filter to approximate lateral dose falloff.
Empirical equation for depth dose deposition.
Data acquired using the experimental setup.

## Slide 15: Bayesian Formulation

## Slide 16: Limitations of DL – Prediction without confidence
?
Hard to interpret (black box);
Limited view reconstruction remains ill-posed the model can “hallucinate”;
Produces a single dose estimate without confidence level we must “blindly” trust the model.
Limitations
Bayesian neural network
Predicts both dose and uncertainty;
Identifies regions with lower reconstruction confidence.
Bayesian formulation – Motivation and Introduction

## Slide 17: Bayesian Neural Network (BNN)
Deterministic model
is fixed after training
a given input always gives the same output
Blundell et al. Int conf on machine learning. 2015
Bayesian formulation – Motivation and Introduction

## Slide 18: Bayesian Neural Network (BNN)
Deterministic model
Bayesian model
is fixed after training
a given input always gives the same output
In practice:
Blundell et al. Int conf on machine learning. 2015
Bayesian formulation – Motivation and Introduction

## Slide 19: Bayesian Neural Network (BNN)
Doubles the number of parameters ;
Potentially requires more training data;
Makes inference longer (multiple forward passes necessary);
Gives information on the confidence level of the model !
Generally, distributions on weights are approximated by gaussian distributions
Bayesian formulation – Motivation and Introduction

## Slide 20: BNN implementation
Raphaël Brodeur
https://github.com/torchbayesian/torchbayesian
GitHub repository
Deterministic-to-Bayesian conversion of any existing model.
Compatible with any nn.Module – Each weight is reparameterized as a variational distribution.
Direct KL divergence access using analytic computation when available and Monte Carlo estimation otherwise.
PyTorch-native design – Preserves existing workflows and is fully compatible with other PyTorch-based tools such as MONAI.
Bayesian formulation – Experimentation

## Slide 21: Deterministic vs Bayesian Training
Training
Training
Pre-trained model
Synthetic dataset
Training set
Validation set
Pre-training
Fine-tuning
Experimental dataset
Training set
Validation set
Pre-trained model
Test set
Training
Training
Deterministic formulation
Bayesian formulation
Evaluation
Evaluation
Model analysis and comparison
Uncertainty analysis
Bayesian formulation – Experimentation

## Slide 22: Results

## Slide 23: Qualitative correlation
Results – Uncertainty Analysis

## Slide 24: Qualitative correlation
Results – Uncertainty Analysis

## Slide 25: Transverse profiles
?
?
Results – Uncertainty Analysis

## Slide 26: Depth dose curves
Results – Uncertainty Analysis

## Slide 27: Quantitative correlation
Pearson and Spearman correlation coefficients on the complete test dataset
Results – Uncertainty Analysis

## Slide 28: Deterministic model vs BNN performance
Performance difference not statistically significant !
Results – Deterministic vs Bayesian models

## Slide 29: Discussion and Conclusion

## Slide 30: Reconstruction time
Deterministic formulation
Bayesian formulation
x
. . .
(single pass)
(single pass)
(multiple passes)
(in our case)
Discussion and limitations

## Slide 31: Uncertainty calibration
Gawlikowski et al. Artif Intell Rev 56. 2013
Kuleshov et al. Int conf on machine learning. 2018
Discussion and limitations
Calibration is needed to ensure that the derived predictive confidence represents a good approximation of the actual probability of correctness.

## Slide 32: Conclusion
Deep learning reconstruction of a scintillation volume was shown to be a promising approach for real-time 3D dosimetry.
However, classical DL gives prediction without confidence level.
Bayesian neural networks can help provide uncertainty information on the predictions.
Take-home message
The BNN reconstruction framework showed that uncertainty correlates with error.
BNNs allow for confidence prediction without trade-off on performance.
Conclusion

## Slide 33: Next steps
Uncertainty calibration and use:
  How can the uncertainty be used ?
Optimize learning methods:
  Test other architectures;
  Add physics priors in the training (gradients, spatial continuity).
Improve training data:
  Implement a new experimental setup providing more information;
  Use of optical design software for synthetic data generation.
Conclusion

## Slide 34
Medical Physics Research Group, Université Laval, 2026
Uncertainty-Aware Deep Learning Reconstruction for 3D Scintillation Dosimetry
Alexis Horik1,2,3, Raphaël Brodeur1,3, Nicolas Drouin1,3, Simon Thibault1,2, and Louis Archambault1,2,3
1Département de physique, de génie physique et d’optique, Université Laval, Québec, Canada
2Centre d’optique, photonique et laser, Université Laval, Québec, Canada
3Centre de recherche du CHU de Québec-Université Laval et Centre de recherche sur le cancer de l’Université Laval, Université Laval, Québec, Canada
alexis.horik.1@ulaval.ca