## Slide 1: Measuring Small Collimated Proton Fields with a Volumetric Scintillation Detector
A history of innovation leading to a practical application
Abram Bell, Jason Holmes, Sam Beddar, Daniel Robertson
IC3DDose 2026
15 June 2026

## Slide 2: Volumetric scintillation dosimetry
Notes: I coined this term during my PhD to describe a 3D detector that is read in a way that produces <3D data (e.g. 2D projections). Might also call it “2.5-D dosimetry”.
Let’s look at the history and development of this idea over the years that led to the current work.

## Slide 3
Notes: Initial concept of 3D scintillation dosimetry
Eye plaques (effectively infinite time scale of brachytherapy source)
Rotating stage and single camera

## Slide 4
Notes: First application of scintillator volumes to proton beams
Range measurement application

## Slide 5
Notes: First application of volumetric scintillation dosimetry in external beam therapy
Compared dose to TPS

## Slide 6
Notes: Concept of quenching addressed
Quenching modeled, potential to correct for quenching raised

## Slide 7
Notes: Machine QA
Range
Spot position
Relative intensity

## Slide 9
Notes: First scintillator 3D reconstruction of a radiation beam (ELBE electron accelerator)
Also measured a proton Bragg peak, but didn’t apply quenching correction

## Slide 10
Notes: Detailed analysis of optical artifacts for volumetric scintillation dosimetry
Their effects
How to correct for them
2%,2mm Gamma comparison

## Slide 11
Notes: Limited-angle reconstruction of proton pencil beams is possible with only 3 views if one of them is a beam’s-eye view
Individual beams can be summed to image patient treatment plans
Single energy layer of an IMPT plan
Nominal plan
Isocenter shifted by 1 mm in X and Y

## Slide 12
Notes: Built the 3-camera system
Range
Spot position
Lateral profile

## Slide 13
Notes: Triggering scheme to align camera acquisitions with beam pulses

## Slide 14
Notes: Important contribution
LET curve and Bragg peak are sharp – it is hard to accurately align them
Use computed Bragg peak and LET curve to calculate quenching correction factor – this has a less steep gradient
  Small misalignments don’t cause such problematic impact on corrected light distribution

## Slide 16
Notes: improved the application of Birks’ model quenching corrections
in 3D scintillators by numerically projecting the dose and LET 3D grid to camera projections

## Slide 17
Notes: Recent work of a colleague I mentored during his PhD – applying VSD to photon beams
Work on photon SRS multi-met fields for his PhD

## Slide 18: Proton stereotactic radiosurgery (srs)

## Slide 19: Proton Stereotactic Radiosurgery (SRS)
Overview
High dose (10-20 Gy)
Field-specific brass aperture
Small target volume
Optional range shifter

## Slide 21: Proton Stereotactic Radiosurgery (SRS)
Patient-Specific Quality Assurance
Demands high spatial resolution
Steep dose gradients, possible non-homogeneity

## Slide 22: Patient-specific qa for proton srs
2D lateral profile
MatriXX ionization chamber array
  7.5 mm chamber spacing
  4.5 mm diameter chambers
  32x32 grid

## Slide 23: Patient-specific qa for proton srs
PDD/Absolute Dose
PTW microdiamond detector
  0.004 mm³ sensitive volume
  Minimal direction/energy dependence
1D scanning water tank
  Limited number of depths/measurements

## Slide 24: Volumetric scintillation dosimetry for Proton srs

## Slide 25: Volumetric Scintillator Detector
Overview
Single lateral projection
~0.1 mm spatial resolution
15 cm x 15cm x 15cm cube
Projection correction omitted due to small field size
PICTURE HERE
Notes: Normally we would need to correct every pencil beam to apply the correct pixel size at its distance from the camera. This was omitted because of the small field size, which makes this correction small.

## Slide 26: Experimental Setup Exploring Detector Viability in Performing SRS PSQA
Phase 1 – Determine optimal quenching model
Implement optical corrections
Birk's model – use Monte-Carlo (MC) dose and linear energy transfer (LET) to predict light
Kelleter's model – use scintillation light to predict total dose

## Slide 27: Birks Model Predicting Quenched light from Dose
Empirical model proposed in 1951 Utilizes track-average LET, A, kB to calculate scintillation light
Variables
  A: Scaling factor
  dE / dx: Localized energy lost, or LET
  kB: Quenching parameter
Voxelized form derived by Dr. Robertson
Volumetric Scintillation Dosimetry for Scanned Proton Beams Robertson, 2014

## Slide 28: Kelleter's Method Predicting Dose from light data
Combines Birk's empirical model and Bortfeld's analytical approximation
Predicts integrated depth dose from light data
  Limited to IDD, rather than voxel data
Quenched Bragg curve for proton therapy Kelleter et. al., 2014

## Slide 29: Phase 1 – Determine optimal Data processing method
Birks Method
Kelleter Method
MC Dose
MC LET
Birks' Formula
Quenched Light
Measured Light
Gamma Analysis
Measured Light
Kelleter's Formula
Predicted Dose
MC Dose
Gamma Analysis

## Slide 30: Model Comparison
Notes: Caution that this was early work, Kelleter could still be potentially useful

## Slide 31: Quenching correction method
Birks Method
Kelleter Method
MC Dose
MC LET
Birks' Formula
Quenched Light
Measured Light
Gamma Analysis
Measured Light
Kelleter's Formula
Predicted Dose
MC Dose
Gamma Analysis
Notes: Note that Kelleter limited to 1D correction. Birks chosen. Show change to dose comparison instead of light

## Slide 32: Quenching correction method
Birks Method
Birks Quenching Correction
MC Dose
MC LET
Birks' Formula
Quenched Light
Measured Light
Gamma Analysis
MC Dose
MC LET
Birks' Formula
Quenched Light
Measured Light
Measured Dose
Quenching Correction Factors
Notes: Note that Kelleter limited to 1D correction. Birks chosen. Show change to dose comparison instead of light

## Slide 33: Testing Detector PSQA performance
Introduce errors in patient plan beam deliveries
Range shift (1, 3 mm)
No aperture
90 Percent MU
Deleted Spot
Compare Measured and Calculated Dose
Gamma analysis (2% 2mm)

## Slide 34: Proton srs patient field
Five-field SRS treatment of an arteriovenous malformation
Patient-specific brass apertures mounted on the SRS nozzle
Dose distribution within the aperture was modulated laterally and in depth
Maximum lateral extent of the aperture < 2 cm

## Slide 35: T0G125 – Control

## Slide 36: T0G125 – Control

## Slide 37: T180G153 – Control

## Slide 38: T270G90 – Control

## Slide 39: T180G60 – Control

## Slide 40: T331G95 – Control

## Slide 41: T0G125 – 1mm range shift

## Slide 42: T331G95 – 1mm range shift

## Slide 43: T0G125 – 3mm range shift

## Slide 44: T331G95 – 3mm range shift

## Slide 45: T0G125 – 90% MU

## Slide 46: T331G95 – 90% MU

## Slide 47: T0G125 – no aperture

## Slide 48: T331G95 – no aperture

## Slide 49: Meningioma Plan - 2%2mm Gamma Analysis
Field | Control | 1mm RS | 3mm RS | No Aperture | 90% MU
T331 G95 | 99.8 | 96.5 | 80.7 | 19.9 | 81.5
T0 G125 | 98.3 | 90.1 | 82.6 | 12.4 | 81.0
T180 G153 | 99.0 |  |  |  | 
T180 G60 | 91.3 |  |  |  | 
T270 G90 | 99.1 |  |  |  | 

## Slide 50: Conclusion
Volumetric scintillation dosimetry shows promise for PSQA of small, collimated proton fields
  High spatial resolution
  Entire field measured in a single delivery
  Ability to detect plan/delivery errors
Some further optimization required before clinical use

## Slide 51: acknowledgments
MD Anderson Cancer Center
Sam Beddar
Beckett Hui
Chinmay Darne
Fahed Alsanea
Mayo Clinic
Jason Holmes
Jason Shen
Colleagues
Students
Abram Bell
Kyle Durrant

## Slide 52
Notes: A group we collaborate with developed this system based on plenoptic camera and flat panel imager
3D reconstruction