## Slide 1: Development of a 3D Electronic Array for Radiotherapy Quality Assurance
Dante E. Roa, PhD

## Slide 3: Prototype v.3.0
QR-Code
Linearity
Dose Rate

## Slide 4: Prototype v.4.5 – Active Matrix (AM) Design
Sensitive Area – 18 x 18 cm2
KEMET Capacitor
MOSFET
Pixel = Diode+Capacitor+MOSFET

## Slide 5: Prototype v.4.5

## Slide 6
Prototype v.4.5 – Service (Diagnostics) Software
Configurable data acquisition
Horizontal and vertical profile visualization (off-axis)
Pixel signal analysis
Visualization as a function of time (instantaneous and cumulative)

## Slide 7
Prototype v.4.5 – Web-based User Software

## Slide 8
Prototype v.4.5 – Characterization Factor - SETUP
Each setup is designed to determine the correction factors for each pixel in the active matrix, thereby homogenizing the detector response.
Linac
SDD = 111.5 cm

## Slide 9
Prototype v.4.5 – Characterization Factor
Correction factors () are generated to normalize all pixels response:
Where:
: Integrated signal of pixel (i,j) in matrix k during the 22x22 field acquisition.
: Integrated signal of the central pixel in matrix k during the 22x22 field acquisition.
: Monte Carlo correction factor for inter-matrix response normalization.
Each pixel measurement is multiplied by its corresponding correction factor.
Greer PB, Popescu CC. Amorphous silicon EPID calibration for dosimetric applications: comparison of a method based on Monte Carlo prediction of response with existing techniques. Medical Physics, 2006.

## Slide 10
Prototype v.4.5 – Characterization Factor – DATA/RESULTS
Elekta Infinity 10×10 cm² 50 MU 300 MU/min 6 MV
Application of correction factors

## Slide 11
Prototype v.4.5 – Angular Dependence – SETUP Measurements
Elekta Infinity:
6 MV,
22×22 cm²,
50 MU,
300 MU/min

## Slide 12
Prototype v.4.5 – Angular Dependence – DATA/RESULTS

## Slide 13
Prototype v.4.5 – Angular Dependence – continues

## Slide 14
Prototype v.4.5 – Angular Dependence – DATA/RESULTS
Elekta Infinity:
6 MV,
22×22 cm²,
50 MU,
300 MU/min

## Slide 15
Prototype v.4.5 – Angular Dependence – DATA/RESULTS
Elekta Infinity:
6 MV,
22×22 cm²,
50 MU,
300 MU/min

## Slide 16
Prototype v.4.5 – Beam Pulse Quality – DATA/RESULTS
Pixel signal increases with higher dose rate.
Specific beam pulse patterns can be identified.
Beam ramp-up and de-tuning can be measured.
Varian Trilogy
Elekta Infinity

## Slide 17
Prototype v.4.5 – Beam Pulse Quality – DATA/RESULTS #2
Interlocks, slight dose rate differences, and other anomalies can be detected.
Interlock

## Slide 18
Prototype v.4.5 – PDD & Profile Results
E = 6 MV,
FS = 10 x 10 cm

## Slide 19
Prototype v.X.X – Towards an Advanced Clinical Model

## Slide 20
Prototype v.X.X – High Spatial Resolution AM Schematic
Sensitive Area – 42 x 42 cm2

## Slide 21
Prototype v.X.X – Vertical High Spatial Resolution AM Schematic
5 mm spatial resolution
5 mm spatial resolution

## Slide 22
Prototype v.X.X – High Spatial Resolution AM Schematic

## Slide 23
Prototype v.X.X – MC Data
1.5 cm - 1.8% diff
6.0 cm – 3.6% diff
10.0 cm – 1.7% diff
14.0 cm – 3.3% diff
22.0 cm – 3.1% diff
30.0 cm – 2.3% diff

## Slide 24: Water Scans versus 3D Detector Commissioning Data Acquisition Times
All Commissioning Data could be acquired in days

## Slide 25: 3D Electronic Detector Array Modular Design – Three different applications in ONE
Full Assembly – 20 layers
Linac commissioning
Monthly QA
Patient-specific plan QA, if needed
Partial Assembly – 14 layers
Patient-specific plan QA with 3D data acquisition
Basic Assembly – 1 layer
Daily linac QA
~ 188 lbs (85 Kg)
~ 74 lbs (34 Kg)
~ 32 lbs (15 Kg)

## Slide 26
$100K - $120K*
$50K - $70K*
$10K - $20K*
$70K - $100K
*Separate purchase
Daily Linac QA
Patient Specific QA
Commissioning & Monthly
~$43K*

## Slide 27
Conclusion
Measured and reliable 3D dosimetric data with an electronic detector array is possible.
Linearity, dose rate, PDD, and profile results support this statement
The detector has sufficient temporal resolution to evaluate beam pulse quality from different linac vendors. Useful in beam diagnostics and an alert to the user.
Can provide 3D dose distribution of patient-specific treatment plan QA minutes after irradiation.
Full implementation of this technology could lead to significant time-savings in linac commissioning and monthly QA and cost savings in PSQA and Daily Linac QA.

## Slide 28
The Team
Dante E. Roa, PhD 1,2*, Miguel Risco-Castillo1,4, Renzo Romero1,4. Enzo Aucca1,3, Renzo Ocampo1,4, William de Souza Santos, PhD6, Mirko Salomón Alva Sánchez, PhD7, Carmen Sandra Guzmán Calcina, PhD1,5, Modesto Montoya, PhD4, Roger Challco, EE1,4, Gabriel López1,3, Erick Paniagua, EE4,8
1 BrightDOSE USA LLC, Mission Viejo, CA, USA
2 Leonard Cancer Institute, Mission Viejo, CA, USA
3 Ponti icia Universidad Católica del Perú , Electronic Engineering Department, Lima, Perú
4 Universidad Nacional de Ingenierı́a, Facultad de Ciencias, Lima, Perú
5 Instituto de Investigaciones en Ciencias Biomédicas, Universidad Ricardo Palma, Lima, Perú
6 Departamento de Fı́sica, Universidade Federal de Sergipe, São Cristóvão, Brasil
7 Federal University of Health Sciences of Porto Alegre, Porto Alegre, Brazil
8 Universidad Nacional de San Antonio Abad del Cusco, Electronic Engineering Department, Cusco, Perú
*E-mail: danteeroa@gmail.com

## Slide 29: Questions?
Visible Light Tests with model v.4.5