## Slide 1
Development of a 3D Slicer Extension for MR-Based Gel Dosimetry Analysis
G S Lee1, K M Alexander1,2,3, and JD Jutras1,4
1 Department of Physics and Astronomy, University of Calgary, Calgary, Alberta, Canada
2 Department of Oncology, Queen’s University, Kingston, Ontario, Canada
3 Medical Physics Department, Cancer Centre of Southeastern Ontario at Kingston Health Sciences Centre, Kingston, Ontario, Canada
4 Department of Medical Physics, Arthur JE Child Comprehensive Cancer Centre, Calgary, Alberta, Canada

## Slide 2
The Return of MRI in Gel Dosimetry
1. Development of
Fricke Dosimetry
2. MR Readout of
Fricke Dosimetry
3. Adoption of
Optical CT Readout
4. Renewed Interest
in MR Readout
Chemical dosimetry based on the oxidation of Fe2+ to Fe3+ by ionizing radiation.
MRI was the preferred readout method for dose quantification.
Optical CT became more accessible and practical, leading to widespread adoption and extensive Fricke gel studies.
Advances in MRI technology and MR-guided radiation therapy have renewed interest in MR-based Fricke gel dosimetry.
No open-source software exists for MR gel dosimetry analysis!
1

## Slide 3
Existing Optical CT Workflow in 3D Slicer
Alexander KM, et al. Streamlined open-source gel dosimetry analysis in 3D slicer. Biomed Phys Eng Express. 2018;4(4). doi: 10.1088/2057-1976/aad0cf.
2

## Slide 4
What’s New in MR Workflow?
R1 or R2 Maps
  Pre- and post-irradiation volume registration
  Denoising
  Pre- and post-irradiation volume subtraction
MR Dose Calibration
  User defined sampling axis
  Adjustable averaging radius
3

## Slide 5
Experimental Case Study
Fricke Gel Formulation
Gelatin: 6 wt%
Sulfuric acid: 65 mM
Ferrous ammonium sulfate: 0.42 mM
Xylenol Orange: 0.05 mM
Two gel jars prepared:
  Measured
  Calibration
Clinical Plan
Liver SBRT
Varian TrueBeam linac
10 MV FFF beam
3 Gy scaled down prescription dose
Cylindrical phantom
4

## Slide 6
Imaging Readout Methods for FXO Gel
Magnetic Resonance Imaging (MRI)
3 T Siemens MAGNETOM Vida scanner
Measures R1 relaxation rates via MP2RAGE pulse sequence
Reconstructed at a resolution of 0.50.50.5 mm3
Modus QA Vista 16 Optical CT Scanner with 590 nm light source
Measures optical attenuation
Reconstructed at a resolution of 0.50.50.5 mm3
Optical Computed Tomography (Optical CT)
5

## Slide 7
MR Workflow Demonstration
Load data
ΔR1 mapping: registration, denoising, subtraction
2. Registration
3. Dose calibration
4. Dose comparison
Planning → IGRT & Measured gel volume → IGRT
PDD-based calibration: ΔR1 → dose
Gamma analysis & line profiles
6

## Slide 8
1.1 Load DICOM Data
7

## Slide 9
1.2 Load non-DICOM Data
8

## Slide 10
ΔR1 Mapping Workflow
Registration
9

## Slide 11
ΔR1 Mapping Workflow
A. Registration – Manual Transformation & Resampling
10

## Slide 12
A. Registration – Apply Transform to R1 Map
ΔR1 Mapping Workflow
11

## Slide 13
ΔR1 Mapping Workflow
Registration
12

## Slide 14
B. Denoising
Before Denoising
After Denoising
ΔR1 Mapping Workflow
13

## Slide 15
C. Volume Subtraction
ΔR1 Mapping Workflow
ΔR1 mapping is performed for both measured and calibration gel volumes!
14

## Slide 16
2. Registration
15

## Slide 17
2.1 Registration
16

## Slide 18
2.2 Registration
17

## Slide 19
2.2 Registration
18

## Slide 20
2.2 Registration
18

## Slide 21
2.2 Registration
19

## Slide 22
Determines the relationship between ΔR1 and dose
Uses a PDD-based calibration method along a user-defined axis
Aligns the ΔR1 depth profile with commissioning PDD data
3. Dose Calibration
20

## Slide 23
3.1 Dose Calibration
21

## Slide 24
3.1 Dose Calibration
22

## Slide 25
3.2 Apply Calibration
23

## Slide 26
4. Dose Comparison
24

## Slide 27
Line Profile
25

## Slide 28
Results
MR Dose Map
Gamma Volume
Line Profile
26
Imaging Modality | Mean Dose inside 95% Isodose Line |  |  | Gamma Analysis at 2%/2mm (%)
 | Planned Dose (Gy) | Measured Volume (Gy) | % Difference (%) | 
MR | 3.22 | 3.13 | 2.8 | 96.05
oCT | 3.22 | 3.16 | 1.9 | 97.88

## Slide 29
Developed an MR-based gel dosimetry workflow in 3D Slicer
MR measurements showed good agreement with planned dose and optical CT
Accessible through an open-source platform
Supports MR-guided radiation therapy applications
Conclusion
27

## Slide 30
Acknowledgements
School of Computing, Queen’s University
  Andras Lasso
  Kyle Sunderland
  Gabor Fichtinger
Design Lab
  Adam Yarschenko
  William Gill
28