## Slide 1: Evaluation of a detector design to measure the superficial 2D dose distribution for skin cancer treated by Protontherapy DeCuPro project
Giulia Tosetti1
tosetti@lpccaen.in2p3.fr
Jean-Marc Fontbonne1, Dorothée Lebhertz2, Cyril Moignier2 and Rémi Lafaye1
1Université de Caen Normandie, ENSICAEN, CNRS/IN2P3, LPC Caen UMR6534, F-14000 Caen, France
2 Centre François Baclesse, service de radiophysique, Caen, France

## Slide 2
Without Bolus
With Bolus
Conventional radiotherapy is an available treatment option for NMSC.
The maximum dose deposited by the photon beam is achieved thanks to a tissue equivalent material: the bolus.
The shrinking of the tumor will lead to air gaps between the bolus and the tumor volume, that could impact the total dose given at the tumor [1,2].
I
Tumor
Tumor
Bolus
Introduction
Photon percentage depth dose (PDD) curve without (a) and with bolus (b)
(a)
(b)
Patient’s surface
Patient’s surface

## Slide 3
In the case presented at PTCOG 2023 [3], the patient has been treated with a photon VMAT plan treatment using 3 Gy/fr with a total of 21 fractions (PD=63Gy) but a relapse occurred.
Tumor has evolved during the treatment, creating some air cavities up to 1.2 cm between the skin and bolus.
I
Introduction
Variation of the CTV dose coverage with daily fraction – (a) beginning of treatment (b) middle of treatment (c) end of treatment
(a)
(b)
(c)
Tumor
Bolus

## Slide 4: Motivation
I
In hadron-therapy, particles like protons or ions, deposits most of its energy at the Bragg peak: the dose is delivered precisely at the tumor site. The need for a bolus is eliminated.
Proton percentage depth dose (PDD) curve
Schematics of the patient treatment (TPS) + clinical proton beams from a ProteusOne (IBA)
Pencil beam spot
Proton pencil beam
Tumor
Treatment simulation with TPS

## Slide 5: DOSION
II
DOSION [4] is a parallel plate air ionization chambers consisted in two sub-chambers measuring the beam intensity and two stripped sub-chambers measuring the position and the size of the beam along X and Y axes.
Signal measurement were conducted using the in-house build acquisition system FASTER [5] with a two daughter cards setup for a total of 64 electrometer channels.
Schematics of DOSION
9cm

## Slide 6: Experimental set up
DOSION is position at the center of the gantry aperture and each stripped chamber is connected to an electrometer unit.
Different experiments are carried out:
Linearity
Patient Specific Quality Assurance (PSQA)
II
DOSION under the ProteusOne at Cyclhad, Caen, France

## Slide 7: Data Analysis
To detect each peak, the signal is analyzed using the following steps:
    Background subtraction
Temporal filtering
Spatial filtering
Threshold identification
Peaks detection
II
Signal processing – Signal minus background noise detected by the 15th stripe in (a) X chamber (b) Y chamber, (c) signal filtered in space and time with peak detection
(a)
(b)
(c)
Counts
Signal [au]
Signal [au]
Signal [au]

## Slide 8
Recontruction of the dose maps
II
Gaussian fitting of the time-filtered signals
Quality-control steps applied for peak rejection
(a)
(c)
(b)
Peak rejection criteria - (a) peaks outside the selected area (b) peaks with non-convergent fitting (c) peaks with negative fitted peak height
Channel
Channel
Signal [au]
Signal [au]
O X data
O Y data
X fit
Y fit
O X data
O Y data
X fit
Y fit

## Slide 9
Recontruction of the dose maps
II
2D dose map reconstruction with DOSION
Reconstruction of a single pencil-beam image
Reconstruction of two-dimensional dose maps
Different single pencil beam shoot

## Slide 10: Linearity
To demonstrate the linearity the beam is deliver with different monitor units, ranging from 50 to 5400 MU.
The linear fit residuals have a standard deviation of 0.0025
II
Linear Regression - Fit (a) and Residuals (b)
(a)
(b)
Signal [au/au.max]
Monitor Units
Residuals
Monitor Units

## Slide 11: Patient Specific Quality Assurance
To compare TPS dose calculations with DOSION measurements, the treatment plan was evaluated using both computational and experimental methods.
II
Patient treatment plan on TPS

## Slide 12: Patient Specific Quality Assurance
Dose assessments were performed at three patient depths: (a) at the surface , (b) 7 mm depth and (c) 17 mm depth.
II
X chamber
Y chamber
DOSION
TPS dose calculation lateral view
Water equivalent material
X chamber
Y chamber
DOSION
12cm
9cm
5cm
1cm
Water equivalent material
X chamber
Y chamber
DOSION
(a)
(c)
(b)

## Slide 13: PSQA at 0 mm
a
2D dose map reconstruction with DOSION
TPS dose calculation top view
6.42
6.18
5.88
5.40
4.20
1.50
TPS – DOSION dose map

## Slide 14: PSQA at 7 mm
6.42
6.18
5.88
5.40
4.20
1.50
b
TPS dose calculation top view
TPS – DOSION dose map
2D dose map reconstruction with DOSION

## Slide 15: PSQA at 17 mm
6.42
6.18
5.88
5.40
4.20
1.50
c
TPS dose calculation top view
2D dose map reconstruction with DOSION
TPS – DOSION dose map

## Slide 16: Conclusions and Overlooks
III
From the experiment:
The device response is linear with respect to MU
Ionization chambers provide reliable measurements for PSQA
Future prototype:
An integral ionization chamber will be developed
A camera system couple with a plastic scintillator and a mirror
The chamber will have a larger window
Future prototype design

## Slide 17
THANK YOU VERY MUCH!
Giulia Tosetti1
tosetti@lpccaen.in2p3.fr
Jean-Marc Fontbonne1, Dorothée Lebhertz2, Cyril Moignier2 and Rémi Lafaye1
1Université de Caen Normandie, ENSICAEN, CNRS/IN2P3, LPC Caen UMR6534, F-14000 Caen, France
2 Centre François Baclesse, service de radiophysique, Caen, France

## Slide 18: Bibliography
[1] E. Boman, J. Ojala, M. Rossi, and M. Kapanen, “Monte Carlo investigation on the effect of air gap under bolus in post-mastectomy radiotherapy,” vol. 55, pp. 82–87
[2] M. J. Butson, T. Cheung, P. Yu, and P. Metcalfe, “Effects on skin dose from unwanted air gaps under bolus in photon beam radiotherapy,” Radiation Measurements, vol. 32, no. 3, pp. 201–204, 2000.
[Online]. Available: https://www.sciencedirect.com/science/article/pii/S1350448799002760
[3] Cyril Moigner, et al., “Protontherapy for superficial tumors: validation of the tps dose calculation and comparison to high energy photon techniques”, PTCOG, 2023.
[4] Guillaume Boissonnat, "Chambres d'ionisation en protonthérapie et hadronthérapie", thèse 2015, https://theses.fr/2015CAEN2018
[5] https://faster.in2p3.fr/
III

## Slide 19: Outline
Part I
  Introduction
    Motivation
    Part II
    DOSION
    Experimental set up
    Data Analysis + Experiments
    Part III
    Conclusions & Outlooks
    Bibliography