## Slide 1: FEASIBILITY OF 3D GEL DOSIMETRY FOR CBCT-BASED ISOCENTER VERIFICATION ON ELEKTA SYSTEMS
Methodology
Gel dosimeter application: usage of a 3D polymer gel optimized for CBCT.
Star-shot beam delivery: 2D and 3D patterns on Evo and Versa HD.
Readout: CBCT and MRI for gel dosimetry; EBT4 film for reference.
Calculation: polyGeVero-CT v. 1.3 software.
INTERNATIONAL CONFERENCE ON 3D DOSIMETRY (IC3DDOSE), CAEN, FRANCE, 15-17 JUNE 2026
Key Conclusions
FDK protocol advantage: significantly lower reconstruction noise (∼1.5 HU) vs. Iris HD.
Iris HD limitations: systematic artifacts in homogeneous gels due to DNN scatter estimation.
Readout method agreement: high spatial correspondence (≤ 0.15 mm) between CBCT and MRI.
High spatial accuracy: sub-millimeter correlation (< 0.25 mm) against reference film dosimetry.
Processing procedure efficiency: SNR enhancement via scan averaging and background subtraction.
Figure 1. Experimental setup (A), irradiated polymer gel dosimeter: 2D and 3D star-shot (B) and 2D star-shot irradiation pattern measured with CBCT (C).
Algorithm | Starshot Type | Evo |  |  | Versa HD |  | 
 |  | Rad. vs. Mech. | Rad. vs. Img | Rad. Size | Rad. vs. Mech. | Rad. vs. Img | Rad. Size
M20-Pelvis-HD | 2D | 0.56 | 0.36 | 1.07 |  |  | 
M20-High_Res |  | 0.53 | 0.33 | 1.11 | 1.57 | 1.25 | 0.12
M20-Med_Res |  | 0.49 | 0.29 | 1.07 | 1.42 | 1.14 | 0.12
M20-Pelvis-HD | 3D | 1.69 | 1.21 | 1.30 |  |  | 
M20-High_Res |  | 1.02 | 0.75 | 1.19 | 2.80 | 1.19 | 0.63
M20-Med_Res |  | 1.51 | 1.1 | 1.2 | 2.73 | 1.42 | 0.55
Oskar Madetko1,2, Piotr Maras3,4, Bartłomiej Jasiak2, and Marek Kozicki4,5
1University of Applied Sciences in Tarnow, Tarnow, Poland
2St. Lucas Hospital, Tarnow, Poland
3Copernicus Hospital, Lodz, Poland
4GeVero Co., Lodz, Poland
5Lodz University of Technology, Lodz, Poland
Objectives
Isocenter 3D check: volumetric verification and 3D coincidence of mechanical, imaging, and radiation geometries.
Elekta CBCT feasibility: evaluation of standard on-board imaging protocols for dosimetry.
Reconstruction algorithm analysis: assessment of FDK vs. Iris HD performance in polymer gels.
Table 1. Radiation isocenter analysis using 3D polymer gel dosimetry: coincidence with mechanical and imaging isocenters, and radiation isocenter size. All values: [mm].
Our Department
Our Research Group
A
B
C
MSc Bartłomiej Jasiak
Professor, DSc, PhD Marek Kozicki, Head
PhD Piotr Maras
MSc Oskar Madetko