## Slide 1
Alberto Pullia
INFN - Milano University of Milano
Department of Physics
25th AGATA Week 15-19 September, 2025 GSI Darmstadt, Germany
Digitizers (DIGI-OPT12) Status, Advancements and Production of v3.7.1

## Slide 2
DIGI-OPT12 v3.7.1 cards fully optimized w ADC32J44 chip
 Segment cards v3.7.1 7 realized (7 tested)
 Core cards v3.7.1 8 realized (5 tested)
 Production yield: 83.3%
 Yield after repairs: 100%
No patch needed, fully optimized
Manual patch needed (noisy ch)
v3.7
v3.7.1
First bunch of v3.7.1 cards delivered in March 2025

## Slide 3
Minor problems in v3.7.1 cards
1.3 mm
3 mm
3 mm
 Fix easily implemented
 BUT diagnostics not easy

## Slide 4
Thermal pictures
Segment card v3.6.x
Core card v3.6.x
Core card v3.7.x
P = 1.6 W
Segment card v3.7.x
P = 8.6 W
P = 3.3 W
P = 3.4 W

## Slide 5
Why v3.7.1? Noisy channel of v3.7
Data taken at LNL
by Alain and Andres
22 Jan 2024
v3.7 equipped with AD8030 opamp
v3.6.x equipped with AD8030 opamp
v3.6.x equipped with LTC6247 opamp
MDR cables disconnected
7 MeV range selected
“Ch 10“ on the PCB numbering scheme
Anomalous channel (noisy)

## Slide 6
From v3.7 to v3.7.1
v3.7
v3.7.1
Ch 10
Ch 10
v3.7.1 prototype card delivered in late June 2024
Ch 10
2 mm

## Slide 7
v3.7.1 noise performance – Milano testbench – no detector

## Slide 8
Noise performance – Milano testbench – no detector
DIGI-OPT12 | Opamp | Energy Range | Noise [LSB rms] | ENOB
v3.6.x | AD8030 “standard” | 7 MeV | 2.27 | 11.03
 |  | 20 MeV | 1.43 | 11.69
 | LTC6247 | 7 MeV | 1.63 | 11.50
 |  | 20 MeV | 1.35 | 11.78
v3.7.x | AD8030 “standard” | 7 MeV | 1.73 | 11.42
 |  | 20 MeV | 1.42 | 11.70
 | LTC6247 | 7 MeV | 1.46 | 11.66
 |  | 20 MeV | 1.33 | 11.80
Previous generation
New generation
σ = 1.47
σ = 1.41
<σ>12ch = 1.46 LSB
J
J
12 chs of v3.7.1 card w LTC6247 opamps

## Slide 9
Noise spectral density
New generation
(v3.7.x – AD8030)
kHz
Previous generation
(v3.6.x – AD8030)
Courtesy of D. Bazzacco
without detector
no change vs offset
without detector
with detector
(1)
(2)
(3)
(4)
Offsets
(1) 127 (middle)
(2) 140
(3) 160
(1) 180
kHz

## Slide 10
Fast-reset ON/OFF option
Segment cards may be patched by mounting missing chip
U_NAND_11 chip is NOT installed
in segment digitizer cards v3.6.x- v3.7
U_NAND_11 chip is installed in core digitizer cards
Signal seen at 1st stage output
Signal seen at 2nd stage output
Cosmic ray (~20MeV) Fast reset is active
Cosmic ray (~20MeV)
1.7 MeV event
1.7 MeV event

## Slide 11
Cards patched for on/off option of segment fast-reset

## Slide 12
Summary
New ADC chip ADC32J44 yields substantial reduction in power consumption
v3.7  v3.7.1 cures anomalous noise of ch10
Fine parameter optimization of v3.7.1 completed
Missing NAND chip in segment cards  ON/OFF of fast-reset not working
NAND chip installed in 106 segments cards