## Slide 1: Validation of GateRBEActor for carbon ion pencil beam therapy
Yihan Jia, Favaretto Martina, Andreas F Resch, Markus Stock, Loïc Grevillot
01/04/2025

## Slide 2: Why relative biological effectiveness (RBE)?
In vitro, in vivo
[2] Hartmann L, et al., Scientific reports (2020)
Dose
Radiation quality (LETd)
Cell line (radiosensitivity)
Endpoint (survival fraction, tumor control probability …)
…
Notes: 1. A factor that characterizes the biological effectiveness of a radiation
Murine tumor cell line

## Slide 3: Why relative biological effectiveness (RBE)?
In vitro, in vivo
[1] Ridolfi, R (2018)
[2] Hartmann L, et al., Scientific reports (2020)
Dose
Radiation quality (LETd)
Cell line (radiosensitivity)
Endpoint (survival fraction, tumor control probability …)
…
Notes: 1. A factor that characterizes the biological effectiveness of a radiation
Murine tumor cell line

## Slide 4: Why relative biological effectiveness (RBE)?
In vitro, in vivo
[1] Ridolfi, R (2018)
[2] Hartmann L, et al., Scientific reports (2020)
Dose
Radiation quality (LETd)
Cell line (radiosensitivity)
Endpoint (survival fraction, tumor control probability …)
…
Notes: 1. A factor that characterizes the biological effectiveness of a radiation
Murine tumor cell line

## Slide 5: Why relative biological effectiveness (RBE)?
Predictive models: phenomenological, mechanistic
Clinically used models
  Local Effect Model I (LEM1)
  Modified Microdosimetric Kinetic Model (mMKM)
Clinical experience based on RBE-weighted dose
  Prescription, fractionation
  Constraints for organs at risk
In vitro, in vivo
[1] Ridolfi, R (2018)
[3] modified based on Japanese P-RAP
clinical outcome
Caen(ARCHADE)
Notes: GSI pilot project: 1997-2008
NIRS: 1994-

## Slide 6: How does RBEActor work?
You can read the User Guideline doc
Here’s the scientific rationale:

## Slide 7: How does RBEActor work?
[2] Hartmann L, et al., Scientific reports (2020)
mMKM
LEM1lda
Linear Quadratic model

## Slide 8: How does RBEActor work?
[2] Hartmann L, et al., Scientific reports (2020)
mMKM
LEM1lda
Linear Quadratic model

## Slide 9: How does RBEActor work?
[2] Hartmann L, et al., Scientific reports (2020)
mMKM
LEM1lda
Linear Quadratic model
Linear Quadratic – linear model

## Slide 10: How does RBEActor work?
GetTotalEnergyDeposit
mMKM
LEM1lda
DoseActor

## Slide 11: How does RBEActor work?
SteppingAction
GetTotalEnergyDeposit
mMKM
LEM1lda
Mixed field radiation
constant
DoseActor
GateBeamQualityActor
Saturation-corrected dose-averaged specific energy
Microdosimetric quantity
Sub-nucleus domain
[5] Zeider & Rossi (1980)
[5]

## Slide 12: How does RBEActor work?
GetTotalEnergyDeposit
mMKM
LEM1lda
GetCurrentDEDX
constant
DoseActor
GateBeamQualityActor
SteppingAction
Mixed field radiation
Initial slope of cell survival function
Radiobiological quantity

## Slide 13: How does RBEActor work?
mMKM
LEM1lda
EndSimulationAction
Linear Quadratic model
Linear Quadratic – linear model

## Slide 14: How does RBEActor work?
mMKM
LEM1lda
Linear Quadratic model
Linear Quadratic – linear model
EndSimulationAction

## Slide 15: Generation of RBEActor input tables
PRUDENT: PRojectile μ-Dosimetric Effect in Nucleus evaluation Tool
Calculates the table of mMKM and table of LEM1lda
mMKM
constant

## Slide 16: Generation of RBEActor input tables
PRUDENT: PRojectile μ-Dosimetric Effect in Nucleus evaluation Tool
Calculates the table of mMKM and table of LEM1lda
LEM1lda

## Slide 17: Comparison of central RBE profile against TPS
Treatment planning system (TPS)
  Dose engine: Pencil Beam Algorithm V7.0
  RBE model: mMKM, LEM1 classic, LEM1 low dose approximation

## Slide 18: Comparison of central RBE profile against TPS
Treatment planning system (TPS)
  Dose engine: Pencil Beam Algorithm V7.0
  RBE model: mMKM, LEM1 classic, LEM1 low dose approximation
120.0 MeV/u pencil beam: rel. diff. < 3% until R80-1 mm
LEM1lda
mMKM
Simulation settings:
G4_WATER
QGSP_BIC_HP_EMZ
reference input tables
1E5 primaries

## Slide 19: Comparison of central RBE profile against TPS
Energy modulated scanned beams:
  field size 6 cm X 6 cm, modulation length 6 cm
  depth of field center 6 cm
  Uniform LEM1-RBE-weighted dose 2.3 Gy
Agreement with TPS profile better than 2% (reference table) and 3% (PRUDENT) within target region
Simulation settings:
G4_WATER
QGSP_BIC_HP_EMZ
2% statistical uncertainty goal
LEM1lda
mMKM
LEM1 classic
mMKM

## Slide 20: Comparison of central RBE profile against TPS
Physics lists: QGSP_BIC_HP_EMZ, ShieldingLIQMD_EMZ
Impact of nuclear interaction modelling:
  agreement with TPS profile better than 3% within target region
  Larger divergence in the fragmentation tail for mMKM, compared to LEM1lda
Simulation settings:
G4_WATER
Reference table
2% statistical uncertainty goal
LEM1lda
mMKM

## Slide 21: Conclusion
GateRBEActor supports the two clinically used RBE models, modified Microdosimetric Kinetic Model (mMKM) and Local Effect Model I (LEM1)
GateRBEActor is primarily benchmarked against clinical TPS RayStation 2024A
Further benchmarking will involve additional fields in water and fields in patient geometry
GateRBEActor is available in GATE 10.0.2 / GATE-RTiON V2!
Opengate documentation
Courtesy of M. Favaretto, “GateRTion v2: a reference GATE 10 version for ion beam therapy”

## Slide 22: Acknowledgement
Medical Physics team, MedAustron
Michael Dutter and Patrick Simons, MedAustron
Giuseppe Magro, PhD, CNAO
Mansure Schafasand, PhD, MedAustron
The financial support from the “Niederösterreichischen Wirtschafts- und Tourismusfonds” and European union in form of the EFRE fonds is gratefully acknowledged.

## Slide 23: Thank you for your attention!

## Slide 24: References
[1] Ridolfi, Riccardo. "Study of the track reconstruction in the FOOT experiment for Hadrontherapy." (2018).
[2] Hartmann, Laura, et al. "Photon versus carbon ion irradiation: immunomodulatory effects exerted on murine tumor cell lines." Scientific reports 10.1 (2020): 21517.
[3] R&D status in various countries. https://www.particle.or.jp/hirtjapan/en/medical/current/world.html Accessed date: 22/03/2025.
[4] Vector illustration by sasami018/ AC-illust on Page 10-14
[5] Zaider, M., and H. H. Rossi. "The synergistic effects of different radiations." Radiation research (1980): 732-739.

## Slide 25: Electronic stopping power in water (G4_WATER)

## Slide 26: SURVIVAL z*1D table