## Slide 1: OpenGATE for helium commissioning
H. Fuchs, Medical University of Vienna

## Slide 2: MedAustron ion therapy center
Synchrotron facility
  Protons: 60-252 MeV (3-38 cm)
  Carbon Ions: 120-402 MeV/u (3-27 cm)
3 clinical rooms
1 Dedicated research room – IR1
  For research users of Austrian universities
  Dedicated beam time
  Helium ions becoming available

## Slide 3: Helium ion
Accelerator commissioning for Helium ions finished
  2020-2024
Common tuning and acceptance phase
  Accelerator physicists and research users
  Tuning in room beam parameters
  Ensure specifications met
  Beam data measurements
© Thomas Kästenbauer

## Slide 4: Helium ion
Target parameters
  Ranges in water 1-80 cm
  Spot sizes comparable to clinical C12 beam
    ~7 mm FWHM
  Calibration in number of particles
© Thomas Kästenbauer

## Slide 5: Tasks
Range tuning
Optics tuning
Absolute dose calibration
Base data measurements
  Full IDD
  Ranges
  Spot sizes / Lateral fall off
Only ~50 h beamtime for acceptance & commissioning

## Slide 6: Preparatory work - GATE
Monte Carlo modell of MedAustron nozzle
  Existing for Gate 9
  Migrated to OpenGATE 10
    About 2 days (incl. validation)
OpenGATE provides flexibility
  Necessary due to recent beam delivery hardware changes
  Changes ~ 1 day

## Slide 7: Range tuning
Ideally energies with ranges in mm steps
Using GATE
  Prediction of suitable energies
Initial measurements
  16 IDD measured
First He beam measured

## Slide 8: Range tuning
GATE
  Initial energies for calibration
  Fine tuning of energies
  Creation of full energy list
    ~350 energies available
    1-30 cm in mm steps
    Beyond in cm resolution
Experimental verification
  38 energies verified
  Energies within 0.1 mm of target range
First He beam measured

## Slide 9: Absolute dose calibration
Calibration in terms of number of particles not MU
  Measurement of dose area product
  Using stopping power conversion to number of particles
Experimental measurement of dose area product
Stopping power calculation
  Monte Carlo simulation
  Fixed measurement depth: 14 mm

## Slide 10: Absolute dose calibration
Calculation of calibration curve
Experimental verification
  Verified 35 energies
  Agreement
    Within 1.5 %
    Average agreement 0.07%
  From 1 – 80 cm
Chart: initial Energy [MeV/u]
| 54.6 | 65.400000000000006 | 75 | 83.8 | 87.9 | 91.9 | 99.5 | 106.6 | 113.5 | 120 | 126.3 | 129.30000000000001 | 132.4 | 138.19999999999999 | 143.9 | 149.5 | 160.19999999999999 | 165.3 | 170.4 | 180.2 | 189.7 | 198.9 | 203.4 | 207.9 | 212.3 | 220.8 | 261.10000000000002 | 298 | 332.4 | 364.7 | 401 | 402.8
 | -1.3965 | -1.2188 | 0.039221 | 0.34989 | 0.65111 | 0.60767 | 0.51463 | 0.42978 | 0.31753 | 0.13812 | 0.086704 | 0.40772 | 0.08189 | 0.075212 | 0.063746 | 0.018451 | 0.0040284 | 0.074968 | -0.10077 | -0.09185 | 0.1552 | -0.023488 | 0.070628 | 0.3343 | 0.1049 | 0.11526 | 0.10578 | 0.027439 | 0.1226

## Slide 11: Summary & Outlook

## Slide 12: Summary
Helium beam time very restricted
  2 shifts for range tuning
  2 shifts for absolute dose calibration and verification
  Optic measurements still ongoing
Extremely efficient use of time
  Only possible with Monte Carlo simulations
Gate 10 makes changing geometries substantially easier