## Slide 1: Measuring Temperatures Below 1 K Beyond the Metrology Laboratory
Laurent Pitre
GRD-DI2I
2025

## Slide 2
Clement Tauzin est docteur depuis le 27/5/25

## Slide 3: PLTS-2000 the international scale below 1K
Provisional Low Temperature Scale
Made to extend ITS-90 below 0.65 K
Based on a thermophysical property:
  The melting curve of ultra-pure helium-3
Defined from 0.9 mK to 1 K
Formal expression:
Solid
Liquid
Liquid+Solid
Notes: Propritété thermophysique. (independent de l’espace et du temps) Equation base sur des mesures de thermomètres primaire avant 2000 (Liquid+Solid) animation
Knot . 65
Polynomilal with 13 terms

## Slide 4: PLTS-2000
Constructed from primary thermometry data prior to 2000:
  University of Florida - 1995
  PTB - 1996
  NIST - 1998
Aimed to approximate the thermodynamic temperature
from 0.9 mK to 1 K
Adopted by the BIPM in 2000
200 mK
Notes: Propritété thermophysique. Equation base sur des mesures de thermomètres primaire avant 2000

## Slide 5: The Melting Curve Thermometer au LNE
MCT (Melting Curve Thermometer)
Based on the Adams Gauge
Capacitance of a plate capacitor:
3He filling line
Membrane
Electrodes
Silver sinter
3He chamber
0 MPa
4 MPa
Vacuum
Notes: Utilise la propriété d’aller sur la melting curve, rajouter vacuum
Inverse plate spacing
Définir x

## Slide 6: T-T2000
Chart: 
Good agreement with NIST-98
Fair agreement with PTB-96 at higher temperature
UF-95
PTB-96
NIST-98
Comparison of PLTS-2000 and the Magnetic Field Fluctuation Thermometer at LNE-Cnam
Clement Tauzin, Mark Plimmer, Fernando Sparasci, Wenxiang Guo, D Rod White and Laurent Pitre
https://iopscience.iop.org/article/10.1088/1681-7575/ade4d0
8 mK
80 mK
Notes: Rajouter sur le graph les noms des courbes

## Slide 7: SRD1000 a transfer standard?
Chart: T /mK
Contains 10 reference points
Notes: We model the susceptibility with a sigmoid to capture the sharp change around the critical temperature. This gives us a reliable fit for extracting Tc.

## Slide 8: SRD1000
Uses the Meissner effect
Chart: Temperature
| 0 | 0.02 | 0.04 | 0.06 | 0.08 | 0.1 | 0.12 | 0.14000000000000001 | 0.16 | 0.18 | 0.2 | 0.22 | 0.24 | 0.26 | 0.28000000000000003 | 0.3 | 0.32 | 0.34 | 0.36 | 0.38 | 0.4 | 0.42 | 0.44 | 0.46 | 0.48 | 0.5 | 0.52 | 0.54 | 0.56000000000000005 | 0.57999999999999996 | 0.6 | 0.62 | 0.64 | 0.66 | 0.68 | 0.7 | 0.72 | 0.74 | 0.76 | 0.78 | 0.8 | 0.82 | 0.84 | 0.86 | 0.88 | 0.9 | 0.92 | 0.94 | 0.96 | 0.98 | 1
 | 1.0306e-10 | 2.2936e-10 | 5.1045e-10 | 1.136e-09 | 2.5283e-09 | 5.6268e-09 | 1.2523e-08 | 2.787e-08 | 6.2025e-08 | 1.3804e-07 | 3.0721e-07 | 6.837e-07 | 1.5216e-06 | 3.3862e-06 | 7.5355e-06 | 1.6768e-05 | 3.7301e-05 | 8.294e-05 | 0.00018421 | 0.00040813 | 0.00089931 | 0.0019583 | 0.0041586 | 0.0083991 | 0.015501 | 0.025 | 0.034499 | 0.041601 | 0.045841 | 0.048042 | 0.049101 | 0.049592 | 0.049816 | 0.049917 | 0.049963 | 0.049983 | 0.049992 | 0.049997 | 0.049998 | 0.049999 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
Notes: Chaque point c’est 30 min de mesures

## Slide 9: SRD1000
Each point on the graph requires about 30 min
Increasing and decreasing temperature to highlight possible hysteresis
Notes: Tout est automatisé !

## Slide 10: Determination of Tc
Modelling with a sigmoidal function:
More robust, lower uncertainty than the double plateau method
Chart: Temperature
| 0 | 0.02 | 0.04 | 0.06 | 0.08 | 0.1 | 0.12 | 0.14000000000000001 | 0.16 | 0.18 | 0.2 | 0.22 | 0.24 | 0.26 | 0.28000000000000003 | 0.3 | 0.32 | 0.34 | 0.36 | 0.38 | 0.4 | 0.42 | 0.44 | 0.46 | 0.48 | 0.5 | 0.52 | 0.54 | 0.56000000000000005 | 0.57999999999999996 | 0.6 | 0.62 | 0.64 | 0.66 | 0.68 | 0.7 | 0.72 | 0.74 | 0.76 | 0.78 | 0.8 | 0.82 | 0.84 | 0.86 | 0.88 | 0.9 | 0.92 | 0.94 | 0.96 | 0.98 | 1
 | 1.0306e-10 | 2.2936e-10 | 5.1045e-10 | 1.136e-09 | 2.5283e-09 | 5.6268e-09 | 1.2523e-08 | 2.787e-08 | 6.2025e-08 | 1.3804e-07 | 3.0721e-07 | 6.837e-07 | 1.5216e-06 | 3.3862e-06 | 7.5355e-06 | 1.6768e-05 | 3.7301e-05 | 8.294e-05 | 0.00018421 | 0.00040813 | 0.00089931 | 0.0019583 | 0.0041586 | 0.0083991 | 0.015501 | 0.025 | 0.034499 | 0.041601 | 0.045841 | 0.048042 | 0.049101 | 0.049592 | 0.049816 | 0.049917 | 0.049963 | 0.049983 | 0.049992 | 0.049997 | 0.049998 | 0.049999 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
Notes: We model the susceptibility with a sigmoid to capture the sharp change around the critical temperature. This gives us a reliable fit for extracting Tc.

## Slide 11: SRD1000
The fitting function can be adjusted to fit more complex shapes:
Ir
Notes: We model the susceptibility with a sigmoid to capture the sharp change around the critical temperature. This gives us a reliable fit for extracting Tc. Trait pour le Zéro des résidus Rajouter les élements

## Slide 12: SRD1000
TC and transition shape depend on material properties:
Material purity
Mechanical stress
Grain size and crystal structure
Residual magnetic fields
Chart: T /mK
| 97.118524999999991 | 98.280999999999992 | 98.441000000000003 | 98.643765217391291 | 98.642108333333297 | 98.628 | 98.801000000000002 | 98.933999999999997 | 98.966999999999999 | 99.105000000000004 | 99.328000000000003 | 99.499000000000009 | 99.682000000000002 | 99.853999999999999 | 100.048 | 100.39700000000001 | 100.429 | 100.39700000000001 | 100.12899999999999 | 99.688999999999993 | 99.48 | 99.222000000000008 | 99.071999999999989
SN005 | 0 | 0.01034 | 0.019095 | 0.046449 | 0.045257 | 0.043535 | 0.087858 | 0.48558 | 0.55736 | 0.89363 | 0.97815 | 0.99071 | 0.99556 | 0.99779 | 0.999 | 0.99997 | 0.99934 | 1 | 0.99902 | 0.9953 | 0.98868 | 0.9713 | 0.83921
SN236 | 0.046602 | 0.046043 | 0.056781 | 0.038676 | 0.024981 | 0.0026325 | 0 | 0.038039 | 0.1199 | 0.28428 | 0.5899 | 0.72012 | 0.83464 | 0.90957 | 0.95811 | 0.98839 | 0.99326 | 0.996 | 0.99583 | 0.99645 | 0.99429 | 0.99635 | 0.99782 | 0.99654 | 0.99673 | 0.99792 | 0.99903 | 0.99829 | 0.99996 | 1 | 0.99829 | 0.99865
SN237 | 0.0093543 | 0.000741 | 0.00080523 | 0.015462 | 0.062386 | 0.23612 | 0.46721 | 0.64519 | 0.92924 | 0.98955 | 1 | 0.99898 | 0.95879 | 0.82981 | 0.51032 | 0.38746 | 0.11764 | 0.039482 | 0.0040947 | 0.0077607 | 0
Ir
Each sample has to be calibrated individually
Notes: We model the susceptibility with a sigmoid to capture the sharp change around the critical temperature. This gives us a reliable fit for extracting Tc.

## Slide 13: Long term stability
SN 005 was tested in 2002 and 2023
Different cryostat
Different wiring
Excellent agreement
Ir
Notes: We model the susceptibility with a sigmoid to capture the sharp change around the critical temperature. This gives us a reliable fit for extracting Tc. Parcequ’on a T2000

## Slide 14: Long term stability
Notes: Rajouter elements We model the susceptibility with a sigmoid to capture the sharp change around the critical temperature. This gives us a reliable fit for extracting Tc. Parcequ’on a T2000

## Slide 15: Calibrating Cryoconcept’s MFFT on site using the transfer standard from LNE-Cnam
Need of a reference temperature:
  SRD1000 fixed temperatures
Chart: T /mK
Notes: Clore T-2000 et reprendre sur le sujet

## Slide 16
At least 2 years to get this product
MFFT from Magnicon at Cryoconcept

## Slide 17: Thermal noise
An electron in thermal motion produces a magnetic field
Random electron movements cause fluctuating magnetic fields
Many electrons randomly moving generate a measurable thermal magnetic noise with a SQUID
x
y
z
Electron
Sensing coil
Notes: Tiny signal/ ordre de grandeur

## Slide 18: The Thermal Noise
The Thermal Magnetic Noise
Magnetic noise is generated by thermal motion of electrons in any conductor.
The intensity of this noise is directly proportional to temperature.
Notes: Changer l’équation

## Slide 19
The Thermal Magnetic Noise
Magnetic noise is generated by thermal motion of electrons in any conductor.
The intensity of this noise is directly proportional to temperature.
The Thermal Noise

## Slide 20
Chart: F /Hz
The Thermal Noise
Chart: F /Hz
Power Spectral Density
Notes: Poser le basse bas

## Slide 21: Magnetic Field Fluctuation Thermometer at LNE
Chart: Residuals
Chart: Fréquence /Hz
Notes: Par rapport aux 130 Mesures select temperatures in multiples of 2 over 130 Measurements
Plus d’un ordre de grandeur

## Slide 22: The Working Standard at Cryoconcept with issues
2 MFFT:
Unit 308
Unit 311
Very Noisy
Given with supplier calibration file
13.5 mK
Notes: Rajouter le logo de la compagnie Température given in our lab with the SRD237

## Slide 23: Comparaison avec T2000
Comparaison at the transition of SRD1000:
13.5 mK
20.77 mK
31.75 mK
66.05 mK
96.1 mK
Chart: T2000 /mK
| 13.5 | 20.77 | 31.75 | 66.05 | 96.1
308 | 2.52 | 2.21 | 2.04 | 2.54 | 2.27
311 | 0.38 | 0.08 | -0.02 | 0.67 | 0.44 | 0
+ 2.32 mK
The unit 311 calibration is good
The unit 308 is not good
  2.5 mK at 14 mK
  -7 mK at 520 mK
Too much unwanted noise in the spectrums

## Slide 24: Comparison against PLTS-2000 at Cryoconcept
The unit 311 calibration is good
The unit 308 is not good
  2.5 mK at 14 mK
  -7 mK at 520 mK
Too much unwanted noise in the spectrums
Chart: T2000 /mK
| 13.5 | 20.77 | 31.75 | 66.05 | 96.1 | 518.1
308 | 2.33 | 1.86 | 1.52 | 1.48 | 0.75 | -7.23
311 | 0.38 | 0.08 | -0.02 | 0.67 | 0.44 | 0.82
Notes: 17% d’écart

## Slide 25
Measuring temperature is easy… until it really matters