## Slide 1: Design of a sub-0.5 mm resolution PET using GATE with staggered 3-layer DOI detectors
Han Gyu Kang1, Hideaki Tashima1, and Taiga Yamaya1
            1. National Institutes for Quantum Science and Technology (QST), Japan
kang.hangyu@qst.go.jp
2025 IEEE NSS/MIC in Yokohama
GATE User Meeting
Date: Thursday, 6-Nov-2025
Time: 13:15-13:30 (12 min presentation + 3 min Q&A)
Place: G216 on Level 2, PACIFICO Yokohama North building
Session: GATE User Meeting
Notes: Good afternoon. I’m Han Gyu Kang, a researcher at QST in Japan.
It is my great pleasure to present my research works in this session.
The title of my presentation is “Design of sub-0.5 mm resolution PET using GATE with staggered 3-layer DOI detectors”.

## Slide 2: Outline
Factors affecting PET resolution
Optimal design of a sub-0.5 mm resolution PET
Preliminary experimental results
Conclusions
Notes: Here is the outline.
First, I’ll explain the factors affecting PET resolution.
Second, I’ll present the optimal design of a sub-0.5 mm resolution PET.
Third, I’ll be presenting the preliminary experiment results such as front-end electronics and PET gantry design.
Lastly, I’ll summarize my presentation and discuss the future plan

## Slide 3: Factors affecting the PET resolution
Multiplexing
Photo sensor
Crystal
Scintillation detector
④ Crystal pitch
e-
Crystal
Photo sensor
β+
① Positron range
② Acollinearity
③ Parallax error
β+
Positron
e-
Electron
Annihilation（511 keV）
Ring diameter
Small ring diameter
(e.g. <50 mm)
Fine crystal pitch
(e.g. <1 mm)
Depth encoding
(e.g. 3-layer)
Opaque reflector
(e.g. BaSO4)
Solution>
⑤ Block effect
ESR
BaSO4
A PET scanner that considered these factors
0.1 mm FWHM
18F
Notes: This slide illustrates the factors affecting the PET resolution.
PET resolution is defined as a convolution of several different factors as shown in this equation.
The first factor is the positron range which has about 0.1 mm in full width at half maximum for F-18.
The second factor is acollinearity which can be reduced by using small ring diameter of less than 50 mm.
The third factor is parallax error, in which the radial resolution is blurred at the periphery of a PET FOV.
The parallax error can be reduced by using depth encoding detectors such as staggered 3-layer configuration.
The most important factor is the crystal pitch, which dominates the overall PET resolution.
The PET resolution can be substantially enhanced by using fine crystal pitch of 1 mm or even smaller.
To encode a large amount of finely pixelated crystal array, multiplexing readout is commonly used.
However, for multiplexing readout, the block effect arises thereby degrading the crystal decoding accuracy.
The block effect can be effectively reduced by using optically opaque reflector such as barium sulfate powder rather than using the ESR film.
Last year, we developed a PET scanner that addressed these issues to achieve submillimeter-resolution.

## Slide 4: SR-PET achieved 0.55 mm resolution (2022 MIC)
Axial FOV
=51.5 mm
[2] Kang et al., 2023 Journal of Nuclear Medicine
0.85
0.8
0.75
0.55
0.5
0.45
PET image
SR-PET (Submillimeter-Resolution PET)[2]
SR-PET achieved the 0.55 mm resolution
Sub-0.5 mm resolution was NOT achieved due to
Large crystal pitch
Unoptimized crystal layer
Crystal pitch = 1 mm
Crystal layer = 4+4+7 mm (1st+2nd+3rd)-> DOI
1st layer
2nd layer
3rd layer
Pitch=1 mm
Goal: Design of a Sub-0.5 mm Resolution PET with 3-layer DOI detectors
Notes: The submillimeter-resolution PET, which is named as SR-PET has a 1 mm crystal pitch.
The 3-layer LYSO crystal arrays were stacked with staggered configuration to encode DOI information.
With the SR-PET, a resolution approaching 0.5 mm was achieved in a phantom study.
However, sub-0.5 mm resolution was not achieved due to the large crystal pitch and unoptimized crystal layer.
Therefore, the goal of this study is to design a sub-0.5 mm resolution PET with 3-layer DOI detectors.

## Slide 5
SPECThead
SPECThead
crystal
pixel
ECAT
ecat
block
crystal
CylindricalPET
cylindricalPET
rsector
module
crystal
layer0
layer1
Optical System
crystal
pixel
OpticalSystem
CTscanner
CTscanner
module
cluster
pixel
PETscanner
PETscanner
level1
level2
level3
level4
level5
Scanner
scanner
level1
level2
level3
level4
level5
Predefined PET systems in GATE
No DOI
2-layer DOI
3-layer DOI
Notes: GATE provides various types of predefined PET systems such as ECAT, CylindricalPET and PET scanner.

## Slide 6
SPECThead
SPECThead
crystal
pixel
ECAT
ecat
block
crystal
CylindricalPET
cylindricalPET
rsector
module
crystal
layer0
layer1
Optical System
crystal
pixel
OpticalSystem
CTscanner
CTscanner
module
cluster
pixel
PETscanner
PETscanner
level1
level2
level3
level4
level5
Scanner
scanner
level1
level2
level3
level4
level5
PETscanner was used for 3-layer DOI PET
Predefined PET systems in GATE
Notes: Here, we used the “PETscanner” for the simulation of the 3-layer DOI PET scanner.

## Slide 7
GATE PET simulation
From GATE simulation to image reconstruction
ROOT output
Custom-written code
E.W=440-560 keV
List-mode data
Block ID
Crystal ID
DOI ID
Image Reconstruction
(In house software*)
*Developed by Dr. Hideaki Tashima (QST)
1st (level2ID)
2nd (level3ID)
3rd (level4ID)
Block (level1ID)
Reconstructed PET image
0.85
0.8
0.75
0.55
0.5
0.45 mm
10 mm
GATEv9.1
Notes: This slide shows the procedure from the GATE simulation to image reconstruction.
First, we conducted the GATE PET simulation with 3-layer DOI detector configuration.
Second, the ROOT output file was imported by using a custom-written ROOT code to generate list-mode data.
An energy window of 440 to 560 keV was used.
Lastly, the list-mode data were reconstructed by using an in-house software developed in our lab.

## Slide 8: Sub-0.5 mm resolution PET design
Notes: First, let me describe PET design.

## Slide 9: Sub-0.5 mm resolution PET
ID=48 mm
2nd layer
1st layer
3rd layer
SiPM (5 × 5 array)
Phantom
<3D view of the PET>
<Staggered 3-layer DOI detector>
14×14
13×13
13×12
Crystal pitch
=0.8 mm
LYSO array
2.4 mm
23.4 mm
Lightguide
Notes: The PET scanner has an inner diameter of 48 mm and axial coverage of 23.4 mm.
The PET scanner has two rings, each of which has 16 DOI detectors.
Each DOI detector consists of 3-layer LYSO crystal array to encode DOI information.
The crystal pitch was 0.8 mm to achieve sub-0.5 mm resolution.
The 3-layer LYSO array was optically coupled to an SiPM 5x5 array with a pitch of 2.4 mm.
A 1 mm thick lightguide was inserted between the crystal and SiPM.

## Slide 10: Crystal layer designs
3 mm
5 mm
3 mm
4 mm
3 mm
4 mm
7 mm
4 mm
4 mm
1st layer
2nd layer
3rd layer
(A) 4+4+7 mm
SiPM 5×5 array
1st layer
2nd layer
3rd layer
(B) 3+4+4 mm
SiPM 5×5 array
1st layer
2nd layer
3rd layer
(C) 3+3+5 mm
SiPM 5×5 array
15 mm
11 mm
11 mm
Notes: We used three different crystal layer designs.
The first design-A has the crystal thickness of 4, 4 and 7 mm for the 1st, 2nd, and 3rd layers thereby resulting a total thickness of 15 mm.
The second design-B has the crystal thickness of 3, 4 and 4 mm, resulting in a total thickness of 11 mm.
The design-C has the crystal thickness of 3, 3 and 5 mm for the 1st, 2nd, and 3rd layers.
The relatively short crystal thickness was used for the design B and C to improve the DOI resolution at the expense of sensitivity.
The crystal thickness of each layer was chosen to make a uniform detection efficiency for each layer.

## Slide 11: Spatial resolution
Notes: Let me move on to spatial resolution

## Slide 12: Spatial resolution with the NEMA 22Na sources
52.5 mm
(A) 4+4+7 mm
Sensitivity
15 mm
1.61%
22Na
15 mm
48 mm
(B) 3+4+4 mm
11 mm
1.29%
22Na source
15 mm
48 mm
(C) 3+3+5 mm
11 mm
1.32%
22Na source
15 mm
5 mm
E.W=440-560 keV
Notes: The spatial resolution was evaluated by using a NEMA Na-22 source with a diameter was 0.25 mm.
The Na-22 sources were placed from the center to radial offset position of 15 mm with a spacing of 5 mm.
The spatial resolution was evaluated for three different crystal layer designs of A, B, and C.

## Slide 13: Spatial resolution (FBP)
Chart: Radial offset [mm]
| 0 | 5 | 10 | 15
A (4+4+7 mm) | 0.72145 | 0.71407 | 0.70874 | 1.2769
B (3+4+4 mm) | 0.694 | 0.67861 | 0.68188 | 0.78424
C (3+3+5 mm) | 0.69186 | 0.68707 | 0.72911 | 0.70369
(a) Radial FWHM
Chart: Radial offset [mm]
| 0 | 5 | 10 | 15
A (4+4+7 mm) | 1.2531 | 1.2572 | 1.3326 | 2.31
B (3+4+4 mm) | 1.2013 | 1.1158 | 1.0557 | 1.3711
C (3+3+5 mm) | 1.1923 | 1.1284 | 1.1673 | 1.266
(b) Radial FWTM
Design C showed the most uniform spatial resolution up to the radial offset of 15 mm
Notes: This slide shows the spatial resolution as a function of radial offset.
Design C showed the most spatial resolution up to the radial offset of 15 mm.
The full width at tenth maximum was over 1 mm for all designs.
Design C provided uniform FWTM across the radial offset.

## Slide 14: Resolution phantom imaging
Notes: Let me move on to the resolution phantom imaging.

## Slide 15: Resolution phantom imaging
ID=48 mm
Phantom
<3D view of the PET>
23.4 mm
11.9 mm
11.9 mm
Resolution phantom
Acquisition time = 10 min
Total activity of 18F= 0.95 MBq
10 mm
0.85
0.8
0.75
0.55
0.5
0.45 mm
Notes: To evaluate imaging performance, a resolution phantom was placed at the center of the FOV.
The phantom has an outer diameter of 11.9 mm, and axial length of 10 mm.
The resolution phantom has 6 different sections, and the smallest rod diameter was 0.45 mm.
All the rods of the phantom were filled with 18F of 0.95 MBq.
The acquisition time was 10 min.

## Slide 16: Normalization by rotating a 68Ge source
68Ge
Normalization
68Ge cylindrical source
Diameter: 2 mm
Length: 23.4 mm
Activity: 0.72 MBq
Scan time: 2 hr
20.4 mm
Notes: For normalization a cylindrical 68Ge source with an activity of 0.72 MBq was used.
Coincidence events were obtained for 120 min during the rotation of the source.

## Slide 17: Resolution Phantom (OSEM)
0.85
0.8
0.75
0.55
0.5
0.45
A (4+4+7 mm)
2
10
Resolvability=83.3%
0
12
Resolvability=100%
Subset=8
Iterations=50
5 mm
B (3+4+4 mm)
0.85
0.8
0.75
0.55
0.5
0.45
5 mm
C (3+3+5 mm)
0.85
0.8
0.75
0.55
0.5
0.45
5 mm
Resolvability=100%
0
12
We chose design C considering the uniform spatial resolution across the FOV
Notes: Here you see is the reconstructed PET image using OSEM with design A.
All the submillimeter rod structures were clearly resolved.
For quantitative analysis, line profiles were drawn in three different directions.
Subsequently, valley-to-peak ratio histogram was obrtained.
The resolvability was defined as the ratio of the number of VPR under 0.735 to total number of VPRs.
The resolvability of design A was 83.3% as two line profiles were above the 0.735.
With designs B and C, the resolvability of the same rod structure were 100% thanks to the reduced crystal layer thickness.
We chose design C considering the uniform spatial resolution across the FOV.
(A) 4+4+7 mm: 7,908,790
(B) 3+4+4 mm: 8,006,160
(C) 3+3+5 mm: 8,217,194

## Slide 18: PET system development
Notes: Based on the simulation results, we recently developed a prototype PET scanner.

## Slide 19: Half-millimeter Resolution PET (HR-PET)[2]
48 mm
23.4 mm
<Staggered 3-layer DOI detector>
Pitch=0.8 mm
LYSO 3-layer
1st layer
2nd layer
3rd layer
Front-end board
Amplifier board
[2] Kang et al., 2025 IEEE TMI
Notes: We developed a half-millimeter resolution PET, which is named as HR-PET.
The HR-PET has an inner diameter of 48 mm, and the PET consists of front-end and amplifier boards.
Each front-end board has 3-layer LYSO crystal arrays with a crystal pitch of 0.8 mm.
The axial FOV is 23.4 mm.

## Slide 20: Resolution phantom (Experiment[2] vs. Simulation)
[2] Kang et al., 2025 IEEE TMI
0.85
0.8
0.75
0.55
0.5
0.45
11.9 mm
Resolution phantom
18F = 1.2 MBq
Scan=60 min
Iteration=50
0.85
0.8
0.75
0.55
0.5
0.45
Experiment
11
1
Resolvability
= 91.7%
Simulation
0.85
0.8
0.75
0.55
0.5
0.45
0
12
Resolvability=100%
Experimental results showed slightly worse resolvability than simulation mainly because of the crystal decoding error and other unmodeled factors
Notes: For imaging performance evaluation, a resolution phantom was filled with 18F of 1.2 MBq and scanned for 60 min.
The images were reconstructed using OSEM with 50 iterations.
With experiments, the 0.45 mm rod structure was identified with a resolvability of 91.7%, which was slightly worse than simulation results.

## Slide 21: Conclusions
We designed a sub-0.5 mm resolution PET with GATE
  Design (C) 3+3+5 mm was chosen due to the uniform resolution over the FOV
  The 0.45 mm rod structure can be resolved clearly in simulation and experiment
1st layer
2nd layer
3rd layer
Design (C)
0.85
0.8
0.75
0.55
0.5
0.45
Phantom image (GATE simulation)
3 mm
5 mm
3 mm
Notes: Let me summarize my presentation.
We designed a sub-0.5 mm resolution PET with GATE.
Design C was chosen due to the uniform resolution over the FOV.
The 0.45 mm rod structure can be resolved clearly in both simulation and experiment.

## Slide 22: Future works
Full validation of GATE simulation results with experimental data
Physical and in vivo imaging performance evaluation of the HR-PET (Talk on Friday 10 AM, M-13-01)
10 mm
MIP image of mouse brain
(18F-FITM)
Matrix= 240×240×117
Voxel= 0.2×0.2×0.2 mm3
DRF=1 mm, Iter=10
Notes: Future works includes the validation of GATE simulation results with experimenta data
Tomorrow I will give a talk about the physical and in vivo imaging performance evaluation of the HR-PET.

## Slide 23: Thank you for your kind attention!
2024 Imaging Physics Group
April-2024 in QST
HanGyu
Tashima
Yamaya
17/17
Notes: This is end of my presentation.
Thank you for your kind attention.

## Slide 24: Progress of the PET development
ID=48 mm
Phantom
23.4 mm
GATE simulation
Optimize the crystal layer
Front-end development
Optimize the PET detector
Front-end
Amplifier
3D CAD gantry design
Notes: Based on the GATE simulation, a PET detector was developed and gantry design was finished.
Next year, I would like to show you the imaging performance of the proposed PET scanner.

## Slide 25
FBP
FBP+GF
Chart: Radial offset [mm]
Chart: Radial offset [mm]
| 0 | 5 | 10 | 15
FBP (4+4+7 mm) | 0.72145 | 0.71407 | 0.70874 | 1.2769
FBP+GF (4+4+7 mm) | 0.7514 | 0.72161 | 0.73031 | 1.3793
Chart: Radial offset [mm]
Chart: Radial offset [mm]
| 0 | 5 | 10 | 15
FBP (3+4+4 mm) | 0.694 | 0.67861 | 0.68188 | 0.78424
FBP+GF (3+4+4 mm) | 0.70573 | 0.67069 | 0.69846 | 0.77148
Chart: Radial offset [mm]
Chart: Radial offset [mm]
| 0 | 5 | 10 | 15
FBP (3+3+5 mm) | 0.69186 | 0.68707 | 0.72911 | 0.70369
FBP+GF (3+3+5 mm) | 0.70884 | 0.66084 | 0.72528 | 0.75414
A (4+4+7 mm)
B (3+4+4 mm)
C (3+3+5 mm)
(a)
(b)
Fig. 4

## Slide 26: Design of the HR-PET
OD=92 mm
44 mm
23.4 mm
Signal (18 pin)
HV&TMP(18 pin)
(a) Front view
FOV
(b) Cutaway side views
Front-end board
Amplifier board
LYSO
Mini-HDMI
OP-AMP
3D printed jig
1
2
48 mm
Notes: The SR-PET scanner has an inner diameter of 52.5 mm, and the crystal thickness of 15 mm.
The number of blocks is 16, each of which consists of 3-layer LYSO array.
The number of rings is 4, thereby resulting in an axial FOV of 51.5 mm.
The bore diameter is 48 mm, and the outer diameter is 107 mm.
The PET scanner consists of front-end board and amplifier board.
The SiPM signals and power lines were transferred by using flexible flat cables.

## Slide 27
(c) Norm. sinogram
(f) Sinogram (Phantom)
(b) Norm. activity sinogram
(a) Norm. activity image
(d) Sensitivity mask
(e) Norm. Factor
(g) Norm. Factor
(h) Normalized sinogram
(i) Interpolated sinogram
(j) FBP with GF
FBP
P
GF
20.4 mm
2 mm
r
φ
r
φ
x
y
x
y
r
φ
r
φ
r
φ
r
φ
r
φ
r
φ
Fig. 3
10 mm

## Slide 28: Resolution phantom (FBP)
0.85
0.8
0.75
0.55
0.5
0.45
0.85
0.8
0.75
0.55
0.5
0.45
0.85
0.8
0.75
0.55
0.5
0.45
5 mm
5 mm
5 mm
A (4+4+7 mm)
B (3+4+4 mm)
C (3+3+5 mm)
VPR=0.928±0.026
VPR=0.876±0.081
VPR=0.864±0.054
Notes: The resolution phantom images were reconstructed using FBP with three different crystal layer designs.
Design C provided best valley-to-peak ratio of 0.864, which was over the Rayleigh criterion of 0.752.

## Slide 29: Preclinical PET for translational research
Microscope
Find molecular biomarkers (e.g., gene, protein, metabolism)
Develop new radiotracers
Preclinical research
Rodent (Mouse, rat)
Clinical PET
Clinical study
Human
Cell
Small molecule
Peptide
Antibody
Basic research
Preclinical PET
Preclinical PET is a bridge between the basic research and clinical study
1 cm
Mouse brain
1 cm
Human brain
Notes: Let me briefly introduce about the role of preclinical PET for translational research.
In basic research, microscopes are commonly used to find molecular biomarkers in cell or DNA levels.
Then new radiotracers are developed for molecular imaging with rodent models.
Preclinical PET is used to visualize the functional information of mice for preclinical research to test the efficacy of the radiotracers.
After preclinical research, we can move on to the clinical study with human subjects.
Therefore, preclinical PET is essential for translational research as a bridge between the basic research and clinical study.

## Slide 30: PET resolution prediction and validation
Res.
0.6 mm
Levin and Hoffman, 1999 PMB
Kang et al., 2023 JNM
0.85
0.8
0.75
0.55
0.5
0.45
OSEM (50 iterations)
LYSO pitch
=1.0 mm
If =0.1 mm, D=53 mm, =1.0 mm, =0, then resolution = 0.55 mm
52.5 mm
18F

## Slide 31: Front-end electronics
Front-end board
Top
DAQ
Power
2.4 mm
SiPM5×5
A
B
C
D
Anode 25 Ch
SiPM
SiPM biasing
Cathode (+HV)
Anode (Signal)
SiPM5×5
Position
4×4=16 ch
Timing
4×1=4 ch
Rpz
Rf
-
+
Output
AD8000
Cpz
Rg
50Ω
Input
200 ns
200 mV
Timing signal
Pole-zero-cancellation circuit
Resistive network
Bottom
Amplifier board
Flexible Flat Cable
miniHDMI
14 mm
<150ns
Notes: The front-end board has SiPM5x5 array with a pitch of 2.4 mm.
The SiPM bias was applied to a common cathode and signals were extracted from anode.
The 25 anode signals were multiplexed into 4 positional signals by using a resistive network.
The multiplexed signals were transferred to an amplifier board via flexible flat cable.
The position signals were amplified with quad channel amplifiers.
The timing signals were amplified by using a fast amplifier followed by a pole-zero-cancellation circuit to obtain fast decay time of less than 150 ns.
The amplified signals were transferred to a DAQ via miniHDMI cables.
The power lines were supplied by using the same cables.

## Slide 32: Staggered 3-layer DOI detector
<Staggered 3-layer DOI detector>
Pitch=
2.4 mm
SiPM
5×5
<SiPM with front-end board>
Pitch=0.8 mm
LYSO 3-layer
Acrylic lightguide
(12.2×12.2×1 mm3)
14 mm
12.5 mm
1st layer
2nd layer
3rd layer
Notes: Two LYSO 3-layer arrays were mounted on top of the SiPM.
A 1 mm thick lightguide was inserted between the crystal and SiPM.

## Slide 33: Crystal map (Experiment)
Pitch=0.8 mm
Crystal map
The crystal map of a 3-layer LYSO crystal array with 0.8 mm pitch was well identified.
22Na
Front-end
1st layer
2nd layer
3rd layer
511 keV
1st layer
2nd layer
3rd layer
Notes: Currently, the front-end electronics are under development.
We plan to develop a full-ring PET system within this year.