## Slide 1: Status of Detector Array Infrastructure
Walter Raniero & Roberto Menegazzo
INFN-LNL / INFN-PD
26th AGATA Week – Infrastructure Group

## Slide 2
W. Raniero – AGATA Week 2026
OUTLINE
Status of Detector Agata Array Infrastructure 2025/2026
Detector Agata Array Infrastructure for ATCs > 15
Mechanics for Detector Infrastructure
New ATC laser alignment 2.0
summary

## Slide 3
Detector Array Infrastructure Status 2025/2026
W. Raniero – AGATA Week 2026
Detector and Autofill LN2 it’s ok up to 16 ATCs

## Slide 4
W. Raniero – AGATA Week 2026
Autofill LN2
Description | Available for ATC at 2026 | Number for 2 π (ATC>15) | Note
patch panel on Autofill rack to manifold | 4 + 2 boxes | 4+4 boxes | custom made
patch panel on Autofill to patch box | 15 | 30 | PT dewar, PT capsule, LN2 red-out (15 to be installed)
PT100 cable det. | 30 x 2 (dew+caps) | 30 | (15 to be installed)
PT100 cable valve | 15 | 30 | (15 to be installed)
LN2 level cable | 15 | 30 | (15 to be installed)
Valve power cable | 15 | 37 | (15 to be installed)
Cryogenic valve | 15 | 37 | Being purchased
No returne valve | 15 | 37 | Being purchased
PT 100 on return valve | 15 | 36 | Being purchased
Bayonets | 15 | 30 | Upgrade in progress
Detector Array Infrastructure Status 2025/2026
Cabling to PLC (DSS)
Software for ATC >16
Designed and produced at LNL

## Slide 5
ATC installation issues and future plans
LN2 leak during the filling on AGATA array (filling LN2 pressure <2bar)
pressure Relieve Valve to be replace. tbd. Detector group
Capacitive level read out for LN2. The out cable is glued and with overpressure there is a LN2 leak.
Maintenance needed /new design tbd. detector group
W. Raniero – AGATA Week 2026
A new design phase is being initiated for bayonet couplings to be connected to AGATA’s flexible metallic LN2 transfer lines.

## Slide 6
W. Raniero – AGATA Week 2026
HV system
Description | Available at 2026 | Number for 2 π (ATC>15) | Note
CAEN SY4527 | 3 | 2 | HV crate
CAEN A1560HDPE | 18 | 12 | HV Boards (8 ch)
SHV cable | 45 | 90 | 45 (already available)
BSD cables | 1 | 2 | PLC to HV system
Digitizer
Description | Available for ATC at 2026 | Number for 2 π (ATC>15) | Note
MDR cable | 315 | 630 | 315 (available)
Detector Array Infrastructure Status 2025/2026
HV Bias Shutdown cabling

## Slide 7
W. Raniero – AGATA Week 2026
LVPS
Description | Available for 15 ATC at 2026 | Number for 2 π | Note
AXIS LVPS New (2025) system (8ch) | 4 | 4 | 
Digitizer 48V cable | 45 | 90 | (45 already availble)
±6/12V cable set | 15 | 30 | (20 Already available)
LN2 cards | 20 | 30 | (20 purchased)
Patch box + 3 mech. part | 20 (LNL) | 30 | 20 set mech. part have been made 35 patch boxes will be purchased
Description | Available for 15 ATC at 2026 | Number for 2 π | Note
Detector legs | 18 (set) | 30 | 1 set = 3 legs from CTT
Detector legs adapter | 22 (set) | 30 | 1 set = 3 adapters IKP
Detector holding ring | 30 | 30 | 
Detector Array Infrastructure Status 2025/2026
New heat exchanger to prevent corrosion of Al plates and pipes (digitizer)
Detector Patch Box

## Slide 8
W. Raniero – AGATA Week 2026
LVPS control
Detector Array Infrastructure Status 2025/2026
LN2 software issue:
an emergency fill of an ATC shortly before the scheduled filling time causes the other ATCs to skip their filling cycle
LN2 control

## Slide 9
W. Raniero – AGATA Week 2026
Mechanics for Detector Infrastructure

## Slide 10
Jig Alignment @ LNL-INFN : ver_0.0
W. Raniero – AGATA Week 2026
Fixed aluminum masks that may damage the ATC end cap.
Time-consuming iterative alignment process using spacers.

## Slide 11
Jig Alignment @ LNL-INFN: ver_1.0
Laser Tracker assistance on setting Z (insertion in Honeycomb) coordinate JIG 1.0
Rotation made on JIG mask system
179.5 mm
JIG ref plane
ATC ref plane
Reference in Laser software
W. Raniero – AGATA Week 2026

## Slide 12
Jig Alignment @ LNL-INFN: ver_1.0
W. Raniero – AGATA Week 2026
Mask
Level and height set with laser tracker
Lateral position set with new mask

## Slide 13
Laser Tracker assistance on insertion and angle wit JIG 2.0 with alignment tool as for the AGATA Honeycomb
Jig Alignment @ LNL-INFN: ver_2.0
W. Raniero – AGATA Week 2026
ACT in Honeycomb
JIG 2.0 from below
New LT stand for ground operation
courtesy D. Scarpa

## Slide 14
ATC Legs
W. Raniero – AGATA Week 2026
46 Adapter legs
(IKP) 2026
The legs and adapter are available for all 18 ATCs at LNL (11 on the frame and 7 in the lab)

## Slide 15
Linear SLIDE BUSH
W. Raniero – AGATA Week 2026
Slide bushes were purchased to ensure the availability of 3 dedicated sets, one for each detector ring.

## Slide 16
Galileo filling LN2 system
Galileo system: host up 14 ATCs
W. Raniero – AGATA Week 2026
Maintenance: valves, insulated pipes, purge valves, PT 100

## Slide 17
Summary
The infrastructure is ready for up to 15 ATCs; further work is ongoing to support more than 15 ATCs.
The infrastructure materials required for 2π coverage are already available.
Detector-ring installation has been improved, ensuring more reliable alignment (x-y-z- rotation) between the JIG and the AGATA array using a laser alignment tool (JIG v2.0).
W. Raniero – AGATA Week 2026

## Slide 18
Thank you !!
W. Raniero – AGATA Week 2026

## Slide 19: Status of Detector Array Infrastructure
Walter Raniero & Roberto Menegazzo
INFN-LNL / INFN-PD
26th AGATA Week – Infrastructure Group