## Slide 1: Ion bunching optimization in S3-LEB and commissioning of the new quadrupole bender
Maylis Brun
PhD supervisor :
Antoine de Roubin
Outline :
Status of S3-LEB
Ion bunching optmization
Commmissioning of the new quadrupole bender
Perspectives

## Slide 2
S3-LEB commissioning perspectives
Heavy neutron-deficient refractory element region
Shape-coexistence
High-spin isomers
N = 82 shell closure
Exotic decay modes
Close to the N=Z line
N=Z / Symmetries
Shell evolution
Nucleosynthesis
Heavy (and Super Heavy) element region
Single-particle vs deformation
Octupole deformation
SHE structure (Fm, No, Lr)
http://u.ganil-spiral2.eu/chartbeams/
Notes: The S3-LEB setup will be dedicated to performing laser spectroscopy, mass spectrometry and decay spectroscopy to address different topics of nuclear structure and astrophysics, as can be shown on this slide

## Slide 3
Mini-RFQ
PILGRIM
(MR-TOF-MS)
QMF
Cooler-buncher
MCP1
MCP2
Pulse-up
MCP3 + Camera
S-RFQ
Deflector / quadrupole bender
Exotic ions coming from S3
Gas cell
A : Gaz cell to neutralize and thermalize ions coming from S3
B : Radio frequency quadrupole (RFQ) chain that guide the ions, purify and prepare the beam
C : Pulse up, to speed up ions
D : Quadrupole bender, to deliver ions to different experiments
A
B
C
D
Schéma de S3-LEB
E
E : PILGRIM : electrostatic ion trap for high resolution mass spectrometry
S3-LEB description
To SEASON

## Slide 4
Mini-RFQ
PILGRIM
(MR-TOF-MS)
QMF
Cooler-buncher
MCP1
MCP2
Pulse-up
MCP3 + Camera
S-RFQ
Deflector / quadrupole bender
Exotic ions coming from S3
Gas cell
A : Gaz cell to neutralize and thermalize ions coming from S3
B : Radio frequency quadrupole (RFQ) chain that guide the ions, purify and prepare the beam
C : Pulse up, to speed up ions
D : Quadrupole bender, to deliver ions to different experiments
A
B
C
D
Schéma de S3-LEB
E
E : PILGRIM : electrostatic ion trap for high resolution mass spectrometry
S3-LEB
To SEASON

## Slide 5
Cooler thermalize ions with buffer gas injection (Helium)
S3-LEB cooler buncher
Pulse up
Cooler-buncher RFQ
Cooler
Buncher
Helium (gas)
Notes: Le test 1 s’intéresse à la mise en paquet des ions que je présente plus en détails
Thermalise : met à température de l’hélium (équilibre thermodynamique) – collisions inélastiques
Pourquoi en deux étapes : pilgrim très sensible au gaz donc on peut pas le mettre au niveau du regroupeur (pour que le gaz fuite tôt pour garder un gaz de bonne qualité)

## Slide 6
Cooler : thermalize ions with buffer gas injection (Helium)
Buncher : creates bunches in potentiel well
ToF dispersion : ToF FWHM
Cooler-buncher RFQ
Potentiels applied to the buncher
Buncher electrodes
S3-LEB cooler buncher
Pulse up
Cooler
Buncher
Helium (gas)

## Slide 7
Cooler : thermalize ions with buffer gas injection (Helium)
Buncher : creates bunches in potentiel well
ToF dispersion : ToF FWHM
Cooler-buncher RFQ
Potentiels applied to the buncher
Buncher electrodes
S3-LEB cooler buncher
Pulse up
Cooler
Buncher
Helium (gas)
Notes: Dispersions of the bunch at different points of the line

## Slide 8
 | PILGRIM’s optimal dispersions values [1] | Simulated results [1] | Experimental results [2]
Time-of-flight dispersion (ns) | 50 | 64 | 157(6)
Energy dispersion (eV) | 15-20 | 42 | 13(1)
Resolving power for PILGRIM | 300 000 | ≈230 000 | ~80 000
[1] :HAL Id: tel-01469429, https://in2p3.hal.science/tel-01469429v1
[2] : HAL Id: tel-04483471, https://theses.hal.science/tel-04483471v1
Table of time and energy dispersions for S3-LEB bunches
Motivation : PILGRIM’s resolving power
Notes: Dispersion attendue autour de 100 ns (entre simulation et exp préc),
PILGRIM encore + compétitif
Autres : ISOLTRAP 300 k

## Slide 9
PILGRIM
(MR-TOF-MS)
Cooler
Buncher
MCP3
Helium
Trapping zones in S3-LEB
Ion bunching optimization

## Slide 10
PILGRIM
(MR-TOF-MS)
Cooler
Buncher
MCP3
Proposition
Pre bunching before the buncher
Why ? : identical time spent in the buncher and no more continuous composante of the beam through the buncher when open
Mesurements for different duty cycle duration
Helium
Trapping zones in S3-LEB
Ion bunching optimization

## Slide 11
PILGRIM
(MR-TOF-MS)
Cooler
Buncher
MCP3
Proposition
Pre bunching before the buncher
Why ? : identical time spent in the buncher and no more continuous composante of the beam through the buncher when open
Mesurements for different duty cycle duration
Helium
Trapping zones in S3-LEB
Ion bunching optimization

## Slide 12
PILGRIM
(MR-TOF-MS)
Cooler
Buncher
MCP3
Proposition
Pre bunching before the buncher
Why ? : identical time spent in the buncher and no more continuous composante of the beam through the buncher when open
Mesurements for different cycle duration
Helium
Trapping zones in S3-LEB
Ion bunching optimization

## Slide 13
Comparison of the different trapping methods
Without pre bunching : Time-of-flight dispersion decrease with cycle duration
Pre bunching in cooler : lower fwhm average : longer cooling with gas : ions are better prepared to be bunched
Results : ToF dispersion
FWHM with cycle duration
No pre bunching
Pre bunching in MRFQ
Pre bunching in cooler
Cycle duration (ms)
ToF FWHM (ns)

## Slide 14
Comparison of the energy FWHM for different bunches energy.
Results : Energy dispersion
Energy of the ions (V)
Evolution of the energy FWHM with the ions’ energy
Skyy Pineda analysis

## Slide 15
 | PILGRIM’s optimal dispersions values [1] | Simulated results [1] | Experimental results [2] | Pre bunching results
Time dispersion (ns) | 50 | 64 | 157(6) | 124,1(0,29)
Energy dispersion (eV) | 15-20 | 42 | 13(1) | 4(1)
Resolving power for mass measurements | 300 000 | ≈230 000 | ~80 000 | ≈120 000
Table of time and energy dispersions for S3-LEB bunches
Results
50 ms cycle duration
Notes: Gagne peu en temps de vol à 50 ms mais plus en énergie

## Slide 16
Pictures of the quadrupole bender installed in S3-LEB beam line
Quadrupole bender (deflector) installation
Quadrupole bender
SEASON Arm
Quadrupole bender
SEASON Arm

## Slide 17
Quadrupole bender current design
Deflector
Arm SEASON
Representation of the deflector, Beam Imaging Solutions, Instructions _ Quadrupole deflector

## Slide 18
Transmission efficiency measurement
Installation of discriminators on MCP1 and MCP2
Installation of new *40 amplifiers for all detectors
Setting of the acquisition threshold at 0,2V for all detectors
Notes: Discriminators : prevent double counting in continuous beam without reducing the resolution (not bunched). It was installed carrefully to suppress noise from the count (trigger outside the noise)
*40 amplifiers to reduce the ion count on the acquisition system
200 threshold : no obvious plateau in the curve presented so it was set according to was was seen on the oscilloscope and previous experience with the detectors.

## Slide 19
Transmission efficiency measurement
Mini-RFQ
PILGRIM
(MR-TOF-MS)
Cooler
Buncher
MCP3
MagneToF Pil
MCP injection
MagneToFs SEASON
1
2
MCP1
MCP2
100%
100%
81%
50%
70%
Notes: Qualitative / relative estimation of transmission efficiency : study with radioactive source to be done later in the year.

## Slide 20
Bender’s influence on ToF FWHM
Cycle duration (ms)
Bending towards SEASON
Going straight to PILGRIM
ToF
ToF
Notes: No bunch deformation for ToF (can measure energy)
FWHM wider after pilgrim

## Slide 21
Working towards the real transmission efficiency
Buid a complete SIMION simulation of S3-LEB
Confirm the bender’s design
Perspectives

## Slide 22
Thank you for your attention

## Slide 23: Appendices

## Slide 24: Appendice 1 Detectors installation

## Slide 25: S3 – Super Separator Spectrometer
Fusion-evaporation products
(with refractory elements)
Proton dripline & nuclei
Nuclear structure
          Study the role of correlations
                Deformation – shape coexistence
Exotic decay
Astrophysics rp-process
Fundamental interaction
Heavy and Superheavy Elements
Limit of the nuclear existence
Nuclear structure
Reaction mechanisms
Atomic properties
Courtesy H. Savajols
40Ar+209Bi244-247Md+xn
4evt/day (244Md)
58Ni+50Cr 100Sn + 𝛂 + 4n
1/s
40Ar + 116Sn 152Er +4n
104/s
20Ne+197Au 210Ac + 7n
200/s
http://u.ganil-spiral2.eu/chartbeams/

## Slide 26: Annexe 4 Fit de données sur python

## Slide 27: Annexe 6 Interface EPICS

## Slide 28
S3-LEB duty cycle illustrated without trapping
Without trapping
With trapping
Bunching cycle optimization
Buncher opening
Trap opening
Buncher opening
Cycle duration (ms)
Cycle duration (ms)
Pulse up acceleration
Time spent in the buncher
Trapping duration
S3-LEB duty cycle illustrated with trapping
Animation