## Slide 1
“GPIB”
commissioning for DESIR
Corentin Roumegou
ISOL-France VIII – 16/03/2026
BRIoCHe

## Slide 2
“GPIB”
commissioning for DESIR
Corentin Roumegou
ISOL-France VIII – 16/03/2026
?
?
BRIoCHe

## Slide 3
fusion-evaporation
S3
SPIRAL1
fragmentation
fusion-evaporation
LUMIERE
DETRAP
BESTIOL
Contaminated
Continuous
High-emittance
Low-energy RIB from SPIRAL1 and S3
The DESIR facility
@
Beam is possibly:
SUSHI ?
« GPIB »
Décroissance, Excitation et Stockage d’Ions Radioactifs

## Slide 4
Beam preparation: cooling & bunching
RadioFrequency Quadrupole Cooler - Buncher (RFQCB)
Reduce beam emittance
Bunch the beam
Bunching Radioactive Ions To DESIR
General Purpose Ion Buncher
Buncher of Radioactive Ions Cooled with Helium
BRIoCHe
(
(
)
)

## Slide 5
Bunching Radioactive Ions To…
RFQCB
MR-ToF-MS
PO-ToF-E
(NA)2STARS
* Non-exhaustive list. The exact location of each setup may change.
** The beam could also pass through PIPERADE, Laser Cooling, etc.
BRIOCHe

## Slide 6
RFQ: Paul Trap
Transverse trapping with 4 unsegmented RF rods
Longitudinal trapping with 25 independent DC segments
800 mm
ions
Mechanical design based on ISCOOL (ISOLDE RFQCB)
RF
DC
“GPIB”
General Purpose Ion Buncher

## Slide 7
Trap RF system
400W RF amplifier
Waveform generator
Balun-transformer circuit
Tunable capacitor
S32-32 windings
(exchangeable)
S32-32
S38-38
S24-24

## Slide 8
Cooling
Radioactive ions
From SPIRAL1/S3
Cooled beam
to DESIR
“Buffer-gas cooling” with He (highly purified!)
Up to 90% transmission
He
MoRI-To mode (Continuous)
BuRI-To mode (Bunched)

## Slide 9
ions
Bunching
Continuous beam
from HRS
Cooled & Bunched
ions to DESIR
5 last electrodes can be switched
switchable

## Slide 10
Bunching
Modular trapping/extraction curves
σ < 60ns
(FWHM < 140ns)
Favor either time or energy dispersion
made with
and
σ > 500ns
(FWHM > 1µs)
Ions on MCP detector

## Slide 11
Energy dispersion measurement
MagneTOF with retarding grid
Scan GRID1 voltage
GND
Grid1
Grid2
MagneTOF
Counts vs Grid Voltage
Derivative (point to point difference)
Gaussian Fit
COUNTS ->
VOLTAGE ->
Made with

## Slide 12
( )
Energy dispersion measurements: GRID2
NB: GRID2 is used to limit “micro-lensing” effects on GRID1
MagneTOF
Remark…

## Slide 13
and measure with the
Faraday Cup downstream
Energy dispersion measurement : Issues
Difficulties with MagneTOF measurements
Get 20 eV energy dispersion for any settings…
GND
Grid1
Grid2
MagneTOF
Faraday Cup
Take out MagneTOF
Needs higher source current !
17x attenuation
Made with

## Slide 14
Energy dispersion: measurements with FC
GND
Grid1
Grid2
Faraday Cup
Much better, but still large dispersions…
MoRI-To mode
(Continuous Beam)
Shapes not gaussian, Low energy tail, …
BuRI-To mode
(Bunched Beam)
- 5 bunches/s
(1e7 trapped ions/bunch)
- 32 bunches/s
- CW
Made with

## Slide 15
Cooling: transverse emittance reduction
Measured in Bordeaux with electrical sweep type emittance scanner, for a 30keV continuous beam
Work in progress for bunched beam measurements with a “Pepper-Pot” emittance-meter (<-> HRS)
Hardware: new 2D mask installed, camera trigger, …
Software: produce emittance figures, analysis for absolute emittance value, bunches/images accumulation… [A. Balana]

## Slide 16
Bordeaux
Stable-ion source
Conclusion & Perspectives
The RFQCB is fully operational
soon...
New place, New life, New Name…
Moving to GANIL
Caen
Work is ongoing to finish full characterization and improve bunched beam performance
Stable-ion source
(Alkali)
[GPIB]
Delivers bunched beam 24/7 @LP2iB
for PIPERADE commissioning
Working on DESIR upgrade: HV platforms, mechanical support, automation, …

## Slide 17
Thank you !
P. Alfaurt, P. Ascher, D. Atanasov, A. Balana, B. Blank, L. Daudin, M. Flayol, M. Gerbaux, S. Grévy, G. Guignard, M. Hukkanen, A. Husson, G. Jevelot, B. Lachacinski, S. Lechner, E. Rey-herme, M. Roche, K. Ronxin, A. de Roubin, C. Roumegou
GPIB-PIPERADE Team
No AI in this picture ;)
BRIoCHe

## Slide 18
BACKUP SLIDES

## Slide 19
The DESIR facility
@
Low-energy RIB
from SPIRAL1
and S3
SPIRAL1
SPIRAL2
S3
DESIR
fusion-evaporation
S3
SPIRAL1
fragmentation
fusion-evaporation
New experimental hall in GANIL
09-2025
Now…
…and in the future

## Slide 20
Décroissance, Excitation et Stockage d’Ions Radioactifs
Physics at DESIR
Decay, Excitation and Storage of Radioactive Ions
Decay Spectroscopy
LASER Spectroscopy
Traps
Exotic decay modes
High-precision β-decay studies
TAS measurements
Charge radii
Nuclear moments
Spins
Mass measurement
Fundamental interactions
Trap-assisted spectroscopy
LUMIERE
DETRAP
BESTIOL

## Slide 21
Beam preparation: mass separation
U-shape mass separator
[HRS] SUSHI : SPIRAL U-shape Separator for Heavy Ions
Currently commissioning in Bordeaux
J. Michaud et al., Nucl. Instrum. Methods Phys. Res. B (2023)

## Slide 22
800 mm
The “GPIB”: an RFQ Paul Trap
Transverse trapping with 4 unsegmented RF rods
Longitudinal trapping with 25 independent DC segments
z
y

## Slide 23
Function generator
20 dB
6 dB
Primary
winding
RFQ
Power
amplifier
Adjustable
ground
Tunable
capacitor
Secondary
windings
Directional
coupler
Osc. ch1
Osc. ch2
Osc. ch3
Osc. ch4
attenuator
attenuator
GPIB RF system

## Slide 24
GPIB mechanical design
He
RF feeding
Grounded
electrode
Grounded
electrode
RF rods
DC segments
Pumping
Alignment system

## Slide 25
P3: PythonPlotPot
Plot potential curve on a given axis of beamline
Check current beamline settings and create new ones
Interface with SIMION python PA library
Adjust electrodes voltages

## Slide 26
PTSD
Scan all EPICS variables
Control trapping sequence
Signal acquisition
Position reconstruction
Data gating and fitting
Adaptable to various DESIR detector type
Python Trap Scanner for DESIR
Already compatible with 2 MCP types (simple ToF signal or position-sensitive with delay lines), Faraday Cups and MagneToF - For Continuous or Bunched beam.