## Slide 1: Neutron Field Spectral and Fluence Characterization and BSA optimization for AB-BNCT
Daniel Santos, Sébastien Chabod, Oliver Guillaudin,
Nadine Sauzet, Alexandre Bes
Laboratoire de Physique Subatomique et de Cosmologie (LPSC)
(Université Grenoble-Alpes -CNRS/IN2P3)
Enya Mobio, Amokrane Allaoua
Laboratoire de Métrologie et Dosimétrie Neutronique (LMDN)
ASNR(ex-IRSN)-France

## Slide 2: How can we measure a neutron energy produced by a non-pulsed beam ?
From θ, Eion, and the IQF,
we get the neutron energy En
Neutron
En
Field cage
Cathode
Pixelated anode
HV1
HV2
θ
Nuclear recoil
e -
e -
e -
e -
3D recoil track θ
Energy deposited in ionization : IQF
Eion
20 ns sampling :
Grid
3D-nuclear recoil camera + Ionization Quenching Factor (IQF)

## Slide 3: Demonstration of the directional detection with mono-energetic neutron fields
- The AMANDE facility (ASNR (ex-IRSN)-Cadarache)
provides mono-energetic neutron fields
- The energy of a mono-energetic neutron field is
defined by the angle of each nuclear recoil track with
respect to the neutron direction…
But the ionization energy of the recoil is different of its
kinetic energy…
It is quenched !!

## Slide 4: At low energies… 27 keV neutron field Proton and carbon recoils give the same neutron energy ! C. Beaufort et al 2024 JINST 19 P05052, arXiv:2312.12842
The kinetic energy is reconstructed with the Ionization quenching factor measured by COMIMAC
Angular distribution
Neutron spectrum

## Slide 5: A mono-energetic Epithermal neutron field (8 keV) Neutron spectrum reconstruction from proton recoils
Neutron spectrum
Angular distribution
Directional performances at 8 keV
Cyprien Beaufort et al. JCAP08(2022)057

## Slide 6: MIMAC-FastN - Polyenergetic measurement of 9Be(d(1.45 MeV),n)
Spectrum measured at 0 deg
Spectrum measured at 60 deg
M.E. Capoulat, N.Sauzet et al.
« Neutron spectrometry of the 9Be(d(1.45 MeV),n)10B reaction for accelerator-based BNCT » NIM B, vol. 445, pp. 57-62, 2019
Angular distribution
for an AB-BNCT nuclear reaction
INFN LNL
(Legnaro - Italy)
700 mbar
He/CO2 (5%)
Collaboration with TANDAR (CNEA-Argentina)

## Slide 7: Energy calibration, Thermal neutron fluence and Active BNCT Phantom
Li
840 keV
Li
1,16 MeV
α
1,47 MeV
α
1,78 MeV
Alpha track
Detection of α and 7Li, resulting from thermal neutron capture on a B4C layer
Anode projection : track start
B4C layer
on the cathode
(500 nm thick)

## Slide 8: Experimental Set-up (schematic) to test the Active-Phantom Mode
MIMAC-FastN
D2O-15cm radius Sphere
Moderator
5 cm thick

## Slide 9: Experimental Set-up with a 252Cf source modered by a sphere of D2O
Set-up at the IRSN-Cezane facility at Cadarache (France)
Bruno Asselineau, Michaël Petit, Thibaut Vinchon

## Slide 10: Geant4 Simulation of the MIMAC-FastN-Active Phantom
Clear separation of the alpha and 7Li branches with the both components each
4He
7Li
Track length totally included in the C4H10 + 50% CHF3 at 30 mbar

## Slide 11: Geant4 Simulation showing all the recoils produced and stopped inside the active volume
The proton background is easily discriminated with respect to the alpha and lithium nuclear recoils in C4H10 + 50% CHF3 at 30 mbar
1H

## Slide 12: Neutron Capture on 10B with MIMAC-Active Phantom mode
Data from the MIMAC-Active Phantom using the
Experimental Set-up shown before with
C4H10 + 50% CHF3 at 30 mbar
Track length (mm) vs. Ionization Energy (ADC channel)
4He
7Li
1H
e-
Ionization Energy

## Slide 13: A complex epithermal neutron field characterization
At AMANDE facility, with a near threshold 7Li (p (E=1.881 MeV),n) reaction
Chart: Thickness of the PEHD moderator (cm)
| 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 6
Fluence (n/s) à 1uA à 5cm (fenetre NFM) | 102109.5566 | 64479.7935 | 44990.6568 | 34299.3185 | 26199.6394 | 21406.6554 | 17429.8206 | 11358.1683
Thermique_1881 | 0 | 0 | 0 | 0.01 | 0.03 | 0.07 | 0.15 | 0.24 | 0.35 | 0.45 | 0.54 | 0.69 | 0.8 | 0.88
Thermique_1883 | 0 | 0 | 0 | 0.01 | 0.04 | 0.09 | 0.16 | 0.26 | 0.36 | 0.46 | 0.55 | 0.69 | 0.8 | 0.88
Epithermique_1881 | 0 | 0.01 | 0.03 | 0.08 | 0.16 | 0.26 | 0.35 | 0.41 | 0.42 | 0.41 | 0.37 | 0.28 | 0.19 | 0.12
Epithermique_1883 | 0 | 0.01 | 0.04 | 0.11 | 0.2 | 0.3 | 0.39 | 0.44 | 0.44 | 0.42 | 0.38 | 0.28 | 0.19 | 0.12
Rapide_1881 | 1 | 0.99 | 0.96 | 0.91 | 0.81 | 0.67 | 0.51 | 0.35 | 0.23 | 0.14 | 0.08 | 0.03 | 0.01 | 0.01
Rapide_1883 | 1 | 0.99 | 0.95 | 0.88 | 0.76 | 0.61 | 0.44 | 0.3 | 0.2 | 0.12 | 0.07 | 0.03 | 0.01 | 0.01
C. Beaufort et al.
JINST, P05052 (2024)

## Slide 14: Neutron flux measurements by the Gold foil Activation method and by the Neutron Flux Monitor (NFM) detector
Foil of 197Au + Cd
Foil of 197Au
NFM detector
Dépôt de 10B
13 mm
Distribution of the first measurement of the 3D-tracks
FADC profile of an alpha (1.24 MeV)
event as a function of the
sampling (20 ns)

## Slide 15: Alpha and 7Li energies measured with the NFM
(difference < 1%)
Enya Mobio, Ph.D. Thesis (April 2025)

## Slide 16: Thermal and Epithermal Fluence Results
 |  | Фth (n.cm-².s-1) | C/E | Difference | Фepi (n.cm-².s-1) | C/E | Diff
197Au | Calcul | (7,52 ± 0,32).102 | 0,68 | - 46% | (1,06 ± 0,05).103 | 1,25 | +20%
 | Exp. | (1,10 ± 0,08).103 |  |  | (8,53 ± 0,84).102 |  | 
NFM | Calcul | (5,66 ± 0,25) .102 | 0,94 | - 6 % | (9,08 ± 0,40).102 | 0,94 | - 6 %
 | Exp. | (6,03 ± 0,34).102 |  |  | (9,64 ± 0,55).102 |  | 
From the Activation of Gold foils and the Neutron Flux Monitor detector
compared with the MCNP simulation calculation
Enya Mobio (LMDN (ASNR)), PhD thesis (April 2025)

## Slide 17: NFM (Neutron Flux Monitor) Detector
Giving the neutron flux (up to 1012n/s)
as a function of time, with an evaluation of the dead time
Including a Boron layer very well measured (thickness, surface and homogeneity !!
Multi-layer deposition PAPVD (plasma-assisted)
LPSC – Alexandre BES, Ana LACOSTE

## Slide 18
What criteria does the community use to design BSAs?
Method 1: in the past, BSA often designed so that the neutron field at its exit verified some
in-air FOMs. E.g. fepi > 5e8 cm-2s-1 (intensity), Dfast/fepi < 2e-13 Gy.cm2 (purity), etc.
Issue: in-air FOMs found too restrictive, can eliminate BSAs with best therapeutic performance
IAEA TECDOC 2001: in-air FOM “values […] are specifically not suggested as ‘requirements’
or ‘recommendations’ that have to be achieved by a system provider’
 method 1 becoming obsolete
Method 2: a therapeutic target is chosen, often a GB of a few cm3 placed at a depth of ∼3-4 cm
in a patient’s head. BSA then designed to maximize Dtum/Dbrain or a TCP.
Issue: arbitrariness of the choice of tumor configuration
 expected design suboptimality for treating other tumor configuration
Method 3: BSA designed to maximize dosimetric FOMs, most often AD
(Advantage Depth) or TD (Treatment Depth).
TD gives a range of depth in body where tumor tissue receives at
least 2 times the max. dose in normal tissue
 common methodology, extensively explored in literature
BSA design methodologies
Topology Optimization of Neutron Moderators

## Slide 19
Breakthrough in recent years: recent developments in algorithms (machine learning, topology
optimization, etc.) now allow unbiased computations of the structures of BSAs.
Ex of result: TopOpt computation of the best BSA structure that max TD s.t. TT < 1h for GB
Modern BSA design calculation methods
S. Chabod et al., Phys. Med. Biol. 2025, 70 (3), pp.035008
Topology optimization
= compute the best
material distribution
in each voxel of a BSA
Optimization result
7Li( p(2,5 MeV) ,n) metallic source

## Slide 20
Progress in the dosimetric performance of BSAs over the past 30 years
Results: AD increased by only 1.1 cm in 20+ years, TD by only 0.05 cm.
State of the art of BSA design using method 3
Study | Source | CB (ppm) | AD (cm) | TD (cm)
Bleuel et al. 1998 | BMR Reactor 7Li(p(2.4 MeV),n) | 13 13 | 8.4 9.9 | NA NA
Riley et al. 2003 | MIT Reactor | 18 | 9.7 | 7.8
Kononov et al. 2004 | 7Li(p(2.3 MeV),n) | 18 | 9.1 | 7.0
Kim et al. 2009 | 7Li(p(2.5 MeV),n) | 10 | 8.1 | 6.0
Tanaka et al. 2009 | KU Reactor KU Reactor 9Be(p(30 MeV),n) 9Be(p(30 MeV),n) | 15 18 15 18 | 8.4 8.6 8.9 9.1 | NA NA NA NA
Inoue et al. 2014 | 9Be(p(8.0 MeV),n) | 10 | 7.5 | 5.0
Minsky et al. 2014 | 7Li(p(2.3 MeV),n) | 15 | 9.5 | 7.0
Capoulat et al. 2014 | 9Be(d(1.45 MeV),n) | 15 | 9.6 | 7.8
Capoulat et al. 2017 | 13C(d(1.45 MeV),n) | 15 | 10.0 | 7.8
Koay et al. 2018 | W(p(50 MeV),n) | 18 | 9.3 | NA
Torres‑Sánchez et al. 2021 | 7Li(p(2.1 MeV),n) | 18 | 9.74 | 7.85
Bae et al. 2022 | 9Be(p(10 MeV),n) | 18 | 9.0 | 7.2
Zhang et al. 2023 | 7Li(p(2.3 MeV),n) | 18 | 11 | NA
Rong et al. 2024 | 7Li(p(2.8 MeV),n) | 18 | 8.24 | 6.5
Verdera et al. 2024 | 45Sc(p(2.918 MeV),n) | 18 | 9 | 7
Chabod et al. 2025 | 7Li(p(2.5MeV),n) | 18 | 12.7 | 10.3
CB = Boron-10 concentration in blood
NA = value not published and cannot be
reconstructed in the absence of published
dose profiles

## Slide 21
Modern BSAs generate complex neutron fields j(r,E,W), that combine directionality (W),
space (r) and energy (E) distribution of neutrons to better control tumors & side effects.
Ex: field generated by a BSA that max TD for a 30 mA 7Li(p(2,5 MeV),n) source
General problem: how to characterize such BNCT neutron fields?
Characterization of advanced neutron fields
Neutron
Filter
Neutron
Tunnel
Ring neutron
collimator
Field intensity
Field directionality
Field performances
AD = 12.7 cm
Previous best value
obtained so far : 11 cm
TD = 10.3 cm
Previous best value
obtained so far : 7.8 cm
Dtum @ 4 cm depth
= 206 Gy-Eq/h
S. Chabod et al., Phys. Med. Biol. 2025, 70 (3), pp.035008

## Slide 22: Neutron Field (NF) characterization (E, flux) and the BSA used to define it are now possible !
1- Spectral and fluence measurements of the NF produced by a non-pulsed particle beam
2- Monitoring the fluence as a function of time giving access to the eventual degradation of the target and a real-time dose calculation
3- Standardization of measurements will give access to the comparison of the different NF produced by the different facilities and the doses delivered
4- An active phantom allowing measurements of the number of captures on a controlled amount of 10B at different “depths” is available
5- Optimal Moderators (BSA) are now possible…
A new generation of AB-BNCT can start !!

## Slide 23
Each dose component of the total dose is associated to a different nuclear reaction:
10B(n,α)7Li a neutron capture reaction to the called boron dose DB,
considering the ionization energy released by the alpha particle and lithium nucleus
1H(n,n’)1H an elastic collision reaction to the called neutron dose Dn,
considering the ionization energy released by the hydrogen resulting from
the elastic neutron scattering,
14N(n,p)14C a neutron capture reaction to the called proton dose Dp,
considering the ionization energy released by the proton and the carbon nucleus,
1H(n,γ) 2H a radiative capture reaction on 1H plus the residual gamma from 10B capture reaction
and all the gamma rays coming out from the BSA…
to the called gamma dose Dγ, considering the energy transferred by the photon by pair production, Compton, and photoelectric processes.
Total absorbed dose (DT ) in Gy (Joule/kg) = (Total Ionization released in a volume of mass m) / (mass m)

## Slide 24
Biological Dose in Gy[RBE]
Taking into account the Relative Biological Effectiveness (RBE)
Jeffrey A Coderre et al, Derivations of relative biological effectiveness
for the high-let radiations produced during boron neutron capture irradiations of the
91 rat gliosarcoma in vitro and in vivo.
International Journal of Radiation Oncology* Biology* Physics, 27(5) :1121–1129, 1993.

## Slide 25
Dose [Gy]= Ionization Energy released [Joule] / mass unit [kg]
Eγ = 2.2 MeV
Elastic collisions !!
10B (n, 4He)7Li
Neutron Capture
14N (n, p) 14C
Neutron Capture
« Total absorbed dose »

## Slide 26: Selection of 4He nuclear recoils : D(d(1.8 MeV),n)
16O(n,α)13C
12C and 16O recoils
4He recoils
Protons
Discrimination from protons, 12C, 16O, and (n,α) reactions
IRSN /AMANDE (Cadarache)
700 mbar He/CO2 (5%)

## Slide 27: Monoenergetic measurements : detection of target pollutions
D(d(1.8 MeV,n) : neutrons of 5 MeV
12C(d,n)
D(d(1.8 MeV),n)
NPL /(UK)
700 mbar He/CO2 (5%)
IRSN / AMANDE (Cadarache)
12C(d,n)
16O(d,n)
D(d(3.2 MeV),n)
Counts
Neutron energy (keV)
D(d(3.2 MeV,n) : neutrons of 6.5 MeV
The targets have in general pollutions not well determined generating neutrons not taken into
account in simulations… We have to characterize the neutron field produced at the target

## Slide 28: Spectral distribution after neutron moderation
Spectral distribution (MCNP simulation) of the neutron flux
coming out the moderator (3,5 cm thickness)
Chart: Energie (MeV)
Thermal
Epithermal
Fast
26%
44%
30%
E. Mobio et al.( 2025)

## Slide 29
Multi-layer deposition PAPVD (plasma-assisted)
LPSC – Alexandre BES, Ana LACOSTE
Cathode availables Al et Cu
Roughness between 1 et 3 µm
Ø 10 mm
Ø 20 mm
Ø 80 mm
Masks of 10, 20 and 80 mm diameter
Experimental Set-up
Production of a layer of 200 ± 5 nm of natB
+ 10 nm of Ti on an Aluminium cathode
Ø 80 mm